## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2001-10

Agency: OSAP

Location: Strait of Georgia

Project: SoG

Party Chief: Diane Masson

Platform: CCGS VECTOR

Date: April 17, 2001 – April 21, 2001

Processed by: Germaine Gatien

Date of Processing: June 14, 2001 – 2 August, 2001

Number of original CTD casts: 79

Number of casts processed: 78

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0443) was mounted with Transmissometer 333DR and SeaPoint fluorometer S/N 222. The deck unit was Seabird 11+, s/n 0425.

SUMMARY OF QUALITY AND CONCERNS

There were a lot of spikes in the data. The fluorometer and transmissivity channels were most seriously affected by this. There was little corruption by shed wakes since the descent rate was kept high.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained. Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There is a difference between the pressure calibration and that on file, but the values in the configuration file produce far better results than those on file which lead to surface values significantly less than zero for both down and upcasts. The same con file was used for the whole mission.

The sensor history was found.

- Conversion of Raw Data

All data was converted using configuration file 21100001.con. As noted in the log there is no data present in cast #47 so this was deleted. Cast #48 is the upcast portion of cast #46.

An initial examination shows all expected channels present. Upcasts and downcasts look reasonably similar but there is a lot of noise in the conductivity channels with severe spiking in all channels including pressure in cast #46. The salinity differences look unusually large. Transmissivity and fluorescence signals are present but there are problems with spiking. The dark value for FL is about 0.15g/l which is consistent with the gain entered in the con file (10X).

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

The deck unit used for this mission advances the primary conductivity channel only by +0.073s. The same equipment was used during cruise 2001-13 in early May 2001. At that time it was found that a net alignment of +0.060s worked best for both channels. Tests were done on casts #1 and 55 and it was found that an advancement of +0.050s did the best job of removing salinity spikes for both channels. So the primary conductivity will be adjusted using –0.023s and the secondary by +0.050s for a net advancement of +0.050s for both sensors.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were for all casts except #46 were:  Pass 1, Std. Dev. = 2; Pass 2, Std. Dev. = 5; Points per block = 50

For cast #46 (which contained very bad spikes) three runs were made of WILDEDIT using parameters 2,10,12; 2,10,100; 2,5,50.

7. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

9. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Max Depth | T (C) | C (units) | S (units) | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 170 | ~-.001 | -0.0023 | -0.025 | Steady |
| 2 | 220 | ~+.001 | -0.0024 | -0.026 | Good |
| 7 | 256 | Very noisy | -0.0025 | -0.028 | Very Noisy |
| 42 | 319 | Very noisy | -0.0021 | -0.022 | Good |
| 55 | 278 | Very noisy | -0.0020 | -0.021 | Good |
| 76 | 370 | ~0 | -0.0021 | -0.022 | Good |

The differences in temperature are extremely noisy but average within +/- 0.001C. The differences in conductivity and salinity are much higher than usual. During cruise 2001-06 the same instruments were used at up to 4000db and the differences were found to be –0.0008C, -0.0005units and –0.005units. During 2001-13 the average salinity differences were about –0.005units but the differences were on the order of –0.012 for early casts and lower later. This was attributed to a drift in the secondary conductivity.

10.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Errors in headers in some files had to be corrected before this step was successful.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

11. Checking Headers

The cruise track was plotted and no errors found.

A header summary and header check were produced and two errors in station names corrected.

CLEAN was run to replace pad values in Pressure with interpolated values. The surface check was run. The average surface pressure is 3.0db.

Plots of all casts were examined (on-screen) to check for problems. The transmissivity data looks bad for casts #1-4 with large blocks of near-zero values in the top 100m; the problem occurs in both up and downcasts but not at the same depths. For cast #67 the transmissivity values are all very low. The CTD log notes that there was grease on the lens. Transmissivity will be deleted from casts #1,2,3,4, & 67.

Chl_Fluorescence also shows blocks of near-zero values in the top 10m for many casts, and there are also blocks of full-scale values (15ug/l).

The primary temperature channel looks wrong in the top 10m for cast #14.

12. DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS: A warning for cast #46 was checked and found to correspond to a region of pressure spiking. The worst spikes were removed by WILDEDIT but some problems remain.

The maximum depths of all casts were checked against those recorded in the CTD log. Two casts were then checked to make sure no valuable data had been lost. In one case the log appears to be wrong and in the other the log notes that the CTD hit a muddy bottom, so the data at the bottom was properly deleted.

All DEL files were copied to EDT files.

13. COMPARE

COMPARE was used to find differences between bottles and CTD values for both salinity channels. Values from casts #46 and #58 were rejected, the former because there were problems with spiking and with bottles closing during that cast and the latter because there seemed to be a very large scatter suggesting that problems might have occurred then as well. The trendlines were much flatter after removing those data points. COMPARE indicates that the primary salinity is low by about 0.003units and the secondary by about 0.025units. The scatter in the comparison is large, particularly for the few bottles below 300db. The deepest bottle is at 375db.

There appears to be a temporal drift in the secondary sensor similar to that noted during the cruise that followed (2001-13). The values vary from about –0.03 at the beginning to about –0.015 at the end. The differences at the beginning of 2001-13 were about –0.01 at the beginning drifting to near zero at the end.

The primary sensors have smaller differences and no suggestion of temporal drift. However, they are not consistent with the observations of 2001-13 when there were deeper bottles. At that time the primary sensors were found to produce salinity that was high by +0.0033units.

- DETAILED EDITING

The primary sensors will be used for further processing.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. While there were many spikes to be removed, there were few records corrupted by shed wakes even when the descent rate was very noise. This is probably because the average descent rate was kept quite high.

T and/or S were edited near the surface and/or bottom for the following casts: #2,14,15,19,22-24,27, 52,54,79.

More extensive editing was done to the following casts: #1,5-7,25,31,34,36,37,39,42,45,46,49,50, 55-72,74-78.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Plots of temperature and salinity vs record # for the BOT files were examined to check for any bad values in these files. A few points in secondary salinity were cleaned in casts #7 and #55.

- MEDIAN FILTER

A median filter of width 11 was applied to the fluorescence channel.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

17. Other comparisons

Previous experience with these sensors – Since the latest calibration the secondary sensors were used during 2001-06 in Feb. 2001 and both pairs of sensors were used for 2001-13 in May 2001. For 2001-13 the primary sensors gave salinity high by 0.0033units and the secondary gave values low by 0.0078units. A temporal drift was noted for the secondary sensors from –0.01units of salinity at the beginning to near zero at the end. Those sensors were also marked by a lot of spiking during 2001-13. During 2001-06 the secondary sensors gave salinity high by +0.002units, but there were very few bottles for that mission and a lot of scatter.

Historic ranges – Plots were made of a few casts where ranges were available and all data fell within the ranges except for a few casts in Haro Strait for which the salinity close to the bottom was high. (See plots in Processing section of the report.)

18. Recalibration

As discussed in section 13 it is not clear how to recalibrate the primary sensors. The results for both 2001-10 and 2001-13 seem stable with time and yet the former has salinities low by about 0.003units and the latter high by about the same amount. Above 500db both comparisons show a lot of scatter and the differences for 2001-13 at those depths were not included in COMPARE. There were problems with data spikes, bottles not closing properly and grease from the wire on bottles and on the transmissometer lens during the latter part of 2001-10. Thus it was decided not to recalibrate the primary salinity. Depending on which comparison is correct this would introduce an error in salinity of 0.003units.

File 2110rcal.ccf was used to recalibrate the data (including the rosette files) as follows:

- primary salinity – no change

- secondary salinity – increase by 0.026units

- fluorescence – decrease by 0.15ul/kg

Compare was rerun to check on the recalibration and the average difference from bottles for secondary salinity was –0.00002units.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

20. REMOVE and REORDER

The following channels were removed from all casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The transmissivity channel was removed from casts: 1,2,3,4,67.

The channels were reordered and data format corrected as needed.

The final data files were named CTD.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed.

The final rosette files were renamed RAC.

21. Producing final files

a.) A cross-reference listing was produced.

b.) The sensor history was updated.

### Particulars

Cast #33 – Log notes profiles very different

Cast #37 – Hit bottom, mud.

Cast #46/47 – Problems with spiking during downcast. Bottles did not close on upcast and cast was restarted as cast #47 in an attempt to fix the problem. This did not help. Cast #46 was put through WILDEDIT 3 times to despike it.

Cast #48 – Upcast only. DELETE.

Cast #49 – Problems with bottles; only 4 closed and there could be some confusion over bottle numbers.

Cast #51 – Change to port-side winch.

Cast #50 – spikes, problem with bottles and turning pump off.

Cast #55 – Hit bottom, grease from wire on top of bottles.

Cast #67 – Grease on transmissometer lens

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 10

Dates: Start: 17 April 2001 End: 21 April 2001

Location: Strait of Georgia/ Juan de Fuca Strait

Vessel: CCGS VECTOR

Party Chief: Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2001-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 02/11/00 | Factory |  |  |
| Conductivity | 1763 | 31/10/00 | “ |  |  |
| Secondary Temp. | 2023 | 01/11/00 | “ |  |  |
| Secondary Cond. | 2128 | 31/10/00 | “ |  |  |
| Transmissometer | 333D | 14/07/00 | IOS |  |  |
| Pressure Sensor | 63507 | 11/01/96 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.