## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2001-12

Agency: OSAP

Location: NE Pacific

Project: High Seas Salmon

Chief Scientist: Steve Romaine

Platform: CCGS W. E. Ricker

Date: 25 March 2001- 1 April 2001

Processed by: Germaine Gatien

Date of Processing: 4 December 2001 – 7 December 2001

Number of original CTD casts: 24

Number of casts processed: 22

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was mounted with Transmissometer 197. The deck unit was a SeaBird 11+ S/N 11P18377-0471.

SUMMARY OF QUALITY AND CONCERNS

The quality of this data is lower than usual probably because of a missing bridle on the CTD leading to flow-rate problems. There were no salinity calibration samples. The primary sensors could not be used because of problems with a pump.

Recalibration was based on previous use of this CTD. During 2001-14 the salinity was considered to be within 0.005units.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The CTD Log was obtained.

No salinity calibration data was available.

The configuration file contains an error in the pressure calibration. The same error was found during cruise 2000-16 (and later Ricker cruises). At that time it was noted that the surface pressures seemed too high and correcting the error reduced the pressures. The calibration file 0506CTD.con was created with the corrected pressure calibration.

The sensor history was found.

Notice was received from Gail Jewsbury that the bridle for the CTD was found to be missing in April 2001. This means that the CTD could hang at an angle during descent. It is presumed that it was missing for all missions from January 2001 to late April 2001. The data collected during this time may be affected.

- Conversion of Raw Data

The raw data was converted using conversion file 0506CTD.con.

An initial examination of the data shows all channels present. There are significant differences between the primary and secondary conductivity particularly in the top 30m of many casts. This occurred during 2001-09 which immediately proceeded 2001-12; at that time possible sources of the discrepancy were considered to be the primary pump and/or the absence of the bridle on the CTD.

4. ALIGNCTD

The deck unit used for this mission advances only the primary conductivity channel by 0.073s.

During 2001-09  the primary was advanced by –0.018s (for a net advancement of +0.055s) and the secondary was advanced by +0.055s.

Tests on cast #15 confirm that the same values are suitable for this cruise. All casts were advanced by     –0.018s for the primary and +0.055s for the secondary conductivity channel.

The fact that bi-polar spikes remain after this step is considered diagnostic of flow-rate problems. This is probably due to the lack of a bridle on the CTD, but pump problems might also be a factor.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

Test were run on cast #15 using a variety of settings for the conductivity cell thermal mass correction. The best results for the primary were (alpha = 0.02; 1/beta = 7.0) and for the secondary either of (0.02/7) or (0.02/9). The conductivity cell thermal mass correction was done for both channels using (alpha = 0.02; 1/beta = 7.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of deep casts (# 1,15 and 34) was plotted to check for agreement between the pairs of T and C sensors. The descent rate was quite smooth and reasonably high. The differences in the sensors were so noisy that it is pointless to estimate an average. It is clear that the secondary temperature is generally higher than the primary and the secondary salinity lower than the primary. The conductivity differences appear to be small (on average).

It was noted in the CTD log that the pumps were not working for cast #17 and that cast #18 was run to replace it. Cast #17 will not be processed further.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. There was no rosette used, so there are no rosette files to convert.

10. Checking Headers

The cruise track was plotted and looks reasonable.

A header summary was produced. The only problem noted was the absence of station names. These were added to the IOS files where the entry in the log book was clear. There is no position in the log book for cast #1, so there is no confirmation for the GPS entry, but it looks reasonable on the cruise track.

The surface check was run. The average surface pressure is 3.5db. This seems high but at the end of cast #31 while the CTD was held near the surface for awhile, it recorded a pressure as low as 0.3db with believable values of T and S. No offset will be applied.

The header check was produced and no errors were found.

- DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance:  1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS: The only warning pertained to data very close to the surface. The maximum pressures were checked and are close to those for the original files.

All DEL files were copied to EDT files.

12. Test Plots

Profile plots were checked for problems, and as noted before there are many differences. The secondary sensors behaved better as was the case during 2001-14. Cast #34 looks particularly bad for both sets of sensors.

13. Intercomparisons

Previous experience with these sensors – Since they were last calibrated these sensors were used during cruises 2001-02, 2001-09 and 2001-14.

During 2001-02 there were Niskin bottle calibration samples, but the scatter was large, the depth of the bottles uncertain and the differences varied with depth. The data was considered insufficient for recalibration but suggested that the CTD salinity was low at the surface and high at depth.

During 2001-09 the few bottles that appeared to be reliable suggested that the CTD salinity is low by from –0.0025 to –0.008units. The data was recalibrated by adding 0.005units to the salinity.

During 2001-14 the CTD salinities appeared to be low by about 0.005units but there were only 4 deep bottles.

##### Bottle comparisons – No bottles available

Historic ranges – All casts were plotted with historic ranges superimposed, where available. All data fell within the historic ranges which are very broad in this region.

14. DETAILED EDITING

The secondary sensors looked better for most casts particularly in the top 30m, so these were selected for editing. Page plots were produced using T1,S1 for guidance in editing the data.

CTDEDIT was used to clean noise in salinity for all casts except #19 and #33. Some records were removed from casts #18,22,24,26,30 and 32. Comments were entered in the headers of each file that was edited and notes were made on the test plots which are found in the section Tests and Plots which follows this report.

Cast #34 was found to contain large unstable areas that are not associated with shed wakes. The descent rate was high and steady for this cast. In areas of large temperature gradient the salinity is unreliable and there is evidence of noise in all channels. It would appear that the problems associated with the lack of bridle became severe during this cast. This cast will be deleted.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

15. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

- Recalibration

There are no bottles available to calibrate the salinity. The salinity was recalibrated using file 2112rcal.ccf which applied an offset of +0.005units to the salinity. This is based on the results of 2001-09 and 2001-14. (See section 13 for details.) For those two cruises the salinity was considered to be within 0.005units. There is no basis to estimate the error for 2001-12.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

21120017 – Pumps not working. Deleted. Cast #18 is at the same station. Deleted.

21120034 – Data looks unreliable, large unstable features, spikes. Deleted

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001-12

Dates: Start: 25 March 2001 End: 1 April 2001

Location: NE Pacific

Vessel: CCGS W. E. Ricker

Chief Scientist: Steve Romaine

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2001-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 30Nov2000 | Factory |  |  |
| Conductivity | 2278 | 6Dec2000 | “ |  |  |
| Secondary Temp. | 2710 | 30Nov2000 | “ |  |  |
| Secondary Cond. | 2280 | 6Dec2000 | “ |  |  |
| Transmissometer | 197 | 14Dec2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.