## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2001-13

Agency: OSAP

Location: NE Pacific/Georgia Strait

Project: LaPerouse, Juan de Fuca, Effingham

Party Chief: Tom Juhaśz

Platform: CCGS John P. Tully

Date: May 19, 2001 – May 27, 2001

Processed by: Germaine Gatien

Date of Processing: June 15, 2001 – July 8, 2001

Number of original CTD casts: 61

Number of casts processed: 61

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0443) was mounted with Transmissometer 333D, two Seapoint Fluorometers #2228 #2229, a PAR sensor. The deck unit was a Seabird 11+ (serial number was not recorded).

SUMMARY OF QUALITY AND CONCERNS

The descent rate was very noisy for many of the casts but the average descent rate was generally kept high. There were many spikes in the data from some casts.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There is a small change in the files between casts #14 and #17. It does not appear to affect the conversion.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using configuration file 21130001.con.

A preliminary check shows all expected channels present except for cast #1 for which one fluorometer was not connected and some casts do not have a PAR sensor channel. In the latter part of the cruise there are a lot of large spikes (including the pressure channel) and the secondary T and C channels are generally spiky.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

It is unknown what deck unit was used for this mission. Tests done on cast #7 established that the deck unit had not advanced the secondary conductivity channel.

Tests on casts #7 and 75 using advancements of +0.050s, +0.055s, +0.060s and +0.073s indicate that the best results are obtained with either +0.055 or +0.060s.

All casts were realigned using –0.013s for the primary and +0.060s for the secondary so that the net alignment is +0.060s for both channels.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The descent rate was often very noisy but the average rate was kept reasonably high.

The differences in temperature, conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 17 | 550 | ~.001 | -.0011 | -.013 | Noisy, slow near the bottom |
| 22 | 1470 | ~.0005 | -.00085 | -.012 | Very noisy |
| 39 | 120 | +0.0035 | -.00085 | -.0014 | High and steady |
| 46 | 100 | +0.0015 | -.0008 | -.0010 | High and steady |
| 54 | 2070 | ~0 | -.0004 | -.005 | Very noisy |
| 55 | 1770 | ~0 | -.00035 | -.0045 | Very noisy |
| 64 | 350 | +0.002 | -.0003 | -.005 | High and steady |
| 65 | 350 | +0.0015 | -.00025 | -.005 | High and steady |
| 80 | 1290 | ~.0005 | -.0002 | -.003 | Very noisy |

There appears to be a drift in the conductivity and salinity differences that is not explained by differences in depth or descent rate. The salinity differences are fairly constant with pressure.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

It is noteworthy that although there were many pressure spikes there were no problems with the NMEA headers.

10. Checking Headers

A header summary and header check were produced and a few errors in station names were found and corrected. The dates given for casts #39-47 are different from those in the CTD log. The log dates are not internally consistent so the header dates are believed to be correct.

The surface check was run. The average surface pressure is 2.2db.

The cruise track was plotted and looks reasonable.

- DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS: None given

Maximum pressure check: The maximum pressure in the DEL files were checked to see that valuable data had not been lost. In the case of cast #63 the bottom 3m were lost because the descent rate was low. Since this cast was in protected waters the requirement of a minumum descent rate of 0.3m/s averaged over 15 points was overly restrictive. So DELETE was rerun for this cast only with the Minimum Drop Rate feature turned off.

All DEL files were copied to EDT files.

12. COMPARE

The BOT files were examined for errors and CTDEDIT was used to fix a few bad points in casts 64 and 65. The edited files were renamed BOT.

Salinity bottle data was available in a spreadsheet. This data was converted to separate files for each cast (*.HY1). COMPARE was then run using these files and the BOT files.

The details of the comparison were checked for indications of errors in the hydro files. None were found.

The primary salinity was found to be high by 0.0033units and the secondary low by 0.0078units. The trendline was quite flat when bottles above 500db were excluded. The primary sensors showed no time dependence but the secondary sensor calibration showed significant drift. This fits the observations of section 8.

13. DETAILED EDITING

All casts were examined for differences between the pairs of sensors and differences between up and down casts for evidence of problems with the processing. The secondary salinity is unusually spiky in many casts.

For cast #40 there is a marked difference between the two conductivity channels and some differences between the temperature channels. The conductivities look much closer in the original converted file, though the temperatures are similar to those observed later. The change comes after CELLTM was run. When it was rerun without applying the CELLTM correction to the secondary channel the problem disappears. The primary channel shows no large differences when CELLTM is skipped. Dropping the conductivity channel from WILDEDIT made no significant difference. There were a lot of spikes in this data and presumably something is affecting how CELLTM works. Since it only seems to affect the secondary channel and the primary seems to have given better data for this cruise, the problem will not be pursued. Such problems do point out the necessity of examining each cast late in the processing stream to look for differences in the channels and between the up and down sections of the cast.

Since the calibration is closer for the primary sensors and the data is less spiky they were chosen for further processing. Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

The following casts were not edited: 38,39,44,71,75.

The following casts were edited at the surface and/or bottom only: 46,47,63,65,67,70.

The following casts were edited more extensively:1-37,40-43,45,48-62,64,66,68,69,72-74,76-82.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. FILTER

The two fluorometer channels were put through a median filter of width 11 to reduce spikiness.

15. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging page plots were examined on screen to check whether further editing was needed. Two casts were re-edited (#59 and 62) to remove more surface data after checking upcasts to confirm that the downcast vaues were unreasonable. A few other unstable features were re-examined but not edited since they may be real and there is no indication that they are due to instrumental or processing errors.

16. Other comparisons

Previous experience with these sensors – No data has been processed using the primary conductivy sensor since its last calibration. The secondary sensors were used for 2001-06 when they were found to produce salinity high by 0.002units; there were few bottles and a lot of scatter for that comparison.

Historic ranges – Four deep casts were plotted with the historic ranges superimposed. All data fell within those ranges. (See plots in Processing Section of Report.)

- Recalibration

File 2113rcal.ccf was used to recalibrate the data (including the rosette files) as follows:

- primary salinity – decrease by 0.0033units

- secondary salinity – increase by 0.0078units

- Chl_fluorescence:Seapoint – decrease by 0.12g/l

- Fluorescence:0 – multiply by 3, then decrease by 0.06g/l

The fluorescence offsets were based on a study of  “dark values”. The fluorometer:0 channel was recorded in volts, so was multiplied by 3 to convert it to g/l.

The SAM files were recalibrated and named SAC; these files were used to rerun COMPARE. The results (2113cpm2.xls) indicate that the deep values after recalibration were within 0.00003units of the bottles. (See plots in Processing Section of Report.)

18. REMOVE and REORDER

The following channels were removed from all casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary, Fluorometer:0 and Flag.

The PAR (Irradiance) channel was removed from casts 19-36, 54-61 and 70-82.

The channels were reordered and data format corrected as needed.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

20. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) A separate set of files was prepared for Angelica Pena. These were bin averaged with a 0.25m bin size. They were then recalibrated and reordered as described in sections 17 and 18; however, the Fluorometer:0 channel was not removed.

### Particulars

21130001 – Unpumped fluorometer not connect.

21130002 – Transmissometer not working until about 125db into downcast.

21130019 – PAR removed

21130040 – Two conductivity channels very different in top 25db.

21130064 – CTDEDIT was used to edit one record (T1 and S1) in the rosette file.

21130065 – CTD hit bottom, mud in system. CTDEDIT was used to edit a few points in the BOT file (T1 and S1) around 150 and 300db. There is a lot of drift in both salinity channels at the bottom bottle probably due to hitting bottom. No editing was attempted on that section. These points were too shallow to be included in COMPARE.

21130070 – CTD removed from rosette

21130078 – Spike in data around 115m.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 13

Dates: Start: 19 May 2001 End: 27 May 2001

Location: NE Pacific / Georgia Strait

Vessel: CCGS John P. Tully

Party Chief: Tom Juhaśz

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2001-13

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2106 | 02/11/00 | Factory |  |  |
| Conductivity | 1763 | 31/10/00 | “ |  |  |
| Secondary Temp. | 2023 | 01/11/00 | “ |  |  |
| Secondary Cond. | 2128 | 31/10/00 | “ |  |  |
| Fluorometer -unpumped | 2228 | ? |  |  |  |
| Fluorometer -pumped | 2229 | ? |  |  |  |
| Transmissometer | 333DR | 14/07/00 | IOS |  |  |
| PAR | 4615 | 08/02/01 |  |  |  |
| Pressure Sensor | 63507 | 31/11/96 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.