## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 11-Jan-2006 | Surface loop data was added to the archive. The data was acquired from John Morris at PBS. The bottle salinity data was added to John’s spreadsheet from file “2001-18sal.xls” that was already in the DOC directory. The original spreadsheet file from John and more detailed processing notes can be found in the “Cruise_Data\Documents” directory. Any questions regarding this data should be directed to John Morris. J.L. |

PROCESSING NOTES

Cruise: 2001-18

Agency: Stock Assessment

Location: WCVI/Gulf of Alaska

Project: Hake

Chief Scientist: Mark Saunders

Platform: CCGS W. E. Ricker

Date: 25 July 2001- 20 August 2001

Processed by: Germaine Gatien

Date of Processing: 2 January 2002 – 17 January 2002

Number of original CTD casts: 95

Number of casts processed: 94

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0506) was mounted with Transmissometer 197. The deck unit was a SeaBird 11+ S/N 11P18377-0471.

SUMMARY OF QUALITY AND CONCERNS

There is some uncertainty in times and positions.

Many of the casts contained spikes in signals that led to errors in the headers and corrupted files. Most data were recovered using FIXDAT.

There were very few deep hydro bottles available for comparison, and there was an unexplained depth-dependence in the comparison; hence, the salinity should be considered accurate to only 2 decimal places.

The average descent rate was kept high most of the time, minimizing shed wakes.

Transmissivity is unedited.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The CTD Log was obtained. There was no list of equipment in the log. Note was made of problems encountered with CTD casts.

The configuration file contains an error in the pressure calibration. File 0506CTD.con was created with the corrected pressure calibration.

The sensor history was found.

The Niskin bottle sample data was obtained from Bernard Minkley.

- Conversion of Raw Data

The raw data was converted using conversion file 0506CTD.con. Some rosette files were converted, but all but one contain bad data only. Cast #67 looks reasonable but there is no evidence in the log book that a rosette was used. It appears that these files were created by stray signals associated with spikes.

An initial examination of the data shows that cast #1 contains bad data only (spikes). The SeaBird lat/long/time headers are rotated in the DAT file; these were fixed and the data converted again with no improvement in the spikes. The routine FIXDAT was run on the data without success; an error message said that the program could not determine the scan length and so could not repair the file.

A few other casts were examined and looked reasonable with all expected channels present.

There is a file 20011800xx.DAT which contains only surface data. This is presumed to be a test cast and will not be processed.

4. ALIGNCTD

The salinity channels were removed using the SeaSoft routine STRIP.

The deck unit used for this mission advances only the primary conductivity channel by 0.073s.

For 2001-09 and 2001-12 the primary conductivity was advanced by –0.018s (for a net advancement of +0.055s) and the secondary was advanced by +0.055s. Tests on cast #3 and 18 confirm that those values are best for this cruise and they were applied to all casts.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in Pressure.  Parameters used were:

Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

In the previous use of this CTD the optimal choice of parameters for the thermal mass correction was found to be (alpha = 0.02; 1/beta = 7.0.) Tests were run on cast #3 and 16 using a variety of settings for these parameters. For the secondary sensors the best choice for both casts was (0.02/7). The best results for the primary were different for the two casts (alpha = 0.02; 1/beta = 9.0) and (0.03,7.0) respectively; Upon investigation it became apparent that the depth of the feature studied for cast #16 was in an area corrupted by a large shed wake. Cast #90 was then studied and the choice of (0.02/9) proved best.

The conductivity cell thermal mass correction was done for the primary channels using (alpha = 0.02; 1/beta = 9.0.) and for the secondary channel using (alpha = 0.02; 1/beta = 7.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of deeper casts (# 12, 41, 57 and 76) was plotted to check for agreement between the pairs of T and C sensors. The descent rate was noisy but the average was fairly high. There is a lot of fine-scale noise in the pressure. The primary temperature was lower than the secondary by about 0.001Cº, the primary conductivity and salinity are higher than the secondary channels by about 0.0005units and 0.006units, respectively.

All casts were displayed on-screen. The following casts did not appear to contain any useful data, only spikes: #15,43,54,56,66,78,93,96. The routine FIXDAT was run on the DAT files for these casts and steps 3 through 7 were rerun using the repaired files.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. There was no rosette used, so there are no rosette files to convert.

10. Checking Headers

The cruise track was plotted and looks reasonable.

A header check was run and indicated problems in headers for 8 files. These were rotated entries and absence of hemisphere in positions probably caused by spikes. These were fixed and no further errors were found.

The header summary was produced and checked against the CTD log entries. There are many errors in times and positions; these were probably caused by signal spikes. Unfortunately, the CTD log entries are far from clear so correcting the errors was not straightforward. Variations in the formats used for time (usually in 24-hour format, but occasionally “p.m.” times are written in 12-hour format but without the “p.m.” noted) and positions (usually in degrees and minutes but sometimes decimal degrees.) The headers were corrected using the best guess. It took some time to interpret the entries and it is possible that there are some errors in the final data files. It was impossible to check the ship’s log at the time of processing. Gail Jewsbury will check the positions at a later date and corrections should be made if errors are found.

There are also errors in the names of the first 8 cast files, which were numbered in order of CTD casts; the names were changed to match the consecutive numbers in the CTD log.

The surface check was run. The average surface pressure is 3.4db. This seems high but is the same value found during 2001-12 when investigation suggested the values were accurate. A data point from the upcast section of cast #97 has believable values of T and S corresponding to a pressure of -0.1db. No offset will be applied.

- DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance:  1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS: Most warnings pertained to data close to the surface or bottom. A few mid-depth warnings were investigated and found to occur in regions where the CTD reversed direction during the downcast. The maximum pressures were checked and are close to those for the original files.

All DEL files were copied to EDT files.

12. Test Plots

Profile plots were checked for problems and primary sensors were chosen for further processing.

13. Intercomparisons

Previous experience with these sensors – Since they were last calibrated these sensors were used during cruises 2001-02, 2001-09, 2001-12 and 2001-14.

During 2001-02 there were Niskin bottle calibration samples, but the scatter was large, the depth of the bottles uncertain and the differences varied with depth. The data was considered insufficient for recalibration but suggested that the CTD salinity was low at the surface and high at depth.

During 2001-09 the few bottles that appeared to be reliable suggested that the CTD salinity is low by from –0.0025 to –0.008units. The data was recalibrated by adding 0.005units to the salinity.

There were no salinity samples from 2001-12.

During 2001-14 the primary CTD salinities appeared to be low by about 0.005units but there were only 4 deep bottles.

Previous uses of this CTD in 2001 have all had serious problems. The bridle was missing from the CTD for some and the acquisition system was set to average over 33 scans for one. Neither of these problems occurred during 2001-18.

##### Bottle comparisons  - There were only 6 deep Niskin bottles. The depths are uncertain. At 500m the four bottles all indicate that the CTD salinity is low (by about 0.012units and 0.017units for the primary and secondary sensors, respectively) and at 1000m the primary appears to be high and the secondary low by about 0.003units and 0.002units respectively. It is unusual to have this sort of pressure dependence at depth, especially in both sets of sensors, but there are only 2 bottles at 1000m.

Historic ranges – Four plots were produced with historic ranges superimposed and all temperature and salinity data fell within those ranges.

14. DETAILED EDITING

The primary salinity is smoother than the secondary so has been selected for archiving.

CTDEDIT was used to edit all casts except #10, #78 and #90.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

15. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

- Recalibration

The depth dependence of the bottle comparisons is troubling. This might indicate trouble with the CTD, but could also be a problem with the bottles. There is too little data to make a pressure-dependent recalibration. In previous use of this instrument the secondary salinity was recalibrated by adding 0.005units. While the results from 500m would suggest a larger recalibration, the deep results indicate a smaller one. It seems reasonable to apply the same correction as for 2001-14. The primary salinity is higher than the secondary by about 0.005units, so for consistency the secondary should be raised by 0.01units. The salinity should only be considered good to 2 decimal places.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

- REMOVE

The following channels were removed from all casts: Scan_Number, Primary Temperature, Primary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The data format was corrected as needed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

20011800xx – Data file contains only surface data, will not be processed. Presumed to be a test cast.

21180001 – No good data, full of spikes.

21180003 to 21180011 – File names and event numbers changed to match consecutive  number in CTD log.

21180009 – Time in CTD log listed as 1913?; time in cast file 18:15. Used data file time.

21180011 – Pumps not turned on until 50m, returned to surface then did full cast.

21180020 & 22 – Changed station names from NSTN to C-3 and C-2, resp.

21180026 & 27 – Changed station names from C-1 to NSTN. Should be confirmed later against ships log.

21180043 – Fixed time/positions

21180054 – Fixed time/positions

21180063 & 64 – time/positions differ in log and data file. Should be confirmed later against ships log.

21180066 – Fixed time/positions.  Should be confirmed later against ships log.

21180078 – Fixed time/positions

21180081 – Fixed time/positions.  Should be confirmed later against ships log.

21180093 – Fixed time/positions

21180095 – Fixed time/positions.  Should be confirmed later against ships log.

21180096 – Fixed time/positions

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001-18

Dates: Start: 25 July 2001 End: 20 August 2001

Location: NE Pacific

Vessel: CCGS W. E. Ricker

Chief Scientist: Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2001-18

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 30Nov2000 | Factory |  |  |
| Conductivity | 2278 | 6Dec2000 | “ |  |  |
| Secondary Temp. | 2710 | 30Nov2000 | “ |  |  |
| Secondary Cond. | 2280 | 6Dec2000 | “ |  |  |
| Transmissometer | 197 | 14Dec2000 | IOS |  |  |
| Pressure Sensor | 69698 | 12Nov1997 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.