## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 12 Feb 2025 | Fixed channel name and format for oxygen flag channel – cast #55. G.G. |
| 21Feb2019 | Added flag 3 to 2 DO samples in 2001-19-0055.CHE file; had to add a flag channel first. |

PROCESSING NOTES

Cruise: 2001-19

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca

Project: SoG/JdeF

Party Chief: Diane Masson

Platform: CCGS VECTOR

Date: June 25, 2001 – June 30, 2001

Processed by: Germaine Gatien

Date of Processing: 11 September 2001 – 17 September 2001

Number of original CTD casts: 70

Number of casts processed: 70

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 192D and Seapoint Fluorometer #2229. The deck unit is unknown.

SUMMARY OF QUALITY AND CONCERNS

The data was in good order.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. An error in the pressure calibration was corrected and saved as 2119CTD.con.

The sensor history was found.

The files had long names (ex. 2001190001.*) which were changed to standard 8-digit file names.

- Conversion of Raw Data

The raw data was converted using configuration file 2119CTD.con.

A preliminary check shows all expected channels present.

Dark values were checked for the fluorometer and are generally about 0.15g/l at 350db. These values are consistent with the use of the 10X cable which is what is recorded in the CTD log and the configuration file. The observed maxima ( 2-7g/l) are reasonable when 10X is assumed.

The temperature and salinity traces track reasonably well and up and downcasts are similar. The secondary salinity looks spiky and the salinity differences are on the order of 0.001units at 2000db.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

There is no record of which version of deck unit was used, but it appears that it is one of the newer versions which advance both the primary and secondary conductivity channels. Tests of alignment were made on casts #1 and 66. Spikes caused by mismatch of T and C were a little difficult to pick out since there was so much variability in temperature. Cast #1 is fairly quiet; the deck unit alignment or –0.013s seems best there for the primary conductivity while –0.013s or –0.018s improves the secondary alignment. For cast #66 –0.013s or –0.018s improve the primary and –0.023s looks best for the secondary. In two previous cruises using this equipment the deck unit setting was used for the primary and –0.023s for the secondary in 2001-01 and deck unit settings for both in 2001-06.

All casts were realigned using the following parameters:

-0.013s for the primary conductivity (for a net advancement of +0.060s)

-0.018s for the secondary conductivity (for a net advancement of +0.055s)

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

Tests were run on casts #1,2 and 66  to determine the optimal parameter choice for this routine. The usual choice proved best although alpha=0.02 looked slightly better for the secondary channel of cast #66.

The conductivity cell thermal mass correction was done for both channels (alpha = 0.03; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The differences in temperature and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 32 | 325 | +0.0005 | -0.0005 | -0.006 | Good |
| 48 | 450 | +0.001 | -0.0005 | -0.006 | Good |
| 66 | 340 | +0.0008 | -0.0005 | -0.006 | Good |

All casts were plotted on-screen to check for problems with processing. Reasonable agreement was found between up and down casts and between pairs of sensors.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary and header check were produced. There were a few errors in headers which were  corrected in the CNV files in directory DERIVE. Those files were then reconverted.

The times in the CTD log are written in 12-hour format rather than 24-hour format. Fortunately, the GPS times were all available and since the casts were so frequent there was no room for misinterpretation.

The surface check was run. The average surface pressure is 2.0db.

The cruise track was plotted and looks reasonable.

- DELETE

- CLEAN was run to replace pad values in pressure with interpolated values.

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 10.00

- Pressure Tolerance: 1.0

- Pressure filtered over 15 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.

- Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS: The only warning pertained to surface data. A comparison of maximum depths before and after DELETE led to investigation of one cast (#27); the data lost was near the bottom, the descent rate was very low and the salinity data was very noisy. The removal of this data is considered valid.

- All DEL files were copied to EDT files.

- 12. COMPARE

- The BOT files were examined for errors and no bad data found.

- The BOT files were bin-averaged with Bottle Number as the bin channel.

- Compare was run. There is a lot of noise in the comparison as is expected in an area of such great variability. There is obvious depth-dependence. The averages of the differences in the fit were –0.0044 and –0.010units for the primary and secondary salinity respectively, so both the CTD salinity channels are reading low. Around 350m the primary differences are about –0.002units and the only point deeper than that is very close to zero. The secondary differences around 350m are about –0.008 and for the deepest bottle –0.006units.

- There was one anomalously high bottle salinity value in file 21190030.sal (depth 200m).

- DETAILED EDITING

The primary sensors were chosen for further processing.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

The following casts were edited at the surface or bottom only:11,12,14,16,18,20,29,36,37,40,42,43,46,53, 57.

The following casts were edited more extensively:1-10,13,17,19,21-23,26,27,30-35,39,41,44,45,47,50-52,54-56,61-65,67-71.

The descent rate was noisy for some of the casts in Juan de Fuca Strait but was quite steady elsewhere. Records that were obviously corrupted by shed wakes were removed but in this region it is often unclear which unstable features are due to shed wakes and which are natural features. Where there was doubt records were left unedited.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Pena.

15. BIN AVERAGE

The following Bin Average values were applied to the BOX files for general use:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen to check whether further editing was needed.

15. Other comparisons

Previous experience with these sensors –These sensors were used during 2000-01 when 10 bottles (none deep) indicated that the primary salinity was high by 0.0004 and the secondary low by 0.004units. There were many deep bottles for 2001-06 when the primary salinity was high by 0.0011 and the secondary low by 0.0047units. The bottles in the top 500m looked quite different from the deep ones with negative values for the primary and values on the order of -0.008 for the secondary. The primary sensors were chosen for archiving in both cases.

Historic ranges – A few casts were plotted with the historic ranges superimposed. All data was within the ranges. (See plots in Processing Section of Report.)

- Recalibration

It is unclear how to recalibrate this data. The results from 2001-06 do not agree well with the results of COMPARE for this cruise. The lack of deep bottles, the large scatter and the pressure-dependence in the results causes further doubts. Since there were many deep bottles during 2001-06 the corrections from that cruise will be applied to the salinity channels.

File 2119rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: -0.0011units

- secondary salinity –  offset: +0.0047units

- Chl_fluorescence -  offset: -0.15g/l

- REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and data format corrected as needed.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed from all casts.

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

22. Station name entered wrong. Changed to station 65.

27. Pump turned on at 10m depth.

40. CTD touched muddy bottom. Upcast salinity wrong.

41. CTD stopped a few times because of cable angle. Station name wrong. Changed to station 42.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 19

Dates: Start: 25 June 2001 End: 30 June 2001

Location: Strait of Georgia / Juan de Fuca Strait

Vessel: CCGS VECTOR

Party Chief: Diane Masson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2001-19

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2374 | 02/05/00 | Factory |  |  |
| Conductivity | 1764 | 28/09/00 | “ |  |  |
| Secondary Temp. | 2668 | 07/03/00 | “ |  |  |
| Secondary Cond. | 2102 | 28/09/00 | “ |  |  |
| Fluorometer -pumped | 2229 | 18/06/01 | IOS |  |  |
| Transmissometer | 192D | 03/11/00 | IOS |  |  |
| Pressure Sensor | 75636 | 4/06/99 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.