## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| March 19, 2002 G. Gatien | SBE oxygen data was calibrated in CTD and CHE files using 2120oxy.ccf. Formats were changed and channels added as needed. See oxygen_calib.doc A note was put in headers that oxygen is to be considered ±5%. |

PROCESSING NOTES

Cruise: 2001-20

Agency: OSAP

Location: NE Pacific

Project: Endeavour Ridge / Effingham Inlet

Party Chief: Rick Thomson

Platform: CCGS John P. Tully

Date: July 3, 2001 – July 13, 2001

Processed by: Germaine Gatien

Date of Processing: August 23, 2001 – 11 September, 2001

Number of original CTD casts: 119 (including 1 downcast only,1 upcast only)

Number of casts processed: 118 downcasts (+8 upcasts processed for R. Thomson)

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 192D, Seapoint Fluorometer #2356 and a PAR sensor (4615). The deck unit was a Seabird 11+ (serial number 0424).

SUMMARY OF QUALITY AND CONCERNS

The data was generally in good order.

There is a pressure-dependent variation in the differences between the conductivity sensors; the few salinity samples available can not resolve which sensor is more accurate. Salinity should be considered accurate to within ±0.004 above 300m and ±0.002units below that.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using configuration file 21200002.con.

A preliminary check shows all expected channels present. As noted in the PAR sensor was used only for casts #2,5,39-71,93-121.

Dark values were checked for the fluorometer and are generally about 0.03mg/l at 1000db. These values are consistent with the use of either the 3X cable or the 10X cable; the latter is believed to be the one used for this cruise. The observed maxima (from 2 to 15mg/l) are reasonable when 10X is assumed.

The temperature and salinity traces track reasonably well and up and downcasts are similar. The secondary salinity looks spiky and the salinity differences are on the order of 0.001units at 2000db.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

5. ALIGNCTD

The deck unit did not advance the secondary conductivity channel.

Tests on casts #18 and 82 show that an alignment of conductivity channels by +0.060s minimizes salinity spikes.

A study was made of the alignment of the SBEOXY channel. This is the second data set processed using this instrument. The manual suggests advancing the oxygen channel for a pumped instrument by +2s to +5s. For cruise 2000-08 a value of 8s proved optimal, but it is believed that the plumbing was a little different for the two missions. Casts #18 and 82 were studied in detail by plotting the up and downcasts together, displaying just oxygen and primary temperature. Getting the two traces to overlie each other is the method suggested in the SeaBird manual. In shallow water this might be reasonable although local effects might cause real changes; in deep water the effect of internal waves would cause the traces to not overlie each other. Deep casts were chosen and the difference between up and downcast temperature traces was taken to be the standard so that the internal wave problem is allowed for. For cast #18 the optimal choice appears to be +10s although the drift in value of oxygen complicates this judgement. Cast #82 is consistent with this choice.

All casts were realigned using the following parameters:

–0.013s for the primary conductivity (for a net advancement of +0.060s)

+0.060s for the secondary conductivity

+10s for the SBE oxygen

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   Pass 1    Std Dev = 2		Pass 2    Std Dev = 5		Points per block = 50

6. CELLTM

Tests were run on casts #18 and 82 to determine the optimal parameter choice for this routine.

The conductivity cell thermal mass correction was done for both channels (alpha = 0.02; 1/beta = 9.0.)

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The differences in temperature, conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 7 | 2450 | ~0 | ~-0.0001 | ~-0.001 | Noisy, high except >2000db |
| 35 | 1600 | ~0 | ~-0.0001 | ~-0.001 | Noisy, high except >1200db |
| 81 | 1900 | ~0 | ~-0.0001 | ~-0.001 | Noisy, high |

All casts were plotted on-screen to check for problems with processing. Reasonable agreement was found between up and down casts and between pairs of sensors. Sub-surface maxima in oxygen are noted in the downcasts, but not upcasts.

It was noted that the differences in salinity (S1-S0) at the beginning of casts, in well-mixed surface waters, was on the order of +0.0025units, rather than the –0.001 found at depth. By contrast the results for 2001-08 are similar in well-mixed surface waters and at depth. These differences were not a result of processing steps such as alignment, since the same differences are also seen in the raw data. They are clearly seen in the conductivity differences, but not in temperature. The upcasts do not show the same character; at 500m there is no hint that the differences are changing sign. Some of the surface data has  S0 higher than S1 but this is in a region of very noisy descent rate and large gradients so is not unexpected. There are no bottles above 800m to help resolve this issue.

The variations in salinity differences can only be seen clearly in deep casts, which mostly occurred early in the mission. Two deep casts (#13 & 83) look similar, and there are hints of the same variations in shallower casts.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary and header check were produced and a few errors in station names were found and corrected.

The surface check was run. The average surface pressure is 1.2db which seems a little low. However, on checking a few casts with surface pressure close to or less than 0, it was found some of the conductivity values were very low or –9.99, so this data is indeed from very close to the surface.

The cruise track was plotted and looks reasonable.

REVERSE was applied to casts #6-7 and 9-14 since upcasts will be needed.

The upcast files will be identified by a 9 in the 5th digit.

- DELETE

- CLEAN was run to replace pad values in pressure with interpolated values.

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 20.00

- Pressure Tolerance: 1.0

- Pressure filtered over 15 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.

- Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS: The only warning was for cast #9 which was an upcast only.

- Maximum pressure check: Casts #48,49,50,60 and 62 were missing data from the bottom. Since these casts were in inlets the DELETE LOW DROP RATE feature was turned off and they were put through DELETE again.

- The “upcast” files 9006-9007 and 9009-9014 were put through DELETE.

- All DEL files were copied to EDT files.

- 12. COMPARE

- The BOT files were examined for errors and no bad data found.

- The BOT files were bin-averaged with Bottle Number as the bin channel.

- There were salinity bottle values for only 2 casts, 5 deep bottles in all. A spreadsheet was prepared comparing the salinity bottle values with those in the averaged bottle files. The two sensor pairs give very similar results with the secondary slightly closer to the bottles. The primary was high by an average of 0.0052units and the secondary by 0.0046units. There appears to be some depth-dependence with the lowest values at depth. The bottles at about 1000db for casts 34 and 83 are reasonably close so there is no hint of time-dependence in the errors.

- DETAILED EDITING

The primary sensors were chosen for further processing.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. All casts required some editing except for casts #61 and 63. The descent rate was very noisy for many of the casts; records that were obviously corrupted by shed wakes were removed.  Many questionable records were not removed from the Juan de Fuca casts, since it was not clear which unstable features were due to shed wakes and which were natural features of that region.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Pena.

15. BIN AVERAGE

The following Bin Average values were applied to the BOX files for general use:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen to check whether further editing was needed.

15. Other comparisons

Previous experience with these sensors – No data has been processed using these sensors since the latest recalibration.

Historic ranges – 4 casts were plotted with the historic ranges superimposed. All data was within the ranges. (See plots in Processing Section of Report.)

- Recalibration

File 2120rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: -0.0052units

- secondary salinity –  offset: -0.0046units

- Chl_fluorescence -  offset: -0.03mg/l

- REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The PAR channel was removed for casts #6-38, 72-92 and 122.

The channels were reordered and data format corrected as needed.

The channels for the rosette files were also reordered but only Conductivity:Primary, Conductivity:Secondary and Flag channels were removed from all casts.

The upcast files 21209006-21209014 were metre-averaged, recalibrated and reordered. No editing was done to these files and the fluorometer channel was not filtered.

18. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

8. Downcast only

9. Upcast only

17.Transmissivity dip at 800db

36. Probably hit jellyfish – transmissivity messy

46. Hit bottom.

51. Error in station name - fixed

66. Error in station name - fixed

77. Probably hit jellyfish – tranmissivity messy

106. Error in station name –fixed

101-122 No rosette – CTD vertical (horizontal in earlier casts)

111 – Pump started after acquisition started.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 20

Dates: Start: 3 July 2001 End: 13 July 2001

Location: NEPacific/Effingham

Vessel: CCGS John P. Tully

Party Chief: Rick Thomson

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2001-20

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2449 | 01/11/00 | Factory |  |  |
| Conductivity | 1766 | 31/10/00 | “ |  |  |
| Secondary Temp. | 2038 | 30/11/00 | “ |  |  |
| Secondary Cond. | 2173 | 06/12/00 | “ |  |  |
| Fluorometer -pumped | 2356 | 01/07/01 |  |  |  |
| SBE43 Oxygen sensor | 0047 | 16/05/01 | ? |  |  |
| Transmissometer | 192D | 03/11/00 | IOS |  |  |
| PAR | 4615 | 15/02/01 |  |  |  |
| Pressure Sensor | 75636 | 4/06/99 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.