# Linda White 							Feb. 22, 2002

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Mary O’Brien

MEHS

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Arctic Frozen nutrients 2001-21 and 2001-22

Sediment Trap Supernatants and Rinses AGS00							Nitrate+Nitrite, Silicate and Phosphate

Summary

Frozen seawater samples were thawed and analysed on a Technicon Autoanalyser according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

Once all the reagents and 3.2% NaCl baseline solution are connected the system is conditioned with a Low, Medium and High standard.

A medium check standard was analysed as a sample separating each station.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning of each day and every 30 – 50 samples.  A final standard envelope completes the analysis at the end of the day.

The baseline/wash solution is 3.2% NaCl for marine samples. Double milli q water is analysed as a sample to establish the true baseline.

The samples were run a station at a time and the duplicate set followed.

All labware and nutrient tubes have been acid cleaned, soon to be capped and stowed.

## Phosphate Turbidity

Surface samples of <27o/oo salinity must be corrected for turbidity.

All reagents and 3.2% NaCl baseline solution are disconnected and the sample is analysed with a DMQ baseline.  Absorbance readings are deducted from the phosphate sample value where salinity is <27o/oo.

#### Instrument Repairs

Pump platen tubing has been replaced, rinsed with 3N NaOH and DMQ water.

The carousel has been repaired – Al Thorn replaced the locking pin and inserted a cone guide in the centre hole and widened the pinholes to ease the removal of the carousel.  It’s perfect now.

The wash manifold requires some repair, which will be tended in March.

## Observations

All samples were analysed in duplicate.

Samples were in excellent condition and labelled well.

Analysis went very well.

All data points have been deposited to spreadsheets provided – 2001-22 and 2001-21.xls.

## Quality Assurance data

## Pooled Standard Deviation  Sp = Square root (sum of diff sqrd/2n)

## Samples Sp

Sample Duplicate analysis:	Nitrate		Silicate 		Phosphate

0.33		0.38			0.01

n = 430	    	n = 423			n = 428

### Standards

## Calibration against Sagami

20 µM Nitrate – Linda 0.5% > Sagami

Note: there was no correction to nitrate + nitrite.

50 µM Silicate – Linda 1.81% > Sagami

Note: a correction was made to the standard values before calculating the peaks.  An error was made in correcting too high and re-corrected to the Sagami standard response for silicate.

Silicate values were decreased by 2(1.8%) on the spreadsheet.  Please see working Excel spreadsheet filename Arctic200122nut.xls.

The standard solutions were prepared by Janet Barwell-Clarke in August 2001.

Primary Standards for Nitrate 2500µM and Silicate 7501 µM were diluted for the daily working standards.

A primary solution of 2500 µM Phosphate was diluted to 250 µM Phosphate - secondary standard.

The secondary standard was again diluted to produce the daily working standards.

## Relative Standard Deviation %

Standards		N03 16.22 µM/L		Si 40.72 µM/L		P04 1.60µM/L

# Calibrated

Medium check		16.2+/-0.13		40.9+/-0.21		1.60+/-0.01

Analysed		n = 50			n = 50			n = 46

Relstdev %		0.82%			0.51%			0.6%

This report includes data results on an Excel spreadsheet with the following sheets: data, quality assurance results, peakbase calculated .anz files, turbidity measurements, and plots.

Filename: Arctic200122nut.xls, 2001-22_nutrients.xls2122nutrients