## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 20-Mar-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |
| July 29, 2005 | Added DIC, Alkalinity and pH data from the Haida project to rosette files. Added PC, PN and TOC data, from GLO-FLO bottles, later in August after Lisa Miller had supplied me with dates and locations. The GLO-FLO files have cast numbers > 9000. J.L |
| March 19, 2002 | SBE oxygen data was calibrated in CTD and CHE files using 2131oxy.ccf. Formats were changed and channels added as needed. See oxygen_calib.doc A note was put in headers that oxygen is to be considered ±5%. |

PROCESSING NOTES

Cruise: 2001-31

Agency: OSAP

Location: NE Pacific

Project: Haida Eddies & Nestucca

Party Chief: Dave Mackas

Platform: CCGS John P. Tully

Date: September 17, 2001 – October 2, 2001

Processed by: Germaine Gatien

Date of Processing: October 30, 2001 – November 6, 2001

Number of original CTD casts: 24

Number of casts processed: 24

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with SeaTech Transmissometer 192D, Seapoint Fluorometer #2229 with 10X cable, Oxygen SBE43 sensor #0047 and PAR sensor #4615. The deck unit was a Seabird 11+ (serial number 0424).

SUMMARY OF QUALITY AND CONCERNS

The data was in good order.

PAR is unedited except that records obviously corrupted by shed wakes have been removed.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained. Salinity calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using configuration file 21310001.con.

A preliminary check shows all expected channels present.

The PAR sensor was used only for casts: #1-29,52,91,114.

Dark values were checked for the fluorometer and are generally about +0.12g/l around 2000db. These values are consistent with the use of the 10X cable.

The oxygen values at the surface are very close for up and downcasts.

The temperature and salinity traces track reasonably well and up and downcasts are similar; however, the differences in salinity appear to be about 0.01units which is higher than usual. Similar results were noted during 2001-29 using the same equipment.

4. STRIP

The salinity channels were stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.  The “number of bottles fired” channel was also removed.

5. ALIGNCTD

The deck unit did not advance the secondary conductivity channel.

Tests on cast #17 show that an alignment of the primary conductivity channels by -0.033s (net alignment +0.040s) and secondary conductivity channel by +0.055s minimizes salinity spikes.  These were the values used for cruise 2001-29. All casts were realigned using those values.

A study was made of the alignment of the SBEOXY channel. . The manual suggests advancing the oxygen channel for a pumped instrument by +2s to +5s. This equipment has been used on three previous cruises 2001-08, 2001-20 and 2001-29 when advancement by +8s, +10s and +5s respectively were found optimal. It is known that there were differences in plumbing which may account for the variations in advancement. On some cruises the fluorometer and oxygen sensor were mounted on the same pump, for others one was on the primary pump and the other on the secondary. For this cruise it is believed that the oxygen sensor was on the primary pump and the fluorometer on the secondary pump.

Cast #17 was studied in detail by plotting the up and downcasts together displaying just oxygen and primary temperature. Getting the two traces to overlie each other is the method suggested in the SeaBird manual. In shallow water this might be reasonable although local effects might cause real changes and in deep water the effect of internal waves would cause the traces to not overlie each other. So deep casts were chosen and the offset of downcast and upcast temperature was compared and taken to be the standard so that the internal wave problem is allowed for. Casts with little rosette sampling were selected to simplify the comparison. An advancement of +6s led to optimal results and was applied to all casts.

6. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

7. CELLTM

Tests were run on casts #17 and 91 to determine the optimal parameter choice. Combinations tested were (α,1/β) = (2,7),(2,9),(3,7),(3,9). The choice of (2,7) produced the best results although (2,9) looked just as good for cast #91. The conductivity cell thermal mass correction was done for both channels (α = 0.02; 1/beta = 7.0.)

8. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

9. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The descent rate was often very noisy but the average rate was kept reasonably high for most casts.

The differences between pairs of sensors were:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 1000 | ~0 | +0.001 | +0.0125 | Extremely noisy, high |
| 44 | 2000 | +0.0006 | +0.0009 | +0.011 | Extremely noisy, high |
| 85 | 2000 | +0.0006 | +0.0008 | +0.008 | Extremely noisy |
| 133 | 2000 | +0.0006 | +0.0009 | +0.011 | Extremely noisy, high |

The temporal variations noted during 2001-29 are not observed in 2001-31.

Temperature differences are extremely noisy and show an increase with depth.

10.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

11. Checking Headers

A header summary and header check were produced and no errors were found.

The surface check was run. The average surface pressure is 2.3db.

The cruise track was plotted and looks reasonable.

- DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00   Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.  Sample interval =  .04 seconds.

COMMENTS ON WARNINGS: There were no warnings.

Maximum pressure check: No problems noted.

All DEL files were copied to EDT files.

13. COMPARE

The BOT files were examined for errors. The salinity channels were noisy in casts #50 (bottle #6) and #85 (bottle #23). In each case these are surface bottles and the conductivity and temperature are noisy as well and the CTD appears to have been bobbing up and down repeatedly; the data was left unedited as it probably represents real changes in the water sampled.

COMPARE was run using SAL and BOT files. The details of the comparison were checked for indications of errors in the hydro files and none were found.

The primary salinity was found to be high by +0.004units and the secondary high by +0.014units.  The variations with file pair number do not show any significant time-dependence.

- DETAILED EDITING

All casts were examined for differences between the pairs of sensors and differences between up and down casts for evidence of problems with the processing. None were found.

The primary sensors were chosen for further processing since they are closest to the bottles.

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

All casts were edited using CTDEDIT; this editing was fairly light for casts #1,15,17,22,23,26, 29,50,52,54,58,109,114 & 115. All other casts were edited very heavily as they were corrupted by many shed wakes. Note was made of the editing details in the relevant files. Edited files were copied to *.EDT.

A study was made of two casts (#44 and #70) that contained many shed wakes. Comparing data from 400-500m before and after DELETE and CTDEDIT shows that the two routines were effective in removing most of the obviously corrupted features and that both routines were necessary even in deep water.

15. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness.  Cast #44 was examined before and after this step to ensure that it was effective. BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. These files were then stored in a separate directory for Angelica Peña.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen to check whether further editing was needed.

17. Other comparisons

Previous experience with these sensors –

All sensors were used during 2001-29 when the primary salinity was found to be high by 0.006units, the secondary high by 0.011units and the fluorescence high by 0.11g/l.

Historic ranges – 4 Deep casts were plotted with the historic ranges superimposed. All data fell within those ranges. (See plots in Processing Section of Report.)

18. Recalibration

File 2131rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: -0.004units

- secondary salinity –  offset: -0.014units

- Chl_fluorescence -  offset: -0.12g/l

The SAM files were recalibrated and named SAC; these files were used to rerun COMPARE. The results (2131com2.xls) indicate that the deep primary salinity values after recalibration were within 0.00005units of the bottles. (See plots in Processing Section of Report.)

19. REMOVE CHANNELS and CORRECT FORMAT

The following channels were removed from all casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The PAR channel was removed for casts #: 44,50,54-89,102,109,115-142.

The same channels were removed from the rosette files.

The data format was corrected as needed using HEAD EDIT and the files named *.CTD.

20. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data

21. Producing final files

a.) A cross-reference listing was produced.

b.) The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 31

Dates: Start: 17 September 2001 End: 2 October 2001

Location: NE Pacific

Vessel: CCGS John P. Tully

Party Chief: Dave Mackas

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550

Cruise ID#:		2001-31

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 18/07/01 | Factory |  |  |
| Conductivity | 2102 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2374 | 18/07/01 | “ |  |  |
| Secondary Cond. | 2128 | 17/07/01 | “ |  |  |
| Fluorometer -pumped | 2229 | 01/07/01 | IOS |  |  |
| SBE43 Oxygen sensor | 0047 | 16/05/01 | ? |  |  |
| Transmissometer | 192DR | 26/05/01 | IOS |  |  |
| PAR | 4615 | 15/12/00 |  |  |  |
| Pressure Sensor | 75636 | 04/06/99 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.