## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

PROCESSING NOTES

Cruise: 2001-33

Agency: Aquaculture/OSAP

Location: Broughton Archipelago

Project: Aquaculture - Broughton

Party Chief: Colin Levings

Platform: CCGS VECTOR

Date: September 23, 2001 – October 4, 2001

Processed by: Germaine Gatien

Date of Processing: 11 April 2002 – 7 May 2002

Number of original CTD casts: 46

Number of casts processed: 46

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with Transmissometer 136, Par Sensor #4601 and Seapoint Fluorometer #2350. The deck unit was #0508.

SUMMARY OF QUALITY AND CONCERNS

The data was in good order. There is a lot of fine-scale noise in the secondary conductivity sensor.

Chlorophyll samples were taken but analysis had not been done at the time of processing.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There were errors in the identification of the pressure sensor and the date of calibration. The calibration constants were correct. A new file was prepared with the correct calibrations and was named 2133CTD.con. An error was later found in the PAR calibration – this was corrected in step 16.

The sensor history was found. In previous use the secondary conductivity sensor has displayed a lot of fine-scale noise.

- Conversion of Raw Data

The raw data was converted using configuration file 2133CTD.con. Errors in the station names for casts 19, 20 and 21 were corrected in the CTD and rosette files. A preliminary check shows all expected channels present.

The temperature and salinity traces track reasonably well and up and downcasts are similar. The secondary salinity shows fine-scale noise (on the order of 0.005 units) as was noted in other cruises using these sensors. This is due to noise in the secondary conductivity channel.

4. STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels. The bottle number channel was stripped from the CNV files.

5. ALIGNCTD

The deck unit is one of the newer versions that advance both the primary and secondary conductivity channels by 0.073s. Tests of alignment were made on cast #6 using advancements of –0.013s, -0.018s,     -0.023s, and –0.028s and –0.018s improved the results most. That value was also used for 2000-28, 2000-30 and 2000-32. Align was run on all casts using –0.018s for a net advancement of +0.055s.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   Pass 1    Std Dev = 2	            Pass 2    Std Dev = 5		            Points per block = 50

6. CELLTM

Tests were run on casts #6 to check that the parameter choice used for 2001-32 (alpha = 0.02; 1/beta = 9.0) is appropriate for this data as well. It worked well on both channels so the conductivity cell thermal mass correction was done for both channels on all casts using alpha = 0.02; 1/beta = 9.0. It was noted that the choice of (alpha = 0.02; 1/beta = 7.0) worked equally well.

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 6 | 400 | -0.0006 | +0.00015 | +0.002 | Steady, high |
| 7 | 250 | -0.0006 | +0.00015 | +0.002 | Steady, high |
| 38 | 415 | -0.0006 | ?noisy | +0.002 | Mostly ok, but low at bottom |
| 42 | 475 | ?noisy | +0.00010 | +0.0015 | Extremely noisy |

These values are similar in sign to those found for 2001-32, but the conductivity and salinity differences are a little lower than for that cruise.

All casts were plotted on-screen to check for problems with processing. Reasonable agreement was found between up and down casts and between pairs of sensors, except for the bottom of cast #35. The problems in that cast are at the bottom and upcast; DELETE should produce good results.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

10. Checking Headers

A header summary and a header check were produced.  Errors in positions for casts 19, 20 and 21 were fixed. The average surface pressure is 2.5db

The cruise track was plotted and looks reasonable. One cast looks like it is on land but the map is crude and the log and GPS agree.  Dario Stucchi confirmed that the position is correct.

- DELETE

- CLEAN was run to replace pad values in pressure with interpolated values.

- The following DELETE parameters were used:

- Surface Record Removal: Last Press Min

- Maximum Surface Pressure (relative): 10.00

- Pressure Tolerance: 1.0

- Pressure filtered over 15 points

- Swells deleted. Warning message if pressure difference of 2.00

- Drop rates < 0.3m/s (calculated over 15 points) was deleted.

- Sample interval =  .04 seconds.

- COMMENTS ON WARNINGS:

- There were no warnings. Comparing the maximum pressure before and after DELETE shows that little data was lost from the bottom of casts. (See delete_study.xls.)

- All DEL files were copied to EDT files.

- 12. COMPARE

- The BOT files were examined for errors and no bad data found.

- COMPARE was run. The primary salinity was found to be low by 0.016 units and the secondary low by 0.014 units. This is consistent with the sensor history. There was no obvious drift with time. There continues to be a lot of fine-scale noise in the secondary sensor so the primary sensors will be chosen for further processing.

- DETAILED EDITING

The primary sensors were chosen for further processing because the secondary conductivity is noisy.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

The following casts were edited at the surface or bottom only: 4,18,19,23,24,27,28,30,32,35,38,44,46.

The following casts were edited more extensively: 3,6-8,11-16,39-42,45.

The descent rate was noisy for casts #39-42. The rate was generally kept high. There were problems with shed wakes and poor sensor response near the surface and the bottom, and during the casts with noisy descent rate.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

14. SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. Casts #42, 43 and 44 were examined before and after this step and showed that the filter was effective.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

The dark value is about 0.05g/l. The maximum value for the cruise is about 9g/l for cast #45. Chlorophyll values were not available at the time of processing.

15. BIN AVERAGE

The following Bin Average values were applied to the BOX files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used	Interpolated values are NOT used for empty bins

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

15. Other comparisons

Previous experience with these sensors – Both pairs of sensors were used during 2000-07, 2000-06, 2000-30, 2000-28, 2001-04 and 2001-32 in April, May, September and October of 2000 and March and September of 2001 respectively. The primary sensors were selected for archiving in all cases since the secondary conductivity was full of fine-scale noise. The CTD salinity values were lower than bottles with differences varying from –0.009 to –0.021 units for the primary and –0.004 to –0.017 units for the secondary. The largest differences occurred for the most recent cruises.

Historic ranges – The ranges were available for only 2 casts. The temperature and salinity were within those ranges except near the bottom where the range data is limited.

- Recalibration

File 2133rcal.ccf was used to recalibrate the AVG files (including those in the IOS directory, the HYDRO directory and the FLUOR directory) as follows:

- primary salinity –  offset: +0.016 units

- secondary salinity –  offset: +0.014 units

- Chl_fluorescence -  offset: -0.05g/l

- PAR – slope: 1.008626

The SAM files were recalibrated and named SAC and COMPARE was rerun on the recalibrated data. The results are satisfactory with average differences of 0.00002 and –0.0003 units for the primary and secondary salinity respectively.

The CHE files were recalibrated using file 2133rcal1.ccf since the channel names were slightly different.

17. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

18. REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and data format corrected as needed.

19. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

### Particulars

19 – Error in station name noted in log; corrected after conversion.

20 – Error in station name noted in log; corrected after conversion.

21 – Error in station name noted in log; corrected after conversion.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001- 33

Dates: Start: 23 September 2001 End: 4 October 2001

Location: Broughton Archipelago

Vessel: CCGS VECTOR

Party Chief: Colin Levings

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2001-33

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2663 | 07/03/00 | Factory |  |  |
| Conductivity | 2399 | 03/03/00 | “ |  |  |
| Secondary Temp. | 2371 | 16/06/99 | “ |  |  |
| Secondary Cond. | 2424 | 03/03/00 | “ |  |  |
| Fluorometer -pumped | 2350 | 07/01 | IOS |  |  |
| Par Sensor | 4601 | 09/02/01 |  |  |  |
| Transmissometer | 136 | 13/09/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |