## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2001-39

Agency: OSAP

Location: Knight Inlet

Project: Zooplankton Aggregation at Sills

Party Chief: Dave Mackas

Platform: CCGS VECTOR

Date: 12 - 26 Novemeber, 2002

Processed by: Germaine Gatien

Date of Processing: 25 March 2002 – 27 March 2002

Number of original CTD casts: 17

Number of casts processed: 17

INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD  (#0443) was mounted with Transmissometer #333DR and Seapoint Fluorometer #2229. The CTD deck unit was a SeaBird model 11, s/n 0508.

SUMMARY OF QUALITY AND CONCERNS

The CTD log contained no details about which sensors or fluorometer cable were used.

There were odd jumps in pressure in one cast.

There was insufficient data to recalibrate salinity. It should be considered to be ±0.004.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files have extension DAT.

2. Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration file was obtained and the calibration constants were checked. There was an error in the date of the pressure calibration which was corrected in file 2139CTD.con.

The sensor history was found.

- Conversion of Raw Data

The raw data was converted using conversion file 2139CTD.con

An initial examination of the data shows severe spiking in conductivity in cast #12, but that is at the bottom or during the upcast; most of the downcast is ok. Otherwise the two pairs of sensors appear to track reasonably well.

4. ALIGNCTD

The deck unit used for this mission advances both conductivity channels by 0.073s. Tests on casts #6 and 20 showed little difference when the advancement was varied from 0.083s to 0.045s; the deck unit choice looked best so this step was skipped.

5. WILDEDIT

Program WILDEDIT was used to remove spikes in all channels.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

6. CELLTM

A study was made of the effect of varying the parameter choices for cast #20. The best results were obtained with alpha = 0.02 and 1/beta = 9.0. The conductivity cell thermal mass correction was done for both channels using alpha = 0.02 & 1/beta = 9.0.

7. DERIVE

Program DERIVE was run twice:

- on all casts to calculate primary and secondary salinity.

- on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

8. Test Plots and Channel Check

There were only 6 casts deeper than 400db. The descent rate was very steady, on the order of 1m/s. The differences in temperature, conductivity and salinity were about 0.002C, 0.00022units and 0.0008units, respectively below 400m. The temperature and conductivity differences are higher than those found during 2001-30 when the same equipment was used and the salinity differences are lower but of the opposite sign. However, the observations for 2001-30 are based on only 2 casts for which the descent rate was very noisy.

9.  Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

10. Checking Headers

A header summary was produced and a few errors in station names were found and corrected.

The surface check was run. The average surface pressure is 1.2db.

The header check was produced and no errors were found.

The cruise track was plotted and looks reasonable.

11. Test Plots

Profiles of all casts were examined for any evidence of problems with the processing. The downcasts and upcasts are reasonably close as are the two temperature and salinity traces. There is noise in both channels but the primary salinity seems slightly better than the secondary.

The fluorescence channel has a dark value on the order of 0.145μg/l which corresponds to about 0.05V since it was converted assuming a gain of 10X. The peak values at most casts are about  0.4 to 0.7μg/l, but the peak at cast #1 is almost 15μg/l. On closer examination that peak looks like a spike and the largest values below 2db is about 5μg/l which may be ok for Saanich Inlet.

- DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 20.00

Surface Swell Pressure Tolerance: 0.5

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS

The only warnings were for cast #25 and concern sudden jumps in pressure that are present in the original data. These jumps are seen in both the downcast and upcast at similar pressures which might suggest a problem with the equipment. This was reported to Doug Anderson.

All DEL files were copied to EDT files.

- Median Filter

A median filter with fixed size 11 was applied to the Seapoint fluorometer channel to reduce spikiness. Casts #1 and 23 were examined before and after this step and the filter was found to be effective.

- Initial comparisons

Previous experience with these sensors –

| Mission | Start Date | Sensors 1764/2023 | Sensors1763/2106 |
| --- | --- | --- | --- |
| 2101 | 01/01/01 | 1764 with diff. Temp sensor(+.0004) | No |
| 2106 | 05/02/01 | 1764 with diff. Temp sensor(+.0011) | No |
| 2110 | 17/04/01 | No | -0.003 |
| 2113 | 19/05/01 | No | +.0033 |
| 2108 | 28/05/01 | No | +0.0023 |
| 2119 | 25/06/01 | 1764 with diff. Temp sensor(-.0044) | No |
| 2130 | 2/10/01 | ~0(shallow bottles, large scatter) | ~-0.05 (shallow bottles, large scatter) |

The only one of these cruises for which there were lots of deep bottles was 2001-06.

COMPARE – There were 3 hydro bottles taken at approximate depths 200, 165 and 507m. File 2139comp.xls shows that comparison of the bottles with the CTD files (averaging over 21 points at the estimated depth of the bottle). The gradient was fairly high for the bottle for cast #13. For the other two casts the primary salinity was low by 0.002 and 0.004units and the secondary was low by 0.001 and 0.003units.

Historic Ranges – None available.

- DETAILED EDITING

Page plots were produced using T0,S0 for all casts. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT.

CTDEDIT was used to delete surface records and to clean salinity lightly in all casts except #1, 5 and 7.  A few records were removed from the bottom of cast #12. More extensive cleaning was done on cast #13. Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

16. BIN AVERAGE

The following Bin Average values were used for EDT files:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used

Interpolated values are NOT used for empty bins

17. Final comparisons

Comparison of nearby casts – T-S plots were produced of nearby casts and compared to check for problems with processing. No such problems were found.

- Recalibration

Given the lack of bottle data and the great variability in the sensor history, it was decided not to recalibrate the salinity data.

File 2139rcal.ccf was used to recalibrate the fluorometer data only by subtracting the observed dark value, 0.145μg/l.

19. Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

20. REMOVE and REORDER

The following channels were removed from all casts: Scan_Number, Secondary Temperature, Secondary Salinity, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and data format corrected as needed.

21. Producing final files

a.) The final files were renamed CTD.

b.) A cross-reference listing was produced.

c.) The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2001-39

Dates: Start: 12 November 2001 End: 26 November 2001

Location: Knight Inlet

Vessel: CCGS VECTOR

Chief Scientist: Dave Mackas

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0443 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0443

Cruise ID#:		2001-39

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 01/Nov/00 | Factory |  |  |
| Conductivity | 1764 | 28/Sep/00 | “ |  |  |
| Secondary Temp. | 2106 | 02/Nov/00 | “ |  |  |
| Secondary Cond. | 1763 | 31/Oct/00 | “ |  |  |
| Transmissometer | 333DR | 11/Nov/01 | IOS |  |  |
| Pressure Sensor | 63507 | 11/Jan/96 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.