## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 14 April 2007 | Reprocessed data including fluorescence channel. See note at end. |

## PROCESSING NOTES

Cruise: 2002-04

Agency: OSAP

Location: Broughton Archipelago

Project: Aquaculture

Party Chief: Stucchi, D.

Platform: VECTOR

Date: 22 January 2002- 30 January 2002

Processed by: Germaine Gatien

Date of Processing: 18 June 2002 – 7 August 2002

Number of original CTD casts: 18

Number of casts processed: 18

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#425) was mounted with a transmissometer (S/N498DR), SBE 43 Dissolved Oxygen Sensor (S/N47) and Seapoint Fluorometer with a 10X gain (S/N2356). The serial numbers of the external sensors were taken from the configuration file, but could not be confirmed as they were not recorded in the Daily Log. The deck unit S/N was 0425.

# SUMMARY OF QUALITY AND CONCERNS

The data was generally in good order. Cast #18 was corrupted and required special treatment.

The SBE Oxygen and Seapoint Fluorometer data will not be archived at this time.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Daily Log was obtained. Salinity calibration data was obtained. The cruise summary sheet was completed. The configuration files were obtained and the calibration constants were checked.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 0585ctd.con.

A preliminary check shows all expected channels present. The oxygen data is unreasonably smooth and upcast and downcast look quite different, suggesting the same problems as have been noted in other cruises. The secondary conductivity sensor is full of fine-scale noise as has been noted in other cruises using this equipment. The dark value for the fluorescence is about 0.06g/l.

Cast #18 was full of spikes. FIXDAT was used to repair the data file and the output was called 2002-04-9918. It was then converted and the output of that was renamed 2002-04-0018.

The temperature and salinity traces track reasonably well and up and downcasts are similar.

##### STRIP

The salinity channel was stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   	Pass 1    Std Dev = 2;

Pass 2    Std Dev = 5;

Points per block = 50

##### CELLTM

Tests were run on casts #1, 5 & 17 to find the optimal parameter choice for CELLTM. Runs using (0.02,9), (0.03,9), (0.03,7) and (0.03,7) were used for (alpha, 1/beta). The overall best choice was found to be (0.03,9) but the differences were not large; CELLTM was run on all casts using (0.03,9).

##### ALIGNCTD

Tests were run on casts #1,5,9 and 18 using advancements of +0.045s, 0.050s, 0.055s, 0.060s, 0.073s, 0.090s and 0.10s to the conductivity channels. The best results were found using +0.090s for the primary sensor and +0.060s for secondary. Since the deck unit advanced the primary by +0.073s ALIGNCTD was used to advance the primary by +0.017s and secondary by +0.060s for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The secondary conductivity and salinity are much noisier than the primary. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 7 | 450 | +0.001 | -0.0003 | -0.004 | ~1m/s, steady |
| 14 | 400 | +0.001 | -0.0003 | -0.004 | ~1m/s, steady |
| 16 | 450 | +0.001 | -0.0003 | -0.004 | ~1m/s, steady |
| 17 | 670 | +0.0008 | -0.0003 | -0.0035 | ~1m/s, steady |

These results are similar to those found with the same equipment during 2002-03 in early January 2002.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. None were noted.

##### Checking Headers

A header summary and a header check were produced and no errors found.

The average surface pressure is 1.6db but the surface for cast #18 is 32db. Looking at that cast there is some corruption at the beginning of the file. A text editor was used to remove the first 71 scans from the CNV file and it was then reconverted; the surface pressure was then 1.8db.

The cruise track was plotted and looks reasonable.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted.

Sample interval =  .0417 seconds.

COMMENTS ON WARNINGS:

There were no warnings in the DELETE log.

Comparing the maximum pressure before and after DELETE shows that an average of 0.9m of data was lost. Casts with a loss of more than 2m of data were checked and all the lost data was associated with very low descent rate. (See delete_study.xls.)

All DEL files were copied to EDT files.

##### COMPARE

The BOT files were examined for errors. The primary salinity for bottle #1 of cast #12 is bad. This is due to corruption of the conductivity channel. There is also minor noise in T0, T1 and C1.

COMPARE was run. When only bottles from below 200db are included in the fit the primary salinity was found to be low by about 0.0009 units and the secondary, low by 0.0048 units. Two deep bottles were excluded from the comparison; cast #3 bottle #1 showed a lot of variability and cast #12 bottle #1 was very spiky. There was no notable pressure dependence. There is a hint of time-dependence but given few data points and a great variation in areas sampled, it may not be significant. The CTD values are lower than during 2002-03 and 2002-01 which occurred before and after this cruise so there is no evidence of long-term drift. As found for 2002-01 there differences are small and the primary salinity needs no recalibration. (See 2204comp.xls)

##### DETAILED EDITING

The primary sensors were chosen for further processing because they were very stable and showed much less noise than the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. There were many unstable features not associated with shed wakes; these were assumed to be real and were left unedited. Only light editing was done of suspicious near-surface records and a few small spikes in salinity.

The following casts were edited at the surface and/or bottom only: 1,7,8,10,13,15.

The following casts were edited more extensively: 11,14,17.

The descent rate was kept high so there was no problem with shed wakes except at the surface and bottom.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

The BOT file for cast #12 could not be edited until it was put through CLEAN. CTDEDIT was then used to remove the primary conductivity and salinity from bottle #1 because those channels were full of spikes. The original file was renamed BO1 and the edited file was named BOT. COMPARE was not rerun since this data had not been used in the original comparison.

##### SPECIAL FLUOROMETER PROCESSING

Chlorophyll values were available for cast 16. These were much lower than the fluorometer readings. When plotted against each other the data fell close to a straight line FL = 3.2 * CHL + 0.046.

The fluorometer data will not be processed further at this time; it will not be archived until more information is obtained about the calibration of the instrument.

##### BIN AVERAGE

The following Bin Average values were applied to the EDT files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – All sensors were used in January 2002 (2002-03) when the primary salinity was found to be high by 0.0038 units and the secondary low by 0.0005 units. They were also used in February 2002 when the primary was high by 0.0011 units and the secondary high by 0.0036 units.

Historic ranges – the only ranges available were from the World Ocean Data Base (arranged by Marsden squares); these do not represent this area well.

##### Recalibration

The SBE 43 oxgyen data and the Seapoint fluorescence will not be archived at this time so no attempt was made to recalibrate those channels. Further research may be put into the processing of this data and the channel may be added to the archive in the future.

The CTD and bottles are in reasonable agreement so no recalibration of salinity will be done.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Fluorescence:Seapoint, Oxygen:Dissolved and Flag.

The channels were reordered and data format corrected as needed.

The Standard Check routine was run and the only problem found was that the scientist is not in the standard scientist list. The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

3. Bottle #1 – lot of variability – not used in COMPARE

12. Bottle #1 – spiky primary conductivity – not used in COMPARE

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-04

Dates: Start: 22 January 2002 End: 30 January 2002

Location: Broughton Archipelago

Vessel: Vector

Party Chief: Stucchi, D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2002-04

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 11/09/01 | Factory |  |  |
| Conductivity | 1729 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2968 | 18/09/01 | “ |  |  |
| Secondary Cond. | 2664 | 14/09/01 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| Oxygen SBE43 | 0047 | 11/26/01 | Factory |  |  |
| Transmissometer | 498DR | 01/08/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |

April 14, 2007 – Germaine Gatien

The fluorescence was originally stripped from these files (and from those of some other cruises in 2002) because the comparison with titrated chlorophyll was considered questionable. It was noted at the time that this should be revisited when we knew more about data from this instrument. It is now believed that this data is suitable for archiving, but as always it should be considered nominal.

The fluorescence in the edited files (EDT) were filtered using a median filter, size 11, then the files were put through BIN AVERAGE as described in the original processing. REMOVE was run as before except that Fluorescence was NOT removed.

HEAD EDIT was run to fix formats and channel names and to add a header comment to indicate that both transmissivity and fluorescence are nominal and unedited. (2002-01-hdr.txt)