## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 3-Mar-2014 | Added Mission number to the rosette files. J.L. |
| 14 April 2007 | Reprocessed data including fluorescence channel. See note at end. |

## PROCESSING NOTES

Cruise: 2002-05

Agency: OSAP

Location: North-West Pacific

Project: High Seas Salmon

Party Chief: John Morris

Platform: Ricker

Date: February 25, 2002 – March 18, 2002

Processed by: Germaine Gatien

Date of Processing: 19 August 2002 – 10 September 2002

Number of original CTD casts: 87

Number of casts processed: 85

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer 197, and Seapoint Fluorometer S/N #2356 with a 10X gain. The deck unit S/N was not recorded.

# SUMMARY OF QUALITY AND CONCERNS

Several casts were corrupted by spikes or aborted partway through; the problem was believed to be with slip rings.

For some casts one pair of sensors was corrupted by spikes while the other pair looked ok.

For some casts close to shore it is not clear whether unstable features are caused by poor sensors or are real, so they were left unedited.

Other problems were noted that appear to be caused by inconsistent alignment of sensors or variations in flow rates. This could be due to the CTD tilting out of the vertical and/or the mounting of external sensors in a way that interfered with the flow.

The descent rate was extremely noisy for most casts, but the average was quite high minimizing shed wakes. As has been noted for other CCGS Ricker cruises during which this CTD system was used, there is a jitter in the pressure sensor of about 0.25db.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There are many entries related to problems with the CTD.

According to the log the Niskin bottle was mounted 5m above the CTD for all casts.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. An error in the date of transmissometer calibration was corrected and the resulting file named 0506CTD.con.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 0506CTD.con.

A preliminary check shows all expected channels present. A number of casts have serious data problems, but for most the temperature and salinity traces track reasonably well and up and downcasts are similar. Some of the problem casts have entries in the log explaining what went wrong. The fluorometer data looks reasonable and the dark values for fluorescence are about 0.037g/l.

##### STRIP

The salinity channels were stripped from the CNV files so that DERIVE doesn’t create a 2nd set.

##### ALIGNCTD

The deck unit serial number is unknown. Tests indicate that it is one of the older versions that advance only the primary conductivity channels by 0.073s.

Tests of alignment were made on casts #25, 111 and 249 using net advancements of 0.095s, 0.073s, 0.060s, 0.055s, 0.050s and 0.045s. For each cast a 10db section was found in which the temperature gradient is large. After running ALIGN the results were examined in T-S space to see which produced the smoothest salinity. The best results were obtained with 0.050s – 0.055s for the primary and 0.060 – 0.073s for the secondary.

Align was run on all casts using –0.018s for the primary conductivity and  +0.060s for the secondary  conductivity for net advancements of +0.055s and +0.060s, respectively.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

Tests were run on casts #18, 99 and 195 to find the optimal parameter choice for CELLTM. Runs using (0.0245,9.5), (0.2,9), (0.3,9), (0.3,7) and (0.3,7) were used for (alpha, 1/beta). The overall best choice was found to be (0.3,9) but the differences were not large; CELLTM was run on all casts using (0.3,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present.

The secondary conductivity and salinity are slightly noisier than the primary.

The two sets of sensors are offset in pressure, with the secondary often seeming deeper than the primary by almost 1db. This may be an indication of different response time or flow-rate. There seems to be less fine-scale detail in the secondary conductivity.  The differences between the pairs of sensors is much higher during the upcasts than the downcasts and quite one-sided again suggesting that the flow-rate was different for the two pairs of sensors.

The differences between sensors were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 25 | 1800 | -0.0012 | ~+0.00002 | +0.001 | Extremely noisy, high average |
| 111 | 1000 | -0.0008 | ~0 | +0.0008 | Extremely noisy, low average |
| 249 | 1000 | -0.001 | ~0 | +0.001 | Extremely noisy, medium average |

The fluorometer data looks reasonable with dark values of 0.035g/l and maxima of about 3.6g/l near the mouth of Juan de Fuca Strait and near Estevan Point.

The transmissivity has maximum values of 90.3% in deep water, close to the maximum possible value of 91.3%. There is no sign of pressure-induced hysteresis.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. The event numbers were not in the headers so all casts were put through CLEAN to add event numbers based on the file names. Pad values were replaced by interpolated values, and header min/max values reset at the same time.

All casts were plotted and checked for evidence of problems with the processing or instruments. There are many questionable features; the log noted many problems with slip rings. The data will need careful checking during editing.

##### Checking Headers

The cruise track was plotted and the positions look reasonable.

A header summary and a header check were produced. In checking against the log it was found that two casts were missing (#126 and #145). These were later found and processed in the same way as all others.

The average surface pressure is 4.3db. This is higher than the usual 2.5db for Ricker cruises, so a few casts were examined in detail. At the end of cast #8 the pressure was as low as 0.2db and during #111 about 0.9db, both at the end of the upcast. So the deeper start to the casts looks deliberate rather than due to a pressure calibration problem.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00 Minimum Salinity: 10

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .042 seconds.

COMMENTS ON WARNINGS:

There were warnings in the DELETE log for three casts. They applied to the upcast in one case and to CTD reversals during the downcast for two others; in all cases DELETE did not remove good data.

Comparing the maximum pressure before and after DELETE shows that an average of 0.9db of data was lost at the bottom of the casts. (See delete_study.xls.)

All DEL files were copied to EDT files.

##### COMPARE

There were no rosette bottles but hydro samples were taken at depth for 8 casts and at the end of most upcasts when the CTD was at 15db. The hydro bottle was mounted 5m above the CTD.

The deep bottle values were compared with metre-averaged CTD data from 5db above the bottom. The primary and secondary salinity were both found to be high by 0.003. The largest difference was for a cast that had an extremely noisy descent rate. (Note: The depths recorded for the deep bottles were wrong in 4 cases; Hugh McLean confirms that all deep bottles were done at the bottom of the casts.)

The surface bottles were examined for all casts using metre-averaged CTD data from the upcasts. Only casts with very well-mixed surface waters were selected and the salinity was taken from metre-averaged upcast data files, matching the pressure at which the bottle is believed to have been fired. The average differences were –0.0001units for both primary and secondary. However, if outliers are ignored the CTD appears to be low by about 0.004 and 0.002units.

The data was examined to see if there was any trend in the surface differences with time and none is evident. There was a group of casts in the middle of the mission with higher differences, but this is likely to be related to sea conditions and variability in how well-mixed the surface waters were, rather than a problem in the calibration. There is some hint of depth-dependence in the differences, but there is insufficient data to be confident about it. There is insufficient information to justify recalibrating the data. The salinity should be considered ±0.005. See 2002-05_comp.xls

##### DETAILED EDITING

The primary sensors were chosen for further processing because they were very stable and showed much less noise than the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Other unstable features were left unedited since the data comes from very active mixing regions.

The following casts were edited at the surface or bottom only: 5,8

The following casts were edited more extensively: 14-47,52-81,87,90,96-148,154-270.

The following casts required special treatment:

45 – The two pairs of sensors give very different results both for T and S. The surface bottle agrees with the secondary salinity, but below 10db the secondary looks bad in T-S space. DELETE.

49 – T1,S1 full of spikes; T0,S0 show poor response, large loops in T-S curves, probably trouble with flow rate related to extremely noisy descent rate but not clearly shed wake problem. Cast deleted.

72,78 – The conductivity near the surface was bad in the primary channel. This does not appear to be related to the descent rate. T1,S1 chosen for editing. T1 and S1 are noisy with many small instabilities.

99 – Primary conductivity bad from 20-40db following a large spike in all channels around 20db. T1,S1 chosen for editing.

120 – Pumps off for top 11db. Records removed.

123 – Odd traces in top 27db, alignment of sensors odd, perhaps CTD at an odd angle or something interfering with the pumps. The primary temperature and conductivity are not aligned well, nor are the primary and secondary temperature with similar features occurring up to 1.5 db apart. Sharp T-S gradients can cause misalignment, but these differences seem too large to be accounted for that way. The top 27db of data was removed.

The descent rate, while frequently extremely noisy, was generally kept high minimizing problems with shed wakes.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

Extracted chlorophyll samples were available for many casts from 10db Niskin bottles. A comparison was made with metre-averaged upcast CTD fluorometer data and a good fit was found. (See 2002-05 chl_comp.xls.) Nevertheless there are serious questions about the reliability of the fluorometer data, so that channel will not be archived at this time. It is intended that this data set and those from other cruises using the fluorometer, be revisited over the next year with the hope of developing a process to ensure confidence in the values. The channel will be added to the archive if and when quality can be assured.

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

##### BIN AVERAGE

The following Bin Average values were applied to the BOX files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – No data has been processed using these sensors since the last calibration.

Historic ranges – The three deepest casts (1000m or more) fell within the ranges. Some of the data falls outside the historic ranges, but these are mostly in areas where there is probably little winter data in the range files or only WODB climatology arranged in Marsden squares; those ranges are not likely to be representative of conditions in inlets, near the surface and near shore.

Two casts were investigated in more detail using plots of multiple casts from the same regions. Cast #27 has salinity below the range minimum in the top 20db only. Casts #139 and 142 have temperatures well outside the historic range; they resemble each other, but not the other stations nearest to them. The secondary sensors look the same as the primary. On a T-S surface all casts are stable. There is no obvious source of error so the data is presumed correct.

##### Recalibration

No recalibration was done.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Chlorophyll:Fluorescence:Seapoint and Flag channels were removed from all casts except #72, 78 and 99.

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Chlorophyll:Fluorescence:Seapoint and Flag channels were removed from casts #72, 78 and 99.

The channels were reordered and data format corrected as needed.

The Standard Check routine was run and problems found were corrected using REORDER and EDIT HEADERS.

The final files were named CTD.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

2.   CTD not working.

5.   Pumps not on – for part of the cast

25. Upcast aborted partway up. Fluorescence weird around 80-90db. Salinity bottle value probably bad.

- CTD aborted  not working properly – slip ring problem

- CTD aborted – not working properly

45. The primary and secondary disagree. Secondary closer to bottle, but looks weird. Shallow. DELETE.

49. CTD performed badly – Secondary T,S full of spikes, Primary T-S full of large loops, poor sensor response maybe due to poor flow to sensors. Descent rate extremely noisy but problems not clearly due to shed wakes. DELETE cast.

- Transmissometer failed; upcast aborted. Data below 166db looks bad and was deleted.

62. CTD cast aborted – cast # was used for a net tow

72. Primary conductivity bad for top 25db at last. Secondary sensors selected.

78. CTD left on while bottle was lowered and tripped. Spikes in primary channels in the top 20db. Use secondary channels.

99. Large spike around 20db; primary conductivity bad between 20 and 40db. Use secondary sensors.

102. Niskin bottle lost.

120. Pumps turned on at 10m.

123. Primary and secondary temperature and conductivity separated in pressure – salinity odd, probably due to bad alignment. Below 25db the traces look ok.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-05

Dates: Start: February 25, 2002 End: March 18, 2002

Location: North-West Pacific

Vessel: RICKER

Party Chief: Morris, J.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | no | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2002-05

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 17/12/01 | Factory |  |  |
| Conductivity | 2278 | 13/12/01 | “ |  |  |
| Secondary Temp. | 2095 | 17/12/01 | “ |  |  |
| Secondary Cond. | 2280 | 13/12/01 | “ |  |  |
| Fluorometer –pumped | 2228 | 07/01 | IOS |  |  |
| Transmissometer | 197 | 21/01/01 | IOS |  |  |
| Pressure Sensor | 69698 | 12/11/97 | Factory |  |  |

April 14, 2007 – Germaine Gatien

The fluorescence was originally stripped from these files (and from those of some other cruises in 2002) because the comparison with titrated chlorophyll was considered questionable. It was noted at the time that this should be revisited when we knew more about data from this instrument. It is now believed that this data is suitable for archiving, but as always it should be considered nominal.

The fluorescence in the edited files (EDT) were filtered using a median filter, size 11, then the files were put through BIN AVERAGE as described in the original processing. REMOVE was run as before except that Fluorescence was NOT removed.

HEAD EDIT was run to fix formats and channel.