# Linda White 							July 25, 2002

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Dave Mackas

OSAP

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Coastal Shelf Frozen nutrients Cruise # 2001- 08				Nitrate+Nitrite, Silicate and Phosphate

Summary

Frozen seawater samples were thawed and analysed on a Technicon Autoanalyser according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

A medium check standard was analysed as a sample following each station.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning, middle and end of the analytical run.

All labware and nutrient tubes have been acid cleaned, soon to be capped and stowed.

## Observations

16 replicates were taken.  Laboratory temperature was quite high at 28C throughout most of the day.

## Quality Assurance data

## Standards

Nitrate + nitrite and silicate standards were prepared in the same concentrations as Wako-Sagami for intercalibration.

Linda> Wako by 0.56% for nitrate + nitrite, Linda<Wako by 0.09% for silicate.

Linda < Janet by 0.8% for phosphate.

No corrections were made to the data because the calibrations remain within 1% of Wako standard concentrations for nitrate + nitrite and silicate.

Wako standards analysed as sample: 	Wako	nitrite + nitrate 20µM	Silicate 50µM

Linda	19.68µM		50.65µM

### Coefficient of Variation % for standards analysed as a sample

Medium check 		N03 24.14µM/L		Si 80.01 µM/L		P04 3.21µM/L

(Theoretical)

Analysed		24.18+/-0.3		80.81+/-1.1		3.22+/-0.05

n = 7			n = 7			n = 7

C.V.%			1.2%			1.4%			1.5%

Pooled Standard Deviation

Sample Replicates: 	NO2+NO3 = 0.78	Si = 0.19		PO4 = 0.059					n = 12			n = 10			n = 12

This report includes a final data spreadsheet data, quality assurance results, peakbase calculated .anz files, and plots.

Excel Filename: Shelf200208nutr.xls.