## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 29-Mar-2003 | Archived bottle data. |

## PROCESSING NOTES

Cruise: 2002-10

Agency: OSAP

Location: Seymour & Belize Inlets/ Alison Sound/Frederick Sound

Project: NSERC/DFO-Paleoclimate and Oceanography

Party Chief: Thomson, R.E.

Platform: Vector

Date: 15 April 2002 – 29 April 2002

Processed by: Germaine Gatien

Date of Processing: 17 October 2002 – 28 October 2002

Number of original CTD casts: 52

Number of casts processed: 51

# INSTRUMENT SUMMARY

According to the configuration file a SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (498DR), SBE 43 Dissolved Oxygen Sensor (S/N #0047), PAR sensor #4601 and Seapoint Fluorometer (S/N #2356). The fluorometer had a 10X cable. The deck unit S/N is not listed but on the previous cruise it was #0508.

# SUMMARY OF QUALITY AND CONCERNS

There is no list of equipment on the first page. This cruise immediately followed 2002-12, but there were errors in the list for that cruise. It is assumed that the details in the configuration file are correct.

The year was wrong in the start time for all casts. The month and day were correct.

There is a lot of fine-scale noise in the secondary conductivity channel.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read. Note was made of an error in the time in the headers – the year is 1999 instead of 2002 in two places in each cast; this will be fixed later.

Data files were missing for casts #51 to #55 and #64 and 65; these were found later and processed in the same way as the other casts.

Salinity and dissolved oxygen calibration data were obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There was no list of equipment in the log book. This mission immediately followed 2002-12 and is believed to have used the same equipment. However, there were errors in the list for 2002-12. It will be assumed that the con file is correct but checks should be done later to ensure a calibration error has not been made.

The sensor history was found.

##### Conversion of Raw Data

There are small differences between the con file for cast #1 and all the other con files for this mission. The differences do not affect calibration, but those for casts #2 – #87 match those used for cruise 2002-12 that preceded this cruise.

The raw data was converted using configuration file 2002-10-0002.con; rosette files were converted separately with salinity channels included.

A preliminary check of a few casts shows all expected channels present. The secondary conductivity is full of fine-scale noise as was noted in 2002-14 and 2002-08.

The temperature and conductivity traces track reasonably well and up and downcasts are similar.

An error in station name was fixed for cast #73.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure only using the following parameters:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50;

Keep data within this distance of the mean = 0.

##### CELLTM

Tests were run on casts #11,27,40 and 87 to find the optimal parameter choice for CELLTM. Runs using (0.02,9), (0.03,9), (0.03,7) and (0.03,7) were used for (alpha, 1/beta). The overall best choice was found to be (0.03,9) but the differences were not large. The secondary channel was not checked since the noise in the signal masks any differences. CELLTM was run on all casts for both channels using (0.03,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. As noted during previous uses of this equipment the secondary conductivity and salinity are much noisier than the primary. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 13 | 620 | ~0.001 | ~0.0003 | ~0.002 | High, fairly steady |
| 27 | 365 | ~0.001 | ~0.0003 | ~0.002 | High, fairly steady |
| 50 | 370 | ~0.001 | ~0.0025 | ~0.0015 | High, fairly steady |
| 66 | 370 | ~0.001 | ~0.0025 | ~0.0015 | Low, fairly steady |

The temperature differences are similar to those observed during 2002-01, -03, -04, -08, -12 and –14. The conductivity and salinity differences are similar to those cruises of April and May which used this equipment but quite different from those of January and February.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. On conversion casts #3 and 73 were missing position information. Minor editing of the SeaBird files fixed this (separating time and position info into individual lines in the header), although it is not obvious why this was necessary.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. The primary conductivity shifted during cast #67; note is made of a problem in the log and casts #66 and #68 are at the same site. Cast #67 will be not be processed further. Cast #87 had markedly different up and downcasts but the pairs of sensors were in agreement so this is presumed to be a real feature, not an instrument problem.

##### Checking Headers

A header summary and a header check were produced. The station names for casts #79-85 are different from those in the Daily Log Book but make more sense when seen on a plot, so are presumed to have been changed on purpose.

The average surface pressure is 1.3db. The mixed layer depth indicates that the surface is not sufficiently well-mixed for shallow salinity calibrations to be useful.

The cruise track was plotted and looks reasonable.

##### Cast #11 had an initial partial cast; the CTD was returned to the surface and restarted probably because the pumps were not on. A full cast followed. A text editor was used on file 2002-10-0011.ios to select the full cast, so that DELETE does not patch the two together inappropriately.

##### SHIFT

The routine SHIFT was used to align the conductivity.  The deck unit was one of the newer versions that advance both channels by 0.073s. Tests were run on casts #13 and 27 and any shift from +0.2 records to +0.4 records improves the data. It is difficult to pick which value is best because there is so much fine-scale temperature variation. This test is easiest in areas of rapid but smooth temperature change. A value of +0.3 records was chosen which is roughly equivalent to +0.012s giving an overall correction of +0.085s. This compares reasonably well with the results for cruises 2002-08 and 2002-14 when 0.091s was the net alignment used. However, it is notably different from the results of 2002-01, -03 and –04. Two runs were needed as the shift to primary and secondary conductivity must be done separately.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

DELETE was run first with the minimum drop rate set to 0.3m/s as usual. After studying the results it appeared that too much data was lost. This cruise was mostly in protected waters. The descent rate was often very low, but was steady with virtually no reversals. So DELETE was rerun with the Low Drop Rate minimum set to 0.1m/s.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.1m/s (calculated over 9 points) was deleted.

Sample interval =  .0412 seconds.

COMMENTS ON WARNINGS:

There were no warnings in the DELETE log.

Because the CTD was started very close to the surface for many casts, CLIP was used to remove the top 2db of data. Such surface data is unreliable due to the presence of the ship and the low initial lowering speed of the CTD; CLIP is more efficient than editing almost every cast.

All CLP files were copied to EDT files.

##### 13. COMPARE

The BOT files were examined for errors and none were found.

COMPARE was run. There was a lot of scatter in the results. At first sight there appeared to be a drift with time for both pairs of sensors. Outliers were removed with the aim of producing a flat trend line for differences vs. pressure. When the same points were used in plotting differences against file pair number the temporal drift disappeared. The average difference indicates that primary salinity was high by 0.0007 units and the secondary by about 0.0033 units. (See 2210comp.xls)

- A second run of COMPARE was done using the DOX files (titrated oxygen values). There is no evidence of drift through the cruise and the comparison indicates a fit of Titrated oxygen = 1.1051 * SBE Oxygen – 0.1117. The slope is similar to those found in early 2001 and during 2002-03 and 2002-12. The slope was about 1.125 for 2002-14 and -08.

##### DETAILED EDITING

The primary sensors were chosen for further processing because there was much less noise in the salinity than is seen in the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C or where there were bi-polar spikes that are probably due to flow-rate problems. Unstable features were left unedited unless they were clearly instrumental.

Most casts required only light editing of salinity, some required the removal of a few records near the surface or bottom only.

Temperature and salinity were not edited for the following casts: 3,4,8,13,15,27,32,36,51-64,74-76.

Note was made of the editing details in the relevant files.

Plots were made of PAR, Fluorescence and Transmissivity vs. Pressure to identify any casts that needed editing. CTDEDIT was used to edit PAR, FL and TRANS for casts #51 – 66 only. Oxygen also has spikes, but since the data is considered unusable, no effort was made to edit it.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

The dark value was about 0.165g/l.

There were no chlorophyll samples. This data will not be archived until confidence in the calibration is achieved.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The primary sensors were used in January and February of 2002 (2002-03, -04 and -01) when the primary salinity was high by 0.0038units, low by 0.0009 and high by 0.0011 respectively. The comparison for cruise 2002-12 which ran immediately before this cruise gave salinities high by 0.0013 units. The sensors were used after this mission during 2002-14 and 2002-08 when there were few bottles and noisy CTD data; the comparison indicated salinities low by 0.0016 units.

The secondary sensors have been used only once since the last calibration; previous to that recalibration the conductivity sensor was associated with some variations with time and pressure. During 2002-12 which immediately preceded this cruise the salinities were found to be high by 0.0025 units, but the scatter was large. The sensors were used after this mission during 2002-14 and 2002-08 when there were few bottles and noisy CTD data; the comparison indicated salinities low by 0.0007 units

Historic ranges – There is no useful data available for this region.

Repeat casts – A time series at station SE2A included 5 CTDs. This is an area of great variability and small unstable features are common. Cast #51 is notably different from casts #52-55. The upcast for #51 is also very different from the downcast. This might suggest that the pump was not operating correctly but the transmissivity, which is unpumped, behaved in a similar fashion. The principal investigator is satisfied that these variations are real, not instrumental.

##### Recalibration

The SBE 43 oxgyen data will not be archived at this time due to poor performance. The calibration is reasonable when checked using bottles, but the time response when the CTD is in motion is unacceptable.

The fluorescence will not be archived at this time due to uncertainties about the calibration of the instrument; it will be studied and may be added at a later date.

The salinity will not be recalibrated since the differences from bottles are fairly small and scattered.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE, REORDER and HEADER EDIT

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE, Chlorophyll:Fluorescence:Seapoint and Flag.

The channels were reordered and data format corrected as needed. The units are missing from the PAR channel and were added using HEADER EDIT

The Standard Check routine was run and no problems found.

The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced. The years are wrong. This should be corrected in the final files before they are placed in the archive.

The sensor history was updated.

##### Particulars

9. Hit bottom

- Partial cast before full cast removed with a text editor.

67. Problems with primary conductivity. Cast #68 was run to replace it. Delete.

73. Station name wrong in header; fixed.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-10

Dates: Start: 15 April 2002 End: 29 April 2002

Location: Seymour/Belize Inlets

Vessel: Vector

Party Chief: Thomson R.E.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585		Cruise ID#:		2002-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 11/09/01 | Factory |  |  |
| Conductivity | 1729 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2968 | 18/09/01 | “ |  |  |
| Secondary Cond. | 2173 | 11/09/01 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| PAR | 4601 | 09/02/01 |  |  |  |
| Oxygen SBE43 | 0047 | 11/26/01 | Factory |  |  |
| Transmissometer | 498DR | 01/08/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |