## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2002-11

Agency: OSAP

Location: Broughton Archipelago/Georgia Strait Inlets

Project: Circulation and Oceanography of the Broughton Archipelago

Party Chief: Stucchi D.

Platform: Vector

Date: 19 May 2002 – 27 May 2002

Processed by: Germaine Gatien

Date of Processing: 22 October 2002 – 5 November 2002

Number of original CTD casts: 35 (including one aborted cast)

Number of casts processed: 34

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (498DR), SBE 43 Dissolved Oxygen Sensor (S/N #0047), PAR sensor #4601 and Seapoint Fluorometer (S/N #2356). The external pumped sensors were mounted on the primary pump. The fluorometer had a 10X cable. The deck unit S/N was #0508.

# SUMMARY OF QUALITY AND CONCERNS

There is a lot of fine-scale noise in the secondary conductivity channel.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read.

Salinity, dissolved oxygen and chlorophyll calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked.

The sensor history was found.

##### Conversion of Raw Data

There were a few minor errors (S/N of pressure gauge and date of oxygen sensor calibration) in the configuration files used at sea. Configuration file 2002-11-0585 was created with those errors corrected. The raw data was converted using configuration file 2002-10-0585.con; rosette files were converted separately with salinity and “number of bottles fired” channels included.

A preliminary check of a few casts shows all expected channels present. The secondary conductivity is full of fine-scale noise as has been noted in other uses of this equipment in 2002.

The temperature and conductivity traces track reasonably well and up and downcasts are similar.

Cast #5 was deleted and cast #5A was renamed as 5; the computer had rebooted during #5 and #5A was a repeat cast at the same site.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

Tests were run on casts #1,17 and 34 to find the optimal parameter choice for CELLTM. Runs using (0.02,9), (0.03,9), (0.03,7) and (0.03,7) were used for (alpha, 1/beta). All choices improved the data but the overall best choices were found to be (0.03,7) and (0.03,9). The secondary channel was not checked since the noise in the signal masks any differences. CELLTM was run on all casts for both channels using (0.03,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. As noted during previous uses of this equipment the secondary conductivity and salinity are much noisier than the primary. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 22 | 630 | .0011 | .00025 | .0015-.002 | Fairly high, steady |
| 25 | 674 | .001 | .00025 | .0015-.002 | High, steady |

The temperature differences are similar to those observed during 2002-01, -03, -04, -08, -12 and –14. The conductivity and salinity differences were extremely noisy due to the noisy secondary conductivity; the differences are slightly higher than during other cruises of April and May during which this equipment was used.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. The fluorometer channel is off-scale frequently and noisy for some casts. The PAR channel is empty for some casts.

##### Checking Headers

A header summary and a header check were produced and checked. No errors were found. The mixed layer depth indicates that the surface is not sufficiently well-mixed for shallow salinity calibrations to be useful.

The cruise track was plotted and looks reasonable.

##### SHIFT

The routine SHIFT was used to align the conductivity.  The deck unit was one of the newer versions that advance both channels by 0.073s. Tests were run on casts #15 and 16 using shifts between –0.1 record and +2 records; the best results were with shifts of +0.4 and +0.5. The data has only very small spikes in salinity so the effects are small. A value of +0.45 records was chosen which is roughly equivalent to +0.018s for an overall correction of +0.091s. This is the same result as for cruises 2002-08 and 2002-14 and close to that found for 2002-12 (net 0.085s). The secondary conductivity has not been shifted since it is very noisy and not to be archived.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted.

Sample interval =  .0412 seconds.

COMMENTS ON WARNINGS:

There were no warnings in the DELETE log.

Comparing the maximum pressure before and after DELETE shows that no more than 1db of data was lost from the bottom of each cast. (See 2211delete_study.xls.)

All DEL files were copied to EDT files.

##### 13. COMPARE

The BOT files were examined for errors and none were found.

COMPARE was run. There was a lot of scatter in the results. When outliers were removed the average differences indicate that primary salinity was low by 0.0006 units and the secondary high by about 0.0007 units. When all data was included the primary was high by 0.0018. (See 2211comp.xls)

- A second run of COMPARE was done using the DOX files (titrated oxygen values). There is no evidence of drift through the cruise and the comparison indicates a fit of Titrated oxygen = 1.062 * SBE Oxygen – 0.1642. The slope is similar to those found in early 2001 and during 2002-03, 2002-10 and 2002-12. The slope was about 1.125 for 2002-14 and -08. (There is a note from Bernard Minkley about problems with the oxygen Standard value; he recomputed the oxygen data using a different Standard value. See file  hydro_data2002-11.doc for details.)

- A comparison was also made between the chlorophyll data and the SeaPoint fluorometer data from averaged rosette bottle files. The chlorophyll concentration was high enough that many values were off-scale for the fluorometer. Ignoring any fluorometer values greater than 14.5, a reasonable fit was found for chlorophyll vs. fluorometer. The slope of Fluorometer/CHL varied from 1.03 to 1.24 depending on whether outliers were included or not. Looking at the distribution suggests that the fluorometer is giving a reasonably linear response for concentrations above 2. There is a suspicion that for low values the fluorometer is reading too high. Until further study is made of this phenomenon, researchers are advised to be careful in using data from the low end of the SeaPoint fluorometer scale. (See 2211chl_comp.xls.)

##### DETAILED EDITING

The primary sensors were chosen for further processing because there was much less noise in the salinity than is seen in the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C.

All casts except #20 required light editing; for most, surface records were removed and salinity was cleaned very lightly to remove small spikes. No casts required heavy editing.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

The dark value was about 0.085g/l.  As noted in the COMPARE section the fluorometer was often off-scale. The comparison with bottles is reasonably good for values greater than 2g/l, but caution is urged in using lower values as there are suspicions about the calibration. Some of the casts had very noisy fluorometer data even at depth.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The primary sensors were used in January and February of 2002 (2002-03, -04 and -01) when the primary salinity was high by 0.0038 units, low by 0.0009 and high by 0.0011 respectively. They were used in April 2002 during 2002-10 and -12 when salinities were high by 0.0007 and 0.0013 units. The sensors were also used in May during 2002-14 and 2002-08 when there were few bottles and noisy CTD data; the comparison indicated salinities low by 0.0016 units.

The secondary sensors have been used during 2002-10 and -12 in April 2002 when the salinities were found to be high by 0.0033 and 0.0025 units, but the scatter was large. The sensors were also used during 2002-14 and 2002-08 in early May when there were few bottles and noisy CTD data; the comparison indicated salinities low by 0.0007 units.

Historic ranges – The ranges available for this data set are primarily from the WODB climatology arranged in 1º Marsden squares, so are not useful for inlet observations. The data from the few casts for which local climatology is available all fell within those limits.

##### Recalibration

The SBE 43 oxygen data will not be archived at this time due to poor performance. The calibration is reasonable when checked using bottles, but the time response when the CTD is in motion is unacceptable.

The salinity will not be recalibrated since the differences from bottles are fairly small and scattered.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE and Flag.

The PAR channel was removed for casts 6 to 34 since there was no signal.

The channels were reordered and data format corrected as needed.

The Standard Check routine was run and problems found were corrected using EDIT HEADERS. The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

8. Hit bottom

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-11

Dates: Start: 19 May 2002 End: 27 May 2002

Location: Broughton Archipelago/Georgia Strait Inlets

Vessel: Vector

Party Chief: Stucchi, D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2002-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 11/09/01 | Factory |  |  |
| Conductivity | 1729 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2968 | 18/09/01 | “ |  |  |
| Secondary Cond. | 2173 | 11/09/01 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| PAR | 4601 | 09/02/01 |  |  |  |
| Oxygen SBE43 | 0047 | 11/26/01 | Factory |  |  |
| Transmissometer | 498DR | 01/08/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |