## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS. |
| May 4, 2003 | Added nutrient data to rosette files. |
| 24 Feb 2003 | The scientist’s name and cruise # were wrong in some casts. These were corrected using HEADER EDIT. |
| 7 Jan. 2003 | 24 casts were put through new version of DELETE to capture more surface data; see Step 11 for details of revision. |

## PROCESSING NOTES

Cruise: 2002-12

Agency: OSAP

Location: Strait of Georgia/Juan de Fuca Strait

Project: SoG-JdeF

Party Chief: Masson, D.

Platform: Vector

Date: 9 April 2002 – 13 April 2002

Processed by: Germaine Gatien

Date of Processing: 17 October 2002 – 31 October 2002

Number of original CTD casts: 73

Number of casts processed: 73

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (498DR), SBE 43 dissolved Oxygen Sensor (#0047), Seapoint Fluorometer (#2356) with a 10X cable and PAR sensor (#4601). The log says the external sensors were mounted on the primary pump, but the sensor numbers given for the secondary temperature and conductivity are wrong, so this may be incorrect as well. The deck unit was a Seabird 11+ (#0508).

# SUMMARY OF QUALITY AND CONCERNS

There is a lot of fine-scale noise in the secondary conductivity channel as has been observed during other cruises using this equipment. The primary channels were chosen for the archive. The fluorescence data was very noisy.

There were changes in the alignment and in the differences between pairs of T and C sensors between 2002-01 (February 2002) and this cruise in early April. Later cruises in April and May 2002 have similar characteristics to this cruise.

Some of the stops for bottles were short reducing the quality of the bottle file data.

The log book entries for sensor serial numbers contained many errors.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read.

Salinity and dissolved oxygen data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There are many inconsistencies between the information in the log book and that in the configuration files. The sensors listed in the log book as primary are shown as secondary in the configuration files; those listed as secondary temperature and conductivity in the log book are the primary conductivity and temperature respectively in the con file. It is assumed that the con file is correct, but checks should be done later to see if the results are reasonable.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2002-12-0001.con; rosette files were converted separately with salinity channels included.

A preliminary check shows all expected channels present.

The station names in casts #1, 54 and 63 were fixed.

The temperature and conductivity traces track reasonably well and up and downcasts are similar for the early casts, but later in the cruise at least one cast has upcast and downcast looking very different. This will be investigated later.

The secondary conductivity is full of fine-scale noise which has been observed in other cruises using this equipment in spring 2002 (2002-08, -10 and –14) which confirms that the configuration file is correct.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

Tests were run on casts #6,22 and 73 to find the optimal parameter choice for CELLTM. Runs using (0.0245,9.5), (0.02,9), (0.03,9), (0.03,7) and (0.03,7) were used for (alpha, 1/beta). The overall best choice was found to be (0.03,9). The tests were only run on the primary conductivity since the secondary signal is so noisy. CELLTM was run on all casts using (0.3,9) for both primary and secondary channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. There were no deep casts during this mission. As noted during previous uses of this equipment the secondary conductivity and salinity are much noisier than the primary. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 285 | +.001 | +.0002 | ~0.002 | High, fairly steady |
| 32 | 320 | +.0008 | +.00015 | ~0.001 | High, fairly steady |
| 50 | 420 | +.0007 | +.00025 | ~0.002 | High, slightly noisy |
| 73 | 380 | +.001 | +.0003 | ~0.0015 | High, fairly steady |

The temperature differences are similar to those observed during 2002-01, -03, -04, -08, -10 and –14. The conductivity and salinity differences are similar to those cruises of April and May which used this equipment but quite different from those of January and February.

An initial look at the other channels suggests that the fluorescence is very noisy. This channel was not put through WILDEDIT, so a test run was done to see if that would have helped; no improvement was noted with the use of WILDEDIT using 20, 50 or 100 scans per block.

The dissolved oxygen sensor continues to show poor time response.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. The following casts had problems:

#33 – there is an initial downcast before the full cast; a text editor was used to remove this data so that DELETE will select the full cast.

#52 – the up and down casts are significantly different in the pumped channels and the primary and secondary are very different. Transmissivity and PAR look ok.

##### Checking Headers

A header summary and a header check were produced. Errors were found in the positions for a few casts; the headers in the SeaBird headers were edited (just for format) and those casts reconverted.

The average surface pressure is 2.1 db.

SURFACE was also used to calculate mixed layer depth using a reference pressure of 7 db and maximum variation in the primary salinity of 0.002 units. The surface waters are not very well-mixed and hence not much use for calibration of salinity.

The cruise track was plotted and looks reasonable.

##### SHIFT

The IOS SHELL routine SHIFT was used to align the conductivity in two stages (one for each conductivity channel). The deck unit was one of the newer units that advance both conductivity channels by 0.073s. Tests were done on casts #28 & 72 by looking at salinity spikes in areas of high temperature gradient. The best choice varied from +0.2 records to +0.4 records, so +0.3 records was chosen for both conductivity channels. This is equivalent to an advancement of ~0.01s for a net advancement of 0.085s. This compares reasonably well with the results for cruises 2002-08 and 2002-14 when 0.091s was the net alignment used. However, it is notably different from the results of 2002-01, -03 and –04.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 9 points) was deleted.

Sample interval =  .042 seconds.

COMMENTS ON WARNINGS:

There were no warnings in the DELETE log.

Comparing the maximum pressure before and after DELETE shows that little data was lost from the bottom of casts. (See 2212_delete_study.xls.)

All DEL files were copied to EDT files.

REVISION: January 7, 2003 – New version of DELETE run on casts 1, 5, 6, 7, 8, 14, 15, 18, 19, 29, 31, 38, 40, 41, 42, 44, 46, 48, 53, 54, 56, 61, 71, 73. The low drop rate feature was applied from 10 db to (Max. pressure – 10 db). The log file for this second run is Delete2.log. Steps #14, 15, 16, 19 and 20 were redone for these casts. Editing was done with the view of keeping as much surface data as possible.

##### COMPARE

The BOT files were examined for errors.

Cast #52 looks odd; the IOS file is bad for that cast as well. All pumped channels look bad. The salinity calibration from 250 db is significantly different from both channels. The T and S data should be deleted from this file.

Cast #57 has a large range of values for temperature at the bottom, but looking at the full data file suggests that there really was a large gradient at the bottom.

COMPARE was run. (See 2212comp.xls) There is a large scatter and the differences vary greatly according to which points are rejected from the comparison. The primary salinity is high by about 0.001 units and the secondary by about 0.002 units. There is no obvious drift with time in the primary, but the secondary differences may be increasing slightly with time.

- A second run of COMPARE was done using the DOX files (titrated oxygen values). There is no evidence of drift through the cruise. Cast #52 was eliminated from the fit since all the pumped channels for that cast look bad. The comparison indicates a fit of Titrated oxygen = 1.1104 * SBE Oxygen – 0.0133. The slope is similar to those found in early 2001 and during 2002-03 and 2002-10. The slope was about 1.125 for 2002-14 and 2002-08. While this comparison shows that the calibration is quite stable, the oxygen data continues to show very poor response times and will not be archived except in the bottle files.

##### DETAILED EDITING

The primary sensors were chosen for further processing because they showed less variation with time and there was much less noise in the salinity than is seen in the secondary.

Page plots were produced using T0, S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C.

The following casts were not edited: 21,24,25,52,55.

The following casts were edited at the surface or bottom only: 3,12,14,15,17-19,23,26,30.

All other casts were edited more extensively, mostly for small spikes in salinity and bad surface records.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

REVISION – January 7, 2002 – The files listed in section 11 were all re-edited with the exception of casts #14 and 15 which were found to require no editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

The dark value was about 0.13g/l.

There were no chlorophyll samples.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The primary sensors were used in January and February of 2002 (2002-03, -04 and -01) when the comparisons indicated that the primary salinity was high by 0.0038 units, low by 0.0009 and high by 0.0011, respectively. They were used after this mission during 2002-10 when the primary salinity was high by about 0.0007 units and during 2002-14 and 2002-08 when the salinity was low by 0.0016. The secondary sensors were used for 2002-10 when the salinity was found to be high by 0.0033 units and during 2002-14 and 2002-08 when the comparison indicated that the salinity was low by 0.0007 units.

Historic ranges – Cast #4, near St. Juan Island, had salinity greater than the maximum of the historic range near 150m depth. Casts #3, #31 & #32, in the same general area, had salinity near the maximum at that depth. All other data for which historic ranges are available fell within those limits. The observations near St. Juan Island are likely real, since the salinity calibration was reasonably close and there is no evidence of instrumental problems affecting only those casts.  The principal investigator concurs.

##### Recalibration

The SBE 43 oxygen data will not be archived at this time due to poor performance. The calibration is reasonable when checked using bottles, but the time response when the CTD is in motion is unacceptable.

The fluorescence data is nominal; it has not been recalibrated or despiked. It will be archived.

The salinity will not be recalibrated since there is a lot of scatter in the comparison with bottles; other cruises using this equipment in early 2002 gave highly variable results ranging from salinity being low by –0.002 to high by +0.004.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD and RAC casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE and Flag.

All pumped channels were removed from cast #52 leaving only Pressure, Transmissivity and PAR.

The channels were reordered and data format corrected as needed.

The Standard Check routine was run and problems found were corrected using EDIT HEADERS. The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

1.   Station name missing from header; added after conversion to IOS Headers.

33. Initial downcast to about 8m before full cast; removed with text editor.

44. CTD hit bottom.

48. CTD hit bottom.

50. CTD hit bottom.

52. Up and downcasts very different in pumped channels and primary and secondary very different; pumped channels removed from final files.

54. Error in station name; fixed after conversion to IOS Headers.

60. CTD hit bottom.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-12

Dates: Start: 9 April 2002 End: 13 April 2002

Location: SoF/JdeFuca

Vessel: Vector

Party Chief: Masson,D

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2002-12

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 11/09/01 | Factory |  |  |
| Conductivity | 1729 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2968 | 18/09/01 | “ |  |  |
| Secondary Cond. | 2173 | 11/09/01 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| PAR | 4601 | 09/02/01 |  |  |  |
| Oxygen SBE43 | 0047 | 11/26/01 | Factory |  |  |
| Transmissometer | 498DR | 01/08/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |