# Linda White 							May 16, 2002

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Diane Masson

OSAP

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Coastal Frozen nutrients 2001- 12						Nitrate+Nitrite, Silicate and Phosphate

Summary

Frozen seawater samples were thawed and analysed on a Technicon Autoanalyser according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

A medium check standard was analysed as a sample following each station.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning, middle and end of the analytical run.

All labware and nutrient tubes have been acid cleaned, soon to be capped and stowed.

Pump platen tubing has been replaced, rinsed with 3N NaOH and DMQ water.

## Observations

No replicates were taken.

Station 39 surface sample was –0.6 um/l (for nitrate + nitrite) below the salt water baseline for the analysis. Adding 0.6um/l to each value, to avoid underestimating the nitrate content adjusted the entire profile.

Samples were in excellent condition and labelled well.

There were no duplicates sampled.

Analysis went very well.

## Quality Assurance data

## Standards

Nitrate + nitrite and silicate standards were prepared in the same concentrations as Wako-Sagami for intercalibration.

Linda> Wako by 0.56% for nitrate + nitrite, Linda<Wako by 0.09% for silicate.

Linda < Janet by 0.8% for phosphate.

No corrections were made to the data because the calibrations remained within 1% of Wako standard concentrations for nitrate + nitrite and silicate.

### Coefficient of Variation % for standards analysed as a sample

Medium check 		N03 24.14µM/L		Si 49.95 µM/L		P04 2.01µM/L

(Theoretical)

Analysed		24.2+/-0.23		49.9+/-0.3		2.00+/-0.02

n = 15			n = 15			n = 15

C.V.%			0.96%			0.60%			0.81%

This report includes a final data spreadsheet data, quality assurance results, peakbase calculated .anz files, and plots.

Excel Filename: coastal200212.xls.