## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS. |
| 15 April 2007 | Reprocessed data including fluorescence channel. See note at end. |

## PROCESSING NOTES

Cruise: 2002-14

Agency: OSAP

Location: West Coast Vancouver Island/Queen Charlotte Islands

Project: Nestucca / Optical / Zooplankton time series / Covariability

Party Chief: Mackas, D.

Platform: John P. Tully

Date: 29 April 2002 – 8 May 2002

Processed by: Germaine Gatien

Date of Processing: 2 August 2002 – 22 August 2002

Number of original CTD casts: 56

Number of casts processed: 56

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (498DR), SBE 43 Dissolved Oxygen Sensor (S/N #0047) and Seapoint Fluorometer (S/N #2356) mounted on the secondary pump, PAR sensor #4601 and SPAR #16504 The fluorometer had a 10X cable for some casts, 3X for others. The deck unit S/N was #0508.

# SUMMARY OF QUALITY AND CONCERNS

Note: The remarks in this section apply to cruises 2002-14 and 2002-08 which occurred sequentially using the same equipment; the data were processed together.

There is a lot of fine-scale noise in the secondary conductivity channel.

Spikes occur which are associated with minima in the descent rate, even where those minima are not so low that shed wakes are present. For cast 2002-14-0108 there was severe spiking in the temperature data that was not associated with shed wakes; the problems were seen in all the pumped sensors, but were worst in the secondary sensors. The fluorescence was also affected, but not the transmissivity indicating that the problem is associated with the pumping system. It does not show up in the oxygen data, but the time response was so slow that this is not surprising. The problems were most acute for casts 50 to 108 for 2002-14, but the suggestion that the flow-rate was not steady remained for the rest of this cruise and for 2002-08 which followed. Metre-averaging will remove many of the problems since the noise is often bi-polar in nature. However, small-scale features in the unaveraged data set will not be as reliable as usual.

The descent rate was kept quite high for most casts, but due to the problems noted above more data than usual was lost in editing.

There are gaps in the SPAR and PAR data at depths of very low light level due to the effect of running WILDEDIT.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read.

Salinity, dissolved oxygen and extracted chlorophyll calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The SPAR calibration was changed after cast #47 (16583.74 to 1658.374) and that value was used throughout 2002-08 as well. It is unsure which is correct. Conversion was done using 1658.374. A check should be done later in processing and recalibration done if necessary.

The fluorometer cable was changed from 10X to 3X after cast #73 and back to 10X after cast #121.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 0585A.con for use on casts #1 - 73 and 122 - 135 and 0585B.con for casts 75 - 121. The only difference between the two is the gain for the fluorometer.

There was great difficulty in converting cast #74. FIXDAT did not help. The number of bytes in the header and con files did not match. When “Surface PAR voltage added” was removed from the CON file and the first scan skipped in conversion reasonable results were obtained. There is a note in the log saying “No Spar” but it is unknown whether this means it was not mounted, or if it was a case of no data being logged. The name of the con file used for cast #74 is 0585C.con.

A preliminary check shows all expected channels present.

The temperature and salinity traces track reasonably well and up and downcasts are similar.

##### STRIP

The salinity channels and bottle number channel were stripped from the CNV files so that DERIVE doesn’t create a 2nd set of salinity channels.

##### ALIGNCTD

The deck unit is one of the newer versions that advance both the primary and secondary conductivity channels by 0.073s. Tests done on 3 casts using ALIGNCTD showed optimal results with a choice of from +0.013s to +0.023s for both channels. Both channels were aligned using +0.018s for a net advancement of 0.091s. This alignment is notably different from previous cruises.

At a later stage of processing doubts arose about the alignment. For casts #50 – 53 there were bi-polar spikes in primary salinity that were associated with minima in the descent rate. Those minima (~0.5m/s) were not low enough to cause major problems with shed wakes. Looking at the secondary channel the temperature might be consistent with very minor corruption by shed wakes, but the salinity, while spiky, does not display the usual shed wake signature. The bi-polar spikes are diagnostic of flow-rate problems in well-aligned data. So the question arises – is the data well aligned? It is possible that something changed in the set-up that affected the alignment, but there is no evidence that occurred. Tests were made using cast #51 and 5 widely different alignments to see if the spiking could be removed. Only slight changes resulted and supported the choices made previously. So the problems appear to be with the flow-rate rather than the alignment.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in all channels except scan number.  Parameters used were:   	Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

It was discovered at the end of processing that this routine left gaps in PAR and SPAR at depths of low light level. In future these channels will not be included in the WILDEDIT routine.

WILDEDIT removed most of the spikes from cast #6. The rest can be removed with an editor later.

##### CELLTM

Tests were run on casts #8,10,36,92 and 129 to find the optimal parameter choice for CELLTM. Runs using (0.0245,9.5), (0.2,9), (0.3,9), (0.3,7) and (0.3,7) were used for (alpha, 1/beta). The overall best choice was found to be (0.3,9) but the differences were not large; CELLTM was run on all casts using (0.3,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. As noted during previous uses of this equipment the secondary conductivity and salinity are much noisier than the primary. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 25 | 1000 | +.0012 | +.00014 | ~0 v.noisy | High, mostly ok |
| 27 | 1000 | +.0011 | +.00010 | ~0 v.noisy | High, extremely noisy |
| 48 | 1000 | +.0008 | +.00008 | ~0 v.noisy | High, noisy |
| 97 | 1000 | +.0008 | +.00010 | ~0.0005 | High, very noisy |

The temperature differences are similar to those observed during 2002-01, -03 and –04 but the conductivity and salinity differences have changed signs.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. Cast #108 has extremely noisy temperature data. The primary temperature is good for the downcast only and the secondary is bad throughout the cast.

##### Checking Headers

A header summary and a header check were produced. The instrument summary did not list the sensors for cast #27. There was an extra comment line in the header, but changing this only made things worse, so the instrument summary was copied from the previous cast and pasted into file 2002-14-0027.IOS since there was no change of equipment. There was a note in the log that the SPAR did not work right so there may have been some corruption of the file.

Errors in stations names in the headers of casts #19 and 32 were found and corrected.

The average surface pressure is 2.2db

The cruise track was plotted and looks reasonable.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 15 points) was deleted.

Sample interval =  .04 seconds.

COMMENTS ON WARNINGS:

There were no warnings in the DELETE log.

Comparing the maximum pressure before and after DELETE shows that very little data was lost from the bottom of casts. (See delete_study.xls.)

All DEL files were copied to EDT files.

##### 13. COMPARE

The bottle files from cruises 2002-14 and 2002-08 were combined for this analysis since one followed the other immediately and used the same equipment.

The BOT files were examined for errors and none were found.

COMPARE was run. There are only 6 bottles below 200db for 2002-14 and only 8 for 2002-08. Using only those bottles the CTD primary and secondary salinities are low by 0.0007 and 0.0016 units, respectively; there is some evidence of pressure dependence and time-dependence. The differences are consistent with those found in section 9. (See 2214comp.xls)

- A second run of COMPARE was done using the DOX files (titrated oxygen values). The comparison indicates a fit of Titrated oxygen = 1.1254 * SBE Oxygen – 0.1769. The slope is similar to those found in early 2001 and during 2002-03.

##### DETAILED EDITING

The primary sensors were chosen for further processing because they showed less variation with time and there was much less noise in the salinity than is seen in the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C or where there were bi-polar spikes that are probably due to flow-rate problems.

For casts #50 – 53 and 108 such bi-polar spikes were common and there are areas of generally poor salinity. Where such bad salinity occurs for only a few records it was smoothed; where more extensive the bad salinity records were removed.

For cast #50 the salinity was very bad from 540-560db but the temperature looked ok so only the salinity was deleted.

For cast #108 there was severe spiking in the upcasts and in the secondary downcast; the primary has many spikes but was usable.

After cast #108 there were no large temperature spikes, but for most casts there are fine-scale unstable variations that are unlikely to be real – again associated with minima in the descent rate. The flow-rate appears to be unsteady. The potential errors are not large since most are bi-polar. For many, but it is unclear how to edit. Metre-averaging seems likely to do a better job than interpolation using CTDEDIT so they were left alone. Small-scale features in the unaveraged data set will not be as reliable as usual.

The following casts were edited at the surface or bottom only: 8,17,69

The following casts were edited more extensively: 1,10-13,21-65,70-135.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

BIN AVERAGE with an average interval of 0.25m was applied to the output of the filter. The ¼m-averaged files were stored in a separate directory for the use of Angelica Peña.

The dark value was about 0.06g/l for the 10X cable and about 0.07 - 0.08g/l for the 3X cable. The maximum values were at the top of the scale.

The cable was changed from 10X to 3X and back again during the cruise.

A preliminary comparison with chlorophyll values was made to determine if the instrument was well-calibrated. Problems were found in the data and it will not be archived at this time. Further study will be made and it may be added later.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The primary sensors were used in January and February of 2002 (2002-03, -04 and -01) when the primary salinity was high by 0.0038units, low by 0.0009 and high by 0.0011 respectively. The secondary sensors have not been used since the last calibration; previous to that the conductivity sensor was associated with some variations with time and pressure.

Historic ranges – A number of casts have values outside the historic ranges in the upper 200m, but deep values are as expected and the excursions from the ranges are believed to be real, since they have been noted in several cruises using different equipment. The investigators were advised of the unusual values.

##### Recalibration

The SBE 43 oxgyen data will not be archived at this time due to poor performance. The calibration is reasonable when checked using bottles, but the time response when the CTD is in motion is unacceptable.

The fluorescence will not be archived at this time; it will be studied and may be added at a later date.

The salinity will not be recalibrated since the differences from bottles are fairly small and based on few deep samples; for the two previous uses of this equipment the salinity was not recalibrated.

The Surface PAR had a small error in the calibration; the data was corrected by multiplying by 1.00016.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Dissolved:SBE, Chlorophyll:Fluorescence:Seapoint and Flag.

The channels were reordered and data format corrected as needed.

The Standard Check routine was run and problems found were corrected using EDIT HEADERS. The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

1-73 and 122-end – 10X cable on fluorometer.

74-121 – 3X cable on fluorometer.

50,51 – two-sided salinity spikes

108 – very noisy signal for temperature:primary during upcast and temperature:secondary up and down. Conductivity was not spiky but was offset for part of the upcast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-14

Dates: Start: 29 April 2002 End: 8 May 2002

Location: WCVI/JdeFuca

Vessel: John P.Tully

Party Chief: MACKAS,D

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2002-14

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 11/09/01 | Factory |  |  |
| Conductivity | 1729 | 17/07/01 | “ |  |  |
| Secondary Temp. | 2968 | 18/09/01 | “ |  |  |
| Secondary Cond. | 2173 | 11/09/01 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| PAR | 4601 | 09/02/01 |  |  |  |
| SPAR | 16504 | 25/05/01 |  |  |  |
| Oxygen SBE43 | 0047 | 11/26/01 | Factory |  |  |
| Transmissometer | 498DR | 01/08/01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |

April 15, 2007 – Germaine Gatien

The fluorescence was originally stripped from these files (and from those of some other cruises in 2002) because the comparison with titrated chlorophyll was considered questionable. It was noted at the time that this should be revisited when we knew more about data from this instrument. It is now believed that this data is suitable for archiving, but as always it should be considered nominal.

The fluorescence in the edited files (EDT) were filtered using a median filter, size 11, then the files were put through BIN AVERAGE as described in the original processing. REMOVE was run as before except that Fluorescence was NOT removed.

HEAD EDIT was run to fix formats and channel.