## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2002-19

Agency: OSAP

Location: West Coast Vancouver Island

Party Chief: Mark Saunders

Platform: RICKER

Date: August 4, 2002 – August 14, 2002

Processed by: Germaine Gatien

Date of Processing: 28 April 2003 – 1 May 2003

Number of original CTD casts: 27

Number of casts processed: 26 (1 test cast deleted)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer 197 and Fluorometer #2228. The deck unit S/N was 0471.

# SUMMARY OF QUALITY AND CONCERNS

This CTD system is rarely calibrated well.

There was no signal from the fluorometer.

There was no calibration sampling for the salinity and both the primary and secondary conductivity had spiky sections mostly near the surface. The salinity should be considered ±0.01 from 0 to 25m and ±0.005 below that. For cast #12 the salinity should be considered ±0.01 for the whole cast.

There were clearly problems with the plumbing of the CTD. The primary downcast conductivity data is bad in the top 50m of all casts probably due to a blocked bleed valve.

The upcast secondary conductivity is poor for many upcasts and occasionally the upcast primary temperature also looked wrong. The upcast transmissivity looks quite different from the downcast data.

There was no CTD daily log available; notes about some of the casts were available. Positions and times could not be confirmed for casts 21-35.

The times in the log differ from those in the headers by 1 hour. The headers are presumed correct.

The station names for casts 21-35 were wrong; because the correct information was unavailable the station name entry was left blank.

If the full CTD daily log and/or the Bridge log become available, header entries should be checked and fixed where necessary.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

No Log Book was available but there were notes about 11 of the casts.

There were no salinity calibration samples.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. An error in the date of transmissometer calibration was corrected and the resulting file named 2002-19-ctd.con.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2002-19-ctd.con.

A preliminary check showed no signal in the fluorescence channel. Tests were run adapting the configuration file in case the fluorometer was actually mounted on a different voltage line, but no signal was found in any position. The data was reconverted without the fluorescence channel.

##### The secondary temperature and conductivity are extremely noisy. The transmissivity upcasts are very different from the downcasts. The downcast primary conductivity is odd near the surface.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel.  Parameters used were:

Pass 1    Std Dev = 2;

Pass 2    Std Dev = 5;

Points per block = 50

##### CELLTM

Tests were run on casts # 6 and 27 to find the optimal parameter choice for CELLTM. Runs using  (0.01,7), (0.02,7), (0.02,9), (0.03,7) and (0.03,9) were used for (alpha, 1/beta). The overall best choice was found to be (0.03,9) for both conductivity channels; CELLTM was run on all casts using (0.03,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The secondary conductivity and temperature are noisier than the primary resulting in poor secondary salinity. The downcast data agree reasonably well, but the upcasts are poor. The differences between the pairs of sensors is much higher during the upcasts than downcasts and varies in sign from one depth to another. In previous use of these sensors there were similar problems.

The differences between sensors were extremely noisy and the values that follow are very rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 10 | 375 | ~0 | -0.0035 | -0.004 | Noisy, ~0.8m/s |
| 12 | 900 | ~0 | -0.003 | -0.004 | Noisy, ~0.9m/s |
| 23 | 300 | ~0 | -0.0045 | -0.005 | Very Noisy, ~0.9m/s |
| 24 | 375 | ~0 | -0.0065 | -0.008 | Very Noisy, ~0.75m/s |

All the differences were extremely noisy.

The transmissivity looks very strange with up and downcasts bearing little resemblance to each other.

The upcast conductivity is also far from the downcast at times; this suggests there was a problem with the package orientation during upcasts. It is hoped that the downcast transmissivity is reliable. The maximum values of in deep water are close 89%.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was run to add event numbers based on the last 4 digits of the file names.

##### Checking Headers

A header summary and a header check were produced. The times in the headers are all one hour later then those in the log notes. This is probably due to assuming Standard Time rather than Daylight Savings Time. Since it is not known which are correct the headers will not be changed at this time. If the bridge logs are made available this can be checked later.  Positions and times were available for the first 11 CTD casts only.

The station name for cast 8 was corrected from C2 to C3 as noted in the log notes. The station names for casts 20 to the end were all B6. Cast #20 was confirmed from the log notes, but the others are presumed to be wrong and so were removed leaving that entry blank in the headers. If the information becomes available the station names should be added later.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 2.8db which is reasonable.

All plots were examined on-screen and it was determined that the primary salinity was bad for all casts in the upper 50m. There are many problems with upcast temperature, conductivity and transmissivity.

CLEAN was run a second time to replace pad values in pressure with interpolated values.

##### SHIFT

Tests were run on casts #10 and 23 using SHIFT to align conductivity data to match the temperature with respect to pressure. Sections of data in which T and C have high gradients were examined for salinity spiking. Runs were made using advancements of 1.5 to 1.9 records. The results were examined in T-S space; the best results minimize unstable spiking but don’t over-smooth. The best choice was to advance the secondary by 1.6 records (equivalent to +0.067s). The primary conductivity was advanced by the deck unit by 0.073s; the poor downcast data make it difficult to fine-tune this so it will not be adjusted.

All data was put through SHIFT to advance the secondary conductivity by +1.6 records.

Cast #12 was put through REVERSE since the downcast salinity was bad for both channels.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0 		Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted from 10db to 10db above the maximum.pressure.

Sample interval =  .042 seconds. (taken from header)

COMMENTS ON WARNINGS:

The only warnings in the DELETE log were for upcast records.

All DEL files were copied to EDT files

##### COMPARE

There were no calibration samples for this cruise.

For 2002-35 which immediately followed this cruise the primary salinity was found to be close to the bottles and the secondary was low by 0.007 units.

For this cruise the difference between the pairs of sensors varies from near zero to +0.008 at depth with the primary higher. There is weak evidence that the calibrations were drifting with time so the differences might be a little lower than for 2002-35. Choosing a figure of +0.004 for recalibration of the secondary salinity is probably a reasonable guess, but expected errors would be on the order of ±0.005

##### DETAILED EDITING

The secondary sensors were chosen for further processing because the primary salinity was bad in the top 50m for all casts. Cast #12 had bad salinity for both downcast channels and the secondary upcast. For that cast only, upcast data and primary sensors were selected.

Page plots were produced using T1,S1. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C.

All casts required some editing except for cast #18.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### BIN AVERAGE

The following Bin Average values were applied to the EDT files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – During 2002-05 in February the deep bottles indicated that both pair of sensors were producing salinities that were high by 0.003units; the shallow bottles suggested that the primary salinity is low by 0.004 and the secondary by 0.002units. During 2002-35 in August 2002 the deep samples were considered unreliable because the data was noisy. There were some shallow samples from well-mixed waters that indicated the primary and secondary salinity were low by 0.0003 and 0.007 respectively.

Historic ranges – Salinity at about 25m was slightly low for one station and near the minimum for a few others. Temperatures around 50db were low for some stations and for the one deep cast the temperature was a little high around 700db. This looks like other data collected in spring/early summer 2002 so is not considered indicative of trouble with sensors.

##### Recalibration

File 2219rcal.ccf  was used to add 0.004units to the secondary salinity.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE

The following channels were removed from all casts except #12:  Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Pump:Status, Descent Rate and Flag.

The following channels were removed from cast #12: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Pump:Status, Descent Rate and Flag.

HEADER EDIT was used to fix data formats and to add a warning to the comment section of the headers advising investigators of quality concerns particularly in regards to the salinity and transmissivity data.

The final files were named CTD.

The Standard Check routine was run and no problems found.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

0 – Very shallow test cast. Deleted.

12 – Both downcast conductivity channels bad, upcast secondary poor, upcast primary ok.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-19

Dates: Start: August 4, 2002 End: August 14, 2002

Location: North-West Pacific

Vessel: RICKER

Party Chief: Mark Saunders

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | no | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2002-19

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 17/12/01 | Factory |  |  |
| Conductivity | 2278 | 13/12/01 | “ |  |  |
| Secondary Temp. | 2095 | 17/12/01 | “ |  |  |
| Secondary Cond. | 2280 | 13/12/01 | “ |  |  |
| Transmissometer | 197 | 21/01/01 | IOS |  |  |
| Fluorometer | 2228 |  |  |  |  |
| Pressure Sensor | 69698 | 12/11/97 | Factory |  |  |