## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 15-Dec-2013 | Added units to the nutrient data. J.L. |
| 10 Oct. 2008 | Nutrient data added to CHE files….G.Gatien |
| 21 Dec. 2004 | Salinity recalibration based on post-cruise calibration. See note below |

## PROCESSING NOTES

Cruise: 2002-26

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Location: Strait of Georgia, Juan de Fuca Strait

Project: SoG/JdeF

Party Chief: Masson, D.

Platform: John P. Tully

Date:  15 July 2002 – 19 July 2002

Processed by: Germaine Gatien

Date of Processing: 8 December 2002 – 6 January 2003

Number of original CTD casts: 84

Number of casts processed: 84

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 333DR and Seapoint Fluorometer S/N #2226 with a 10X cable. The deck unit was S/N 0425.

# SUMMARY OF QUALITY AND CONCERNS

The data was in generally good order. There are few deep bottles from this cruise. Also there were difficulties with the salinity sampling from all other cruises using this equipment since the sensors were last calibrated. Hence the salinity must be considered to be ±0.003.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The transmissivity co-ordinates do not agree with those on file at IOS. Upon investigation there were two sets of calibration files in existence and the constants used at sea are correct.

The fluorometer was logged as a WETLAB but was actually a SeaPoint fluorometer.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2002-26-0550.con; this was created by editing the con file used at sea to correct the type of fluorometer to SeaPoint.

A preliminary check shows all expected channels present.

The temperature and conductivity traces track reasonably well, and up and downcasts are as similar as can be expected in this region. There is a lot of noise in some casts.

##### ALIGNCTD

The deck unit was one of the older ones that aligns only the primary conductivity channels by +0.073s.

ALIGNCTD was used to align the secondary channel by +0.073s as well. The alignment of both channels will be fine-tuned using SHIFT after conversion to IOS HEADERS.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and transmissivity channels only.  Parameters used were: 	Pass 1    Std Dev = 2 Pass 2    Std Dev = 5    Points per block = 50

##### CELLTM

Tests were run on casts #52 and 119 to find the optimal parameter choice for CELLTM. Runs using  (0.02,7), (0.02,9), (0.03,7) and (0.03,9) were used for (alpha, 1/beta). The best choice varied from feature to feature but overall (0.02,7) had the best effect for both the primary and the secondary sensors. CELLTM was run on all casts using (0.02,7) for the primary and secondary sensors.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

There were no deep casts. Three of the deepest (#4, 72 and 83) were plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The descent rates were typically around 0.8m/s and fairly steady. The differences between channels were very noisy with averages of near zero for temperature, +0.0003 for conductivity and 0.003 for salinity. These are higher than the differences during 2002-30 and 2002-32 when the same equipment was used, but these casts are much shallower than the ones used for the tests in the other cruises.

The dark values for the fluorometer varied from 0.15 to 0.165 μg/l. However, cast #4 has fairly high values at all depths.

Maximum values for transmissivity were about 88% (over 25cm).

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files and the extensions changed to BOT.

##### Checking Headers

A header summary and a header check were produced and two errors in station names were fixed. A track plot was produced and looks fine.

The average surface pressure is 1.8db.

##### Test Plots

All casts were plotted and checked for evidence of processing and/or instrumental problems. Cast #41 has bad conductivity and temperature on the upcast.

##### SHIFT

The IOS SHELL routine SHIFT was used to fine-tune the alignment. Because the temperature gradient was never very high there were no notable spikes to test the alignment. During 2002-32 when the same equipment was used the best results were found using –0.4 records for both the primary and secondary channels. A test was run on a section of cast #6 using shifts between –0.2 and –0.6 records for both channels. As expected there was not much change but using –0.4 records appears to improve some of the data very slightly without creating significant problems elsewhere.

All casts were put through two runs of SHIFT applying –0.4 records to the primary and secondary conductivity channels.

##### DELETE

All casts were put through CLEAN to replace flags in the pressure channel by interpolation.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted. This feature was not applied in the top 10m or the bottom 10m.

Sample interval =  .042 seconds.

COMMENTS ON WARNINGS: There were no warnings.

The DEL files were copied to EDT files.

##### COMPARE

The BOT files were examined for errors; none were found.

COMPARE was run. The results are highly dependent on what data is included in the fit. At first sight the data appears to be highly pressure dependent, but since there is no deep data available this is probably not significant. During 2002-30 and 2002-32 no pressure dependence was obvious. If outliers are removed until a flat result is achieved the primary salinity is found to be high by 0.0035 and the secondary by 0.0063. However, this leaves only 10 points in the fit. If we use all data from 200db down the average differences are smaller but pressure-dependent. See 2232com1.xls for details.

##### DETAILED EDITING

The primary sensors were chosen for further processing since they are closer to the bottles and were used in other recent cruises using those sensors.

Page plots were produced using T0,S0.

These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C or due to other instrumental spiking. The spikes were largely left unedited if it appeared that metre-averaging would do as good a job. Otherwise an attempt was made to smooth salinity.

Note was made of the editing details in the relevant files.

All files required some editing except for casts: #3,31,36,43,56,58,69,70.

##### SPECIAL FLUOROMETER PROCESSING

A special set of edited files with data from the top 100 db was created using the IOSSHELL routine CLIP. These are for the use of Angelica Peña and will not go into the archives.

For the data to be archived a median filter, size 11, was applied to the fluorometer data to reduce spikiness.

There was no chlorophyll sampling from bottles during this cruise.

No attempt will be made to calibrate the fluorometer at this time.

##### BIN AVERAGE

The following Bin Average values were applied to the edited files:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The data from 3 cruises that used these sensors have been processed since the latest calibration: 2002-30, 2002-32 and 2002-34. There were serious problems with the calibrations during all three. ARGO float data and surface sampling did support the conclusion that the primary sensor was high by 0.0028 and the secondary by 0.0040 during 2002-30. During 2002-32 they were high by 0.0048 and 0.0073 respectively and during 2002-34 by 0.0026 and 0.0056 respectively.

Historic ranges – Temperatures were below and salinity above the historic minima for casts 4, 26, 27, 28 and 30; these casts were south of San Juan Island and Victoria. For most the anomalous data was quite shallow, but for cast #4 it was deep. The temperatures below 100 db were near the historic minima for casts 11, 15 and 16 in the middle of Juan de Fuca Strait. There have been many observations of low temperature water off the coast of North America from spring/summer 2002 and so these results are probably not indicative of instrument calibration problems. Higher than normal salinity values have also been found on the continental shelf this year.

##### Recalibration

It is clear that the CTD salinity is too high, but there is considerable doubt about the size of the error. It is unfortunate that the only other experience with these sensors is from other doubtful comparisons. Of some comfort is the fact that the results for 2002-26 are intermediate between those of the other cruises.

File 2226rcal.ccf was created to recalibrate the primary salinity by subtracting 0.0035 from the primary and 0.0063 from the secondary. This was applied to the SAM, AVG, CLIP and BOT files

COMPARE was rerun to check that the recalibration was done correctly; the average difference after recalibration is +0.00004 for both primary and secondary salinity. (See 2230com2.xls.)

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary and Flag.

The channels were reordered and formats corrected as needed.

The Standard Check routine was run and no problems found.

The final files were named CTD and RAC.

The same steps were taken with the bottle files and clipped CTD files and stored separately for Angelica Peña.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

39. Log Note: First CTD lifted out of water. May have hit bottom.

41. Bad secondary signals on upcast.

50. Log Note: Hit jellyfish; transmissivity very noisy from 240 to bottom.

72. Log Note: Stop during downcast

82. Log Note: Hit jellyfish. Transmissivity very noisy from 280 to bottom.

83. Log Note: Hit jellyfish.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-26

Dates: Start: 15 July 2002 End: 19 July 2002

Location: Georgia Strait / Juan de Fuca Strait

Vessel: John P.TULLY

Party Chief: Masson, D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2002-26

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2371 | 16/04/02 | Factory |  |  |
| Conductivity | 2399 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2663 | 18/04/02 | “ |  |  |
| Secondary Cond. | 2424 | 16/04/02 | “ |  |  |
| Fluorometer –pumped | 2226 |  | IOS |  |  |
| Transmissometer | 333DR | 07/10/02 | IOS |  |  |
| Pressure Sensor | 75636 | 14/06/99 | Factory |  |  |

Dec. 22, 2004: Based on a post-cruise calibration the CTD and CHE files were recalibrated, using file 2002-26-recal2.ccf, to raise the primary salinity by 0.0053psu. It was assumed that the drift was linear with time. G. Gatien