# Linda White 							November 21, 2002

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Diane Masson

OSAP

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Coastal Cruises 2002 – 26					      	       	                                 Nitrate+Nitrite, Silicate and Phosphate

Summary

Frozen seawater samples were thawed and analysed on a Technicon Autoanalyser according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

A medium check standard was analysed periodically throughout the day’s analysis.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning, middle and end of the analytical run.

All labware and nutrient tubes have been acid cleaned, soon to be capped and stowed.

Pump platen tubing has been replaced, rinsed with 3N NaOH and DMQ water.

## Observations

Analysis went very well.

## No replicates taken at sampling.

## Quality Assurance data

## Standards

## Coefficient of Variation %

# Calibrated Standards

Medium check		N03 23.96µmol/L	Si 60.02µmol/L		P04 2.00µmol/L

Analysed		23.9+/-0.1		60.2+/-0.8		1.98+/-0.02

n = 16			n = 14			n = 16

C.V.%			 0.51%			1.6%			0.50%

This report includes a final data spreadsheet, quality assurance results, peakbase calculated .anz files, and plots.

Excel Filename: CoastalDM200226.xls.