# Linda White 							Oct. 22, 2002

7036 Bickstan Place

Brentwood Bay, B.C.

V8M 1G6

## Diane Masson

OSAP

Institute of Ocean Sciences

9860 West Saanich Road

Sidney, B.C.

V8L 4B2

Summary Report: 	Nutrient analysis – Coastal Frozen nutrients 2001- 31						Nitrate+Nitrite, Silicate and Phosphate

Summary

Frozen seawater samples were thawed and analysed on a Technicon Autoanalyser according to procedures stated in “Institute of Ocean Sciences Nutrient Methods and Analysis” by J.Barwell-Clarke and F. Whitney 1996.

A medium check standard was analysed as a sample following each station.

Standard envelopes of 0, Low, Medium and High nutrient concentrations were analysed at the beginning, middle and end of the analytical run.

## Observations

Nutrients were analysed from one plastic tube.  Samples were in excellent condition and labelled well.  Analysis went very well.

## Quality Assurance data

## Standards

No corrections were made to the data because the calibrations remained within 1% of Wako standard concentrations for nitrate + nitrite and silicate.

### Coefficient of Variation % for standards analysed as a sample and Sp for sample replicates

Medium check 		N03 20.13µm/L		Si 60.2 µm/L		P04 2.00µm/L

(Theoretical)

Analysed		20.2+/-0.17		60.7+/-0.3		1.97+/-0.02

n = 15			n = 15			n = 15

C.V.%			0.82%			0.43%			1.26%

Pooled Standard deviation for sample replicates, Sp; 	NO3+NO2 	is 0.68,	n = 14

SiO4		is 0.98,	n = 12

PO4		is 0.07,	n = 14

This report includes a final data spreadsheet data, quality assurance results, peakbase calculated .anz files, and plots.

Excel Filename: Coastal200231.xls.