## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| Dec. 22, 2004 | Recalibrated salinity channel; see note below. |
| May 6, 2003 | Added nutrient and chlorophyll data to the archive. |

## PROCESSING NOTES

Cruise: 2002-32

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Location: WCVI

Project: Nestucca

Party Chief: Yelland, D.

Platform: John P. Tully

Date:  16 September 2002 – 24 September 2002

Processed by: Germaine Gatien

Date of Processing: 12 November 2002 – 9 December 2002

Number of original CTD casts: 59 (including 4 casts with only surface data)

Number of casts processed: 55

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 498DR and Seapoint Fluorometer S/N #2229 with a 10X cable. The deck unit was S/N 0508.

# SUMMARY OF QUALITY AND CONCERNS

The descent rate was at times extremely noisy and sometimes fairly low; the combination meant a lot of lost data due to shed wakes.

The bottle comparison showed a lot of scatter, so the calibration should be considered ±0.003.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The calibrations date and parameters given for the transmissometer do not agree with those on file at IOS. The same thing was noted for cruise 2002-30 when it was determined that the con file was probably correct. The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2002-32-0000.con.

A preliminary check shows all expected channels present; PAR was not used for all casts – maybe none. This will be checked later.

The temperature and conductivity traces track reasonably well and up and downcasts are similar.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Tests were run on casts #52 and 119 to find the optimal parameter choice for CELLTM. Runs using (0.01,7), (0.02,6), (0.02,7), (0.02,9), (0.03,7) and (0.03,9) were used for (alpha, 1/beta). The best choice was clearly (0.02,7) for the primary sensors but for the secondary all values led to worse results. CELLTM was run on all casts using (0.02,7) for the primary and (0.0,0) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three deep casts (#3, 58 and 65) were plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. There is a PAR channel for every cast, but it was not always mounted; files will be checked later and the channel removed where there is no data.

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 3 | 2000 | ~0.0002 | ~0.00024 | ~0.0027 | Extremely noisy, high |
| 58 | 2000 | ~0.0001 | ~0.00025 | ~0.0027 | Extremely noisy, high |
| 65 | 1750 | ~0.0003 | ~0.00025 | ~0.0027 | Noisy, high |

These differences were about twice as high as during the previous cruise which used the same equipment. This may indicate some sensor calibration drift.

The dark values for the fluorometer were on the order of 0.13μg/l.

Transmissivity at depth is about 89% with some variability with pressure.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files and the extensions changed to BOT.

##### Checking Headers

A header summary and a header check were produced and some errors found.

There were errors in positions and times due to GPS problems.

The positions were wrong for casts #14, 19 & 20. These were fixed using the log positions.

The times were local for cast #3. Many other times are also a little different from the times in the log. During the previous cruise there was apparently some drift in the computer time. Cast #3 was changed to UTC and the other times were not changed since it is not clear which time is right and the differences were small.

The station name for cast #52 was changed to agree with the log.

After all the corrections were made the header summary and header check were rerun and track plots produced.

The same corrections had to be made to the bottle files as well.

The average surface pressure is 1.9db. Mixed layer depth was calculated using a reference pressure of 7db and maximum variation in the primary salinity of 0.002units. This was done to identify casts that are very well-mixed in the upper 20db and hence suitable for calibration of the salinity. However, there were no surface salinity bottles available.

##### Test Plots

All casts were plotted and checked for evidence of processing and/or instrumental problems. The primary conductivity data is odd for cast 19.

After conversion transmissivity below 500db has values of about 63.5% and shows no evidence of significant pressure effects.

##### SHIFT

The deck unit was one of the newer ones that aligns both primary and secondary conductivity channels by +0.073s. The IOS SHELL routine SHIFT was used to fine-tune the alignment.

During 2002-30 when the same equipment was used the best results were found using –0.3 records for the primary and –0.7 records for the secondary channel. So an initial test was run on a section of cast #39 using shifts between –0.2 and –0.9 records for both channels. The results were examined in T-S space to see how well spikes in salinity were removed. The best choice varied from feature to feature, but the overall best results were with –0.4 records for both primary and secondary.

A section of cast #75 was also tested and similar results found.

All casts were put through two runs of SHIFT applying –0.4 records to the primary and secondary conductivity channels.

##### DELETE

All casts were put through CLEAN to replace flags in the pressure channel by interpolation.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 10   	Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted.

Sample interval =  .042 seconds.

COMMENTS ON WARNINGS: There were no warnings.

A comparison was made of the maximum pressures before and after running DELETE to ensure that a large amount of data was not lost at the bottom as sometimes occurs due to misinterpretation of pressure spikes. There was no more than 1db lost at any station. (See 2232delete_study.csv in the processing section of the full report.)

##### COMPARE

The BOT files were examined for errors; there is a lot of noise in the salinity channel at the surface for cast #24. CTDEDIT was used to edit or remove bad primary salinity from bottles #5, 6 and 7. The output file ED1 was renamed as BOT.

Cast #17 has a surface bottle followed by the usual upcast bottles.

COMPARE was run.  After rejecting points with differences of more than 0.1 salinity units there were only 14 points left. Two of the bottle samples from cast #74 appear to be reversed; when these entries were exchanged the differences are similar to many others. There was a larger than usual scatter in the differences. They were reasonably flat vs pressure and there is no clear temporal change. However, the inclusion or rejection of just one point makes a big difference in these conclusions. The primary salinity was found to be high by 0.0048 and the secondary by 0.0073units. See 2232com1.xls for details.

Unfortunately there was no surface sampling. In well-mixed surface waters such sampling is very useful.

##### DETAILED EDITING

Both primary and secondary salinity channels were noisy; since the primary was closer to the bottles it  was the primary sensors were chosen for further processing.

Page plots were produced using T0,S0. Since the primary salinity looked bad from 30-50db it was decided to use the secondary channels for that cast only.

These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. The spikes were largely left unedited if it appeared that metre-averaging would do as good a job. Otherwise an attempt was made to smooth salinity.

The descent rate was frequently extremely noisy, and at times the average descent rate was low so that  shed wake effects led to a heavy loss of data; DELETE removed a lot and more was removed in editing.

All casts required editing.

Note was made of the editing details in the relevant files.

##### SPECIAL FLUOROMETER PROCESSING

A set of edited files with data from the top 100db was created using the IOSSHELL routine CLIP.

For the data to go into the archives a median filter, size 11, was applied to the fluorometer data to reduce spikiness.

The results were compared with the extracted chlorophyll from Melanie Quenneville. The chlorophyll values were mostly less than 2μg/l and the scatter was large. A trend line forced through the origin had a slope of 1.4 with the fluorescence values higher,or 1.6 when one outlier was omitted. The scatter was large.

When the ratio of FLUOR to CHL was plotted against hour of day in PST a pattern emerges. The ratio is higher at night. The average ratio is 2 at night and 1.5 during daylight.

No attempt will be made to calibrate the fluorometer at this time.

##### BIN AVERAGE

The following Bin Average values were applied to the edited files:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The only data that has been processed from these sensors since the latest calibration was from 2002-30. There were serious problems with the calibrations during that cruise. ARGO float data and surface sampling did support the conclusion that the primary sensor was high by 0.0028 and the secondary by 0.0040.

Historic ranges – As has been observed for much of the data collected off the west coast of Vancouver Island in 2002, the data frequently fell outside the historic ranges. A number of casts on or near the continental slope had low temperatures – at about 100m depth in the north and 50m in the south. The salinity was low for cast #56 around 150m. Salinities were unusally high for casts 47, 48 and 65; the first two were very close to shore and similar observations were made off Oregon and Washington this year. The temperatures were near the maximum in the 400-1000m range for a number of the deeper casts. Given the unusual conditions of 2002 and given that some casts show higher temperature and salinity than expected and others lower, these observations cannot be considered evidence of poor calibration.

##### Recalibration

It is clear that the CTD is reading high, but there is considerable doubt about the size of the error. It is unfortunate that the only other experience with these sensors is from another doubtful comparison.

File 2232rcal.ccf was created to recalibrate the primary salinity by subtracting 0.0048 from the primary and 0.0073 from the secondary. This correction should be considered to be ±0.003.

COMPARE was rerun to check that the recalibration was done correctly; the average difference after recalibration is -0.00003 for the primary and 0.0003 for the secondary. (See 2230com2.xls.)

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE, REORDER and EDIT HEADERS

The following channels were removed from all CTD casts except #19: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, PAR, Surface PAR and Flag.

The following channels were removed from cast #19: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, PAR, Surface PAR and Flag.

The channels were reordered and formats corrected as needed.

The Standard Check routine was run and problems fixed.

The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

- Surface data only. Deleted.

- Surface data only. Deleted.

- Descent rate exceptionally noisy, many reversals.

4. Surface data only, just before cast #7. Deleted.

14a. Surface data only. Deleted.

14-19. No NMEA in CTD file.

19. Primary conductivity bad from 30 to 50m; use secondary sensors.

24. Surface bottles, primary salinity bad.

67. May have hit bottom; CTD/R rinsed after this cast.

104. Hit bottom.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-32

Dates: Start: 16 September 2002 End: 24 September 2002

Location: NE PACIFIC

Vessel: John P.TULLY

Party Chief: Yelland, D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2002-32

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2371 | 16/04/02 | Factory |  |  |
| Conductivity | 2399 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2663 | 18/04/02 | “ |  |  |
| Secondary Cond. | 2424 | 16/04/02 | “ |  |  |
| Fluorometer –pumped | 2229 |  | IOS |  |  |
| Transmissometer | 498DR | 08/23/02 | IOS |  |  |
| Pressure Sensor | 75636 | 14/06/99 | Factory |  |  |

Dec. 22, 2004: Based on post-cruise recalibration of conductivity sensors, the primary salinity channel was recalibrated using file 2002-32-recal2.ccf to add 0.0078psu to the primary salinity. It was assumed that the drift was linear with time. G. Gatien