## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| Dec. 22, 2004 | Recalibrated salinity based on post-cruise calibration; see note below. |
| May 6, 2003 | Added nutrient data to rosette cast 29. |

## PROCESSING NOTES

Cruise: 2002-34

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Location: WCVI, LaPerouse, Juan de Fuca

Project: Covariability, LaPerouse, Effingham, Juan de Fuca

Party Chief: Tom Juhász

Platform: John P. Tully

Date:  24 September 2002 – 30 September 2002

Processed by: Germaine Gatien

Date of Processing: 16 December 2002 – 2 January 2003

Number of original CTD casts: 36

Number of casts processed: 36

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 498DR, PAR sensor #4601 and Seapoint Fluorometer S/N #2229 with a 10X cable. The deck unit was S/N 0508. There was a surface PAR (#16504) listed in the configuration files, but only null data was logged.

# SUMMARY OF QUALITY AND CONCERNS

One data file had been corrupted and required help from SeaBird for conversion. This appears to have been the result of opening the data file using an inappropriate editor or word processing program.

The bottle comparison showed a lot of scatter, so the calibration should be considered ±0.002.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The calibrations date and parameters given for the transmissometer do not agree with those on file at IOS. The same thing was noted for cruise 2002-30 and 2002-32 when it was determined that the con file was probably correct. The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2002-34-0001.con.

A preliminary check shows all expected channels present; PAR was not used for all casts.

The temperature and conductivity traces track reasonably well and up and downcasts are similar.

Cast #11 contained garbage when converted. Running FIXDAT produced the error message: “unable to determine scan length in data file”. Attempts were made to change the header file and con file in ways that have helped in the past with unconvertible files, but nothing helped. This file was sent to SEABIRD where John Backas concluded that someone had attempted to edit the data file. Each occurrence of LF was converted to <CR><LF>. He wrote a new program FIXWP that corrects this problem. The word processor had also deleted random bytes (or perhaps the cast had data transmission problems) and FIXDAT V 1.3 did a bad job of fixing the remaining problems, so he made some improvements, creating FIXDAT V 1.4.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Tests were run on casts #4, 20 and 29 to find the optimal parameter choice for CELLTM. Runs using (0.02,7), (0.02,9), (0.03,7) and (0.03,9) were used for (alpha, 1/beta). The best choice was clearly (0.02,7) for the primary and secondary sensors. CELLTM was run on all casts using (0.02,7) for both primary and secondary conductivity.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Three deep casts (#22, 24 and 26) were plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. There is a PAR channel for every cast, but it there is only data for casts 29 to 41. There is a channel for surface PAR but no data present. The data was reconverted without surface PAR.

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 22 | 1700 | Xnoisy~0 | ~0.00025 | ~0.003 | Extremely noisy, high |
| 24 | 1700 | ~0.0002 | ~0.00025 | ~0.003 | Extremely noisy, high |
| 26 | 1200 | ~0.0002 | ~0.00028 | ~0.003 | Noisy, high |

These differences are similar to those found during the previous cruise that used the same equipment.

The dark values for the fluorometer were on the order of 0.13 μg/l.

Transmissivity at depth is about 89%. While there is some variability with pressure it is not systematic and seems to reflect changes in other variables such as temperature.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files and the extensions changed to BOT.

There were some missing event numbers so both IOS and BOT files were put through CLEAN to create event numbers from the last 4 digits of the file name. The IOS files were also put through the routine to replace PAD values and to fix the limits in the headers.

##### Checking Headers

A header summary, a header-check and track plots were produced and no errors found.

The average surface pressure was 0.88 db which is shallower than usual. Checking individual cast files it appears (judging by conductivity values) that the CTD was recording while at the surface, bobbing in and out of the water for some casts. The pressure appears to be correct.

##### Test Plots

All casts were plotted and checked for evidence of processing and/or instrumental problems. Cast #41 has bad data in the downcast to 60 db.

After conversion transmissivity below 500 db has values of about 63%.

##### SHIFT

The deck unit was one of the newer ones that aligns both primary and secondary conductivity channels by +0.073s. The IOS SHELL routine SHIFT was used to fine-tune the alignment.

During 2002-32 when the same equipment was used the best results were found using –0.4 records for both primary and secondary channels. So an initial test was run on a section of cast #29 using shifts between –0.2 and –0.6 records for the primary. The results were examined in T-S space to see how well spikes in salinity were removed. The test confirms that –0.4 records is the best choice for this cruise as well.

All casts were put through two runs of SHIFT applying –0.4 records to the primary and secondary conductivity channels.

##### DELETE

Because of problems with the downcast for cast #41 it was put through reverse. The output was named 2002-34-9941 and put through DELETE. It can be decided later what data to archive for cast 41.

The following DELETE parameters were used for casts 1-12 which were in shallow protected waters:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Low drop rate feature turned off

Sample interval =  .042 seconds

The following DELETE parameters were used for casts 14-41:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Salinity: 20   	Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted.

Sample interval =  .042 seconds.

COMMENTS ON WARNINGS: There were no warnings.

##### COMPARE

The CLN files (created from BOT files) were examined for errors; there was noise in the secondary salinity channel for bottle #11 of cast #29. CTDEDIT was used to clean this data. The output file ED1 was renamed as CLN.

COMPARE was run.  Including points from below 100db and rejecting one outlier, the primary salinity was found to be high by 0.0026 and the secondary by 0.0056. See 2234comp.xls for details. There was a lot of scatter but no obvious pressure or time-dependence.

##### DETAILED EDITING

The primary salinity was much closer to the bottles so it was chosen for further processing.

Page plots were produced using T0, S0.

These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. The spikes were largely left unedited if it appeared that metre-averaging would do as good a job. Otherwise an attempt was made to smooth salinity.

The upcast data was selected for cast #41.

All casts required editing except for cast #8.

Note was made of the editing details in the relevant files.

##### SPECIAL FLUOROMETER PROCESSING

A set of edited files with data from the top 100 db was created using the IOSSHELL routine CLIP.

For the data to go into the archives a median filter, size 11, was applied to the fluorometer data to reduce spikiness.

The results were compared with the extracted chlorophyll from Melanie Quenneville. The average ratio of CHL to FL is 0.97 but the standard deviation is 0.49. A trend line forced through the origin gave CHL = 0.93 * FL. If one point (from 5 db) is rejected the slope rises to 1.06. The ratios for the two bottles from 50 db are very low, 0.14 and 0.21, but the CHL values are extremely low so this is probably not significant.

There were few daylight samples for this cruise, but there does appear to be a trend to lower CHL/FL values during daylight. This was also noted during 2002-32.

No attempt will be made to calibrate the fluorometer at this time.

##### BIN AVERAGE

The following Bin Average values were applied to the edited files:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – The only data that has been processed from these sensors since the latest calibration was from 2002-30 and 2002-32. There were serious problems with the calibrations during 2002-30. ARGO float data and surface sampling did support the conclusion that the primary sensor was high by 0.0028 and the secondary by 0.0040. During 2002-32 there was also a lot of scatter; the primary salinity was found to be high by 0.0048 and the secondary by 0.0073.

Historic ranges – As has been noted for other observations off the west coast of Vancouver Island in 2002, the data frequently fell outside the historic ranges. A number of casts on or near the continental slope had low temperatures at depths of from 35m to 100m, varying from cast to cast. The salinities were within the historic ranges offshore, but in Juan de Fuca salinities were unusually high at depths of about 50m for JFE6 and JFE7 and from 100m down at casts JFD4 and JFD3. High salinities have been observed on the continental shelf off Oregon and Washington this year. Given the unusual conditions of 2002, these observations cannot be considered evidence of poor calibration.

##### Recalibration

Results from 2002-30, -32 and –34 all indicate that the CTD salinity is high. The results for this cruise are quite close to those of 2002-30. File 2234rcal.ccf was created to recalibrate the primary salinity by subtracting 0.0026 from the primary and 0.0056 from the secondary. Given the scatter in the comparison for this cruise and all the others that used this equipment the salinity calibration should be considered ±0.002.

COMPARE was rerun to check that the recalibration was done correctly; the average difference after recalibration is -0.0002 for the primary and 0.0002 for the secondary. (See 2234com2.xls.)

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE, REORDER and EDIT HEADERS

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary and Flag..

PAR was removed from casts #1 to 27.

The channels were reordered and formats corrected as needed.

The Standard Check routine was run and problems fixed.

The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

41. Bad data during downcast from surface to 60db; upcast used for archive.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-34

Dates: Start: 24 September 2002 End: 30 September 2002

Location: NE PACIFIC

Vessel: John P.TULLY

Party Chief: Juhász T.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2002-34

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2371 | 16/04/02 | Factory |  |  |
| Conductivity | 2399 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2663 | 18/04/02 | “ |  |  |
| Secondary Cond. | 2424 | 16/04/02 | “ |  |  |
| Fluorometer –pumped | 2229 |  | IOS |  |  |
| PAR | 4601 | 17/22/02 |  |  |  |
| Surface PAR | 16504 | 07/22/02 |  |  |  |
| Transmissometer | 498DR | 08/23/02 | IOS |  |  |
| Pressure Sensor | 75636 | 14/06/99 | Factory |  |  |

Dec. 22, 2004: Based on post-cruise calibration of conductivity sensor the primary salinity was recalibrated using file 2002-34-recal2.ccf to add 0.0058psu. It was assumed that the drift was linear with time. G. Gatien