## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 12-Aug-2003 | Subtracted 0.0027 from the CTD primary temperature for all the CTD files due to sensor drift suspected by Germaine and confirmed by Seabird. See file “Report on sensors 2023-1763.doc” in the DOC directory. |
| 18-Nov-2003 | Added times and locations to bottle salinity files. |

## PROCESSING NOTES

Cruise: 2002-40

Agency: OSAP

Location: Knight Inlet

Project: Zooplankton Aggregation at Sills

Party Chief: Dave Mackas

Platform: Vector

Date: 13 November 2002 – 25 November 2002

Processed by: Germaine Gatien

Date of Processing: 23 April 2003 – 15 May 2003

Number of original CTD casts: 20 (of which one was run with CTD on deck), 3 missing

Number of casts processed: 18

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (#333DR) and a Seapoint Fluorometer (#2356). The fluorometer had a 10X cable and was mounted on the secondary pump. The deck unit was #0619.

# SUMMARY OF QUALITY AND CONCERNS

The data files for 3 casts could not be found; two of these were calibration casts.

One cast contained bad data due to the presence of a syringe on the pump.

The data was otherwise in good order.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read. There are three files mentioned in the log for which there are no data files: casts 67, 68 and 73.

Salinity calibration data was obtained. The cruise summary sheet was completed.

The data entry in the configuration files was correct except for the serial number of the pressure calibration. File 0585ctd.con was prepared with the corrected serial number.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 0585ctd.con; there were some Niskin bottles but no rosette was used.

The year in all files was changed from 1999 to 2002 using a text editor on the cnv files.

The Daily Log mentioned errors in the headers of casts 24 and 64; these were corrected.

Cast #64 contains no useful data. The CTD was run on deck. This cast will not be processed further.

A preliminary check of a few casts shows all expected channels present.

The temperature and conductivity traces track reasonably well during downcasts and up and downcasts are similar. There is some spikiness in both conductivity channels. The transmissivity looks reasonable but spiky. The Fluorescence has dark values of about 0.08ml/l but goes higher below 200m for some casts. No con file had a PAR sensor listed but no cast was found with PAR data, nor was it mentioned in the Daily Log.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in pressure, conductivity and transmissivity due to spikiness noted in many casts.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### CELLTM

Tests were run on casts #11, 30 and 33 to find the optimal parameter choice for CELLTM. Runs using (0.01,7), (0.02,7), (0.03,7), (0.02,9), (0.03,9) and (0.0245,9.5) were used for (alpha, 1/beta). There was very little difference between any of the results for casts 11 and 33, probably because there was such a small temperature range. For cast #30 with a slightly larger range there were very small differences; (0.02,7) was the best choice overall for the primary conductivity channel. It was impossible to distinguish between a few of the choices for the secondary conductivity because it was so noisy, but (0.02,7) certainly did improve the data. CELLTM was run on all casts for both channels using (0.02,7).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The differences in conductivity and salinity were extremely noisy and the values that follow are very rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 12 | 432 | -0.003 | ~-0.00005 | ~+0.0025 | 0.75, very steady |
| 40 | 517 | -0.003** | ~0 | ~+0.0025 | 0.75, very steady |

** There appears to be some depth dependence with the differences increasing with depth.

The differences are very close to those observed during 2002-39 which preceded this cruise and used the same sensors.
The descent rate was generally about 0.75 and very steady.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to replace pad values in pressure with interpolated values.

All casts were plotted and checked for evidence of problems with the processing or instruments. The secondary conductivity is very noisy resulting in noise in the secondary salinity on the order of 0.004psu at depth.

Cast #24 looks very odd. The two pairs of sensors are very different and up and downcasts are different. The fluorescence looks bad with large values at depth, but the transmissivity looks fine. There is a note in the log that the syringe was left on the secondary system, so it is presumed that the primary data is ok.

##### Checking Headers

A header summary and a header check were produced and checked. No errors were found.

The average surface pressure is 0.6 db.

The cruise track was plotted and looks reasonable.

##### 10. COMPARE

Salinity calibration bottles were taken using Niskin bottles about 2m above the CTD. There are only 4 samples available since two of the calibration cast files cannot be found. The differences found are as follows:

| Cast # | Pressure | Salinity:Bottle | CTD S0 | CTD S1 | S0 - Bottle | S1-Bottle |
| --- | --- | --- | --- | --- | --- | --- |
| 12 | 429 | 31.5498 | 31.570 | 31.573 | 0.020 | 0.023 |
| 25 | 180 | 32.0323 | 32.031 | 32.033 | -0.001 | 0.001 |
| 34 | 373 | 31.5114 | 31.510 | 31.510 | -0.001 | -0.001 |
| 40 | 516 | 31.5219 | 31.516 | 31.518 | -0.006 | -0.004 |

There are doubts about how the bottles were fired. If the depths recorded for firing are correct then some were fired while the CTD was still in motion, so the contents of the bottles did not have time to equilibrate and would likely contain water from farther above the CTD than expected. If they were actually fired while the CTD was at rest at the bottom of the cast, then the CTD salinity would be higher than that recorded in the table. These data are considered too unreliable to be used for recalibration. Cast #12 is particularly problematic since the local variations are large at the depth of firing. (See 2240comp.xls)

##### SHIFT

###### Conductivity

The routine SHIFT was used to align the conductivity.  The deck unit was one of the newer versions that advance both channels by 0.073s. Tests were run on casts #11 and 33 using shifts between –0.5 records and +0.3 records. After shifting, the data was put through DELETE then displayed in T-S space to look for the choice that “just” removes unstable salinity spikes without oversmoothing. There was not a lot of change but the best results were with shifts of –0.4 for both conductivity channels.

###### Fluorescence

A few casts were checked and the offset of upcast and downcast fluorescence compared with that for temperature. The fluorescence is separated by from 2 to 2.5m more than the temperature. Dividing that figure by 2 (since the downcast will be shifted in the opposite direction to the upcast) this amounts to about a 1m shift which is equivalent to about +24 records. A +24 record shift was applied to all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted between 10db and 10db above the maximum pressure.

Sample interval =  .0412 seconds.

COMMENTS ON WARNINGS: There were none.

All DEL files were copied to EDT files.

##### DETAILED EDITING

The primary sensors were chosen for further processing because they were less noisy than the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Salinity was cleaned where large spikes occurred. Small spikes were cleaned only if it was clear they were due to imperfect alignment of T and C.

All casts required light editing only, including the deletion of records at the surface and/or bottom and/or light editing of salinity.

Note was made of the editing details in the relevant files.

Cast #29 has an unstable feature in both pairs of sensors in the downcast but not in the upcast. Given the active mixing in this area and no obvious instrumental problem it was left unedited.

CLIP was used to create files with the top 100db only for the use of Angelica Peña.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen. Cast #29 has a significant unstable feature. This appears in both pairs of sensors in the downcast but not in the upcast. Given the active mixing in this area and no obvious instrumental problem it was not edited any further.

##### Other comparisons

Previous experience with these sensors – The sensors have been used during four other cruises in the summer/autumn of 2002 (2002-16, 2002-28, 2002-31,2002-39). The primary salinity was found to be high by 0.0015, low by 0.0011, high by 0.0008 and high by 0.0004. The secondary salinity was high by 0.005, 0.0019, 0.0037 and 0.003. The second values are probably the most reliable coming from a cruise with the most bottles. The primary was selected for processing in all three cruises. The secondary pump was replaced before 2002-39 so the secondary history may not be useful.

Historic ranges – There was no useful climatology available for this area.

Comparison of nearby casts – Cast #24 looked odd so was compared to #25 which was nearby and to #21 and #26 which were on either side of #24 and #25. The secondary salinity for #24 was about 5psu lower than the other three casts at all depths. This is understandable since the syringe had been left on the secondary pump. But the primary salinity was also low at all depths by about 1psu. So, either the syringe impeded the flow to the primary pump as well, or there is some connection in the tubing. In any case the data is bad and this cast will be deleted.

Casts #30 through 40 were examined and the variations look reasonable.

##### Recalibration

The salinity was not recalibrated since the differences from bottles are small and scattered. Salinity should be considered ±0.002.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE

The following channels were removed from all CTD casts: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Irradiance (PAR), Bottle_number, Status:Pump and Flag.

EDIT HEADERS was used to fix the data format of CTD files and to add a warning to the headers of the CTD files as follows:

Transmissivity: The data are unedited.

Fluorescence:URU:Seapoint – The data are nominal and unedited.

The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

24. Bottom depth changed from 159 to 251 as noted in log book. Syringe left on secondary pump. Salinity bad in both channels. DELETED.

64. File name in header changed from 99 to 64 as noted in log book.

64. Running CTD on deck. DELETED.

67, 68, 73. Data files missing.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-40

Dates: Start: 13 November 2002 End: 25 November 2002

Location: Knight Inlet

Vessel: Vector

Party Chief: Mackas, D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2002-40

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 03/05/02 | Factory |  |  |
| Conductivity | 1763 | 23/04/02 | “ |  |  |
| Secondary Temp. | 2106 | 19/04/02 | “ |  |  |
| Secondary Cond. | 1764 | 23/04/02 | “ |  |  |
| Fluorometer –pumped | 2356 | 07/01 | IOS |  |  |
| PAR | 4615 | 15/12/00 | Factory |  |  |
| Transmissometer | 333DR | 11/10/02 | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/1999 | Factory |  |  |