## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 3-Mar-2014 | Added Mission number to the rosette files. J.L. |
| May 24, 2003 | File 200-41_nut.csv is the edited version of QF2002-41nuts.xls. The .xls file had several errors in cast and sample numbers. New file was placed in the DOC directory. Bernard Minkley. |

## PROCESSING NOTES

Cruise: 2002-41

Agency: OSAP

Location: SoG/JdeF

Project: SoG/JdeF

Party Chief: Cummins, P.

Platform: Vector

Date: 16 December 2002 – 20 December 2002

Processed by: Germaine Gatien

Date of Processing: 24 March 2003 – 7 April 2002

Number of original CTD casts: 71

Number of casts processed: 71

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Transmissometer (498DR at first and  333DR later), SBE 43 Dissolved Oxygen Sensor (S/N #0047), PAR sensor #4615 and Seapoint Fluorometer (S/N #2229 for the first part of the cruise and #2356 later). The DO sensor was mounted on the primary pump and the fluorometer on the secondary. The fluorometer had a 10X. The deck unit S/N was #0619.

# SUMMARY OF QUALITY AND CONCERNS

The data was in good order. There was some confusion over configuration files but the log contained details on equipment changes.

The dissolved oxygen sensor continues to show poor time response. Attempts have been made to correct for some of this error using titrated samples, but the errors are considered to be on the order of:

- ±0.4 ml/l  from 0 to 125m

- ±0.2 ml/l  from 125 to 200m

- ±0.1 ml/l  below 200m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained and read.

Salinity and dissolved oxygen calibration data was obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. A number of errors were found and corrected.

The sensor history was found.

##### Conversion of Raw Data

Due to changes in equipment 3 configuration files were used as follows:

0550ctda.con – 	casts   1 - 36

0550ctdb.con – 	casts 37 - 50

0550ctdc.con – 	casts 51 - 76

The raw data was converted using the above files; rosette files were converted separately with salinity channels included.

A preliminary check of a few casts shows all expected channels present.

The temperature and conductivity traces track reasonably well and up and downcasts are similar although upcasts are much noisier than downcasts.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Cruise 2002-36 which immediately preceded this one used the same equipment. The best choice for CELLTM for 2002-36 was (alpha, 1/beta) = (0.03,9) for both conductivity sensors. Tests were run on casts #1, 47 and 76 and (0.03,9) proved a good choice for #1 and #76. No choice improved #47, but a close look at the data showed major differences between up and downcasts, so this is not a good test. CELLTM was run on all casts for both channels using (0.03,9).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity and descent rate.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The differences were extremely noisy and the values that follow are very rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 26 | 325 | -0.0006 | -0.0003 | -0.003 | ~0.75 steady |
| 42 | 325 | -0.0005 | -0.0003 | -0.003 | ~0.75 very noisy, often <0.5 |
| 75 | 325 | -0.0005 | -0.0003 | -0.0025 | ~1 steady |

The differences are very noisy and slightly higher than those observed during 2002-36, which is to be expected due to the shallowness of the comparisons and the active mixing in this area.

The dark values of fluorescence are about 0.085ug/l.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The ROS files were converted to IOS files, and the extensions were changed to BOT.

CLEAN was used to add event numbers to both IOS and ROS files. The ROS files were named BOT after clean.

All casts were plotted and checked for evidence of problems with the processing or instruments. Cast #58 had bad data at the end of the file. A text editor was used to remove the bad records and the file was put through CLEAN again. Some bad records remain at the beginning but these will be removed by DELETE.

##### Checking Headers

A header summary and a header check were produced and checked. Station names were fixed in casts #6 and #10. There was no surface salinity sampling so the mixed-layer depth was not investigated.

The cruise track was plotted and looks reasonable.

##### 10. COMPARE

The BOT files were examined for errors and none were found.

SALINITY

COMPARE was run. Given the active mixing in this area and the relatively shallow sampling it is not surprising that the comparison contains a huge amount of scatter. When only samples from below 250db are included the primary sensors are found to be high by about 0.0025 and the secondary low by 0.0003. There is no convincing evidence of pressure or time dependence. (See 2241comp.xls)

DISSOLVED OXYGEN (2002-36 and 2002-41 were considered together for this step)

- A three-step approach was taken to determining the correction of the dissolved oxygen channel. First, COMPARE was run using the OXY files (titrated oxygen values) and the SBE_DOX channel from the CTD bottle files. This corrects for the errors in calibration of the sensor. It is assumed that by the time the bottle is fired the sensor has come to equilibrium, but in fact in areas of high gradient of DOX this is far from true, but is the best we can do. As has been found for other data sets the best fit is differences versus dissolved oxygen rather than versus pressure or time. The results are as follows:

- Titrated oxygen = 1.1054*(SBE_DOX) – 0.0565

- For 2002-41 a similar result was found:

- Titrated oxygen = 1.1170*(SBE_DOX) – 0.1051

- Because there is so much scatter it seemed wise to combine the two data sets and the results were then:

- Titrated oxygen = 1.1090*(SBE_DOX) – 0.0704

- Attempts were made to achieve a tighter fit by excluding outliers but the effects were small and hard to justify. As noted with other data sets there is some apparent dependence on pressure, but that is likely a secondary effect since the DOX values are varying with pressure. Similarly a slight hint of time-dependence completely disappears when the two cruises are examined together. Each of the cruises moved through different water types muddying the issue. (See 2241oxy.xls and 36-41combo_DOX.xls.)

- The second step in the analysis involved shifting the oxygen channel to improve the alignment relative to pressure. There is clearly a large delay in the oxygen – this step will attempt to align features in downcast and upcast data so the vertical offset is the same as for temperature. This is a fairly gross estimate since again it depends on descent rate and local gradients; moreover, the upcast temperature for this cruise was particularly noisy and the upcast contained many stops. Nonetheless, since the position of features is one of the more important uses of this data it seems important to try to get the features to a realistic pressure. For 2002-31 a value of +220 records was found after extensive analysis. For 2002-36 estimates of +170, +220 and 250 records were made based on different features. The higher values are in higher gradient zones. There were bottle stops for all casts complicating the issue. From 2002-41 two casts with steady descent rate and no bottle stops suggest that +220 records is the best choice.

- After running SHIFT and DELETE a set of recalibrated, averaged and calibrated downcast files will be compared with the titrated values. At that point a fit of differences vs pressure may be chosen to correct some of the time-response errors that remain after SHIFT.

##### SHIFT

Conductivity

The routine SHIFT was used to align the conductivity.  The deck unit was one of the newer versions that advance both channels by 0.073s. Tests were run on casts #25 and 50 using shifts between –0.6 records and +0.2 records. After shifting, the data was put through DELETE then displayed in T-S space to look for the choice that “just” removes unstable salinity spikes without oversmoothing. The best results were with shifts of –0.3 records for both primary and secondary conductivity. For 2002-36 the best results were with -0.2 records; there may have been slight changes to plumbing. Both channels will be shifted by -0.3 records.

Dissolved Oxygen

Based on the analysis in the previous section the dissolved oxygen sensor was advanced by +220 records.

Fluorescence

- The offset between upcast and downcast fluorescence was compared to that of temperature in the same way as was done for dissolved oxygen to determine if alignment is needed. Noisy upcast temperature traces and low fluorescence values made this difficult and estimates of 0.5s to 3s lags were made. For 2002-36 using the same equipment the estimate was even more difficult due to bottle stops on the upcasts for all casts; lags of 0s, 1s, 1.5s, 1.5s, 1.7s and 3s were found. During 2002-38 using similar equipment a lag of 1s (24 records) was found. The fluorescence will be advanced by 24 records.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0

Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted between 10db and 10db above the maximum pressure.

Sample interval =  .0412 seconds.

COMMENTS ON WARNINGS:

The only warnings in the DELETE log referred to upcast or bottom records.

All DEL files were copied to EDT files.

##### FINAL OXYGEN COMPARISON

The downcast data was averaged in 1m bins and thinned to depths corresponding to bottles. The DOX channel was recalibrated using the results of section 10 (2241rcal1.ccf) which corrects for the errors in the sensor itself. There are two sources of error in what remains - time response errors not fixed by SHIFT and real differences between up and downcast. Since we can not distinguish between the two we just hope the latter error averages out over many casts. A second run of COMPARE was done with the titrated values and the thinned, recalibrated downcast data (2241oxy2.xls). The differences were fairly flat with pressure so a simple offset was applied to the thinned files using calibration control file 2236rcal2.ccf. COMPARE was run again and the average difference between bottles and downcast CTD was 0.0049ml/l (2241oxy3.xls).

##### DETAILED EDITING

The secondary sensors were chosen for further processing because the salinity was less noisy and closer to bottles than the primary salinity; the same result was found for 2002-36. However, for casts #36, #58, #70 and #71 the primary sensors were chosen because of bad secondary data in the top 20db.

Page plots were produced. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Salinity was cleaned where large spikes occurred. Small spikes were cleaned only if it was clear they were due to imperfect alignment of T & C.

All casts except #16, 27 and 34 required some editing; for most, surface and/or bottom records were removed and salinity was cleaned very lightly to remove spikes.

The following casts required heavy editing due to many shed wakes: 2-8, 43, 44, 48.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

The dark value was about 0.085g/l.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

The same values were used for the BOT files except that the Bin Channel = Bottle Number.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors –The sensors have been used only one other time since the last recalibration. The comparison for 2002-36 was poor due to a huge scatter; the primary salinity was found to be high by about 0.0022 units and the secondary low by 0.0003. When only very deep data was used the differences were +0.0033 and +0.0009.

Historic ranges – The data from the deeper casts in Juan de Fuca Strait fall outside the ranges below 100db with lower salinity and higher temperatures. The bottles support the CTD salinity values, so it is assumed that these are real differences from the norm, or that the ranges are very limited at depth. All other data falls within the ranges which are generally very broad. There are no ranges available for the northern Strait of Georgia.

##### Recalibration

Recalibration was done in two steps using 2241rcal1.ccf and 2241rcal2.ccf. The first step is identical to that used for cruise 2002-36 and corrects the SBE dissolved oxygen for sensor calibration errors. The second does a further correction to the dissolved oxygen to correct for time-response problems and corrects the primary salinity by applying an offset of -0.0025.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE and REORDER

The following channels were removed from all CTD casts except #36, 58, 70 and 71: Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Bottle_Number, Descent_Rate and Flag.

The following channels were removed from casts #36, 58, 70 and 71: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Bottle_Number Descent_Rate and Flag.

The same channels were removed from the rosette files except that Bottle_Number was not removed.

The Fluorescence:URU:Seapoint channel was removed from cast #41.

The data format was corrected  and comments added to the headers using EDIT HEADERS.

The final files were named CTD and RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars (including notes from CTD Daily Log)

19. Fluorometer reading 0.

20. Hit bottom

- Transmissometer reading low

36. Secondary conductivity bad in top 20db. Use primary sensors.

51. The fluorescence channel was removed. All values to 80m are zero, spikes around 80m.

58. Serious spiking. Secondary conductivity bad in top 20db. Use primary sensors.

70. Secondary conductivity bad in top 20db. Use primary sensors.

71. Secondary conductivity bad in top 20db. Use primary sensors.

62. Hit bottom

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2002-41

Dates: Start: 16 December 2002 End: 20 December 2002

Location: SoG/JdeF

Vessel: Vector

Party Chief: Cummins, P.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0550

Cruise ID#:		2002-41

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2968 | 9/07/21 | Factory |  |  |
| Conductivity | 2173 | 9/07/21 | “ |  |  |
| Secondary Temp. | 2038 | 18/09/01 | “ |  |  |
| Secondary Cond. | 1729 | 9/07/21 | “ |  |  |
| Fluorometer –pumped | 2229 | 9/07/21 | IOS |  |  |
| Fluorometer –pumped | 2356 | ? | IOS |  |  |
| PAR | 4615 | 15/12/00 |  |  |  |
| Oxygen SBE43 | 0047 | 21/08/02 | Factory |  |  |
| Transmissometer | 498DR | 05/12/02 | IOS |  |  |
| Transmissometer | 333DR | ? | IOS |  |  |
| Pressure Sensor | 75636 | 6/04/1999 | Factory |  |  |