## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 8-Dec-03 | Added nutrient data to rosette files. Noticed that oxygen data for cast 66 was missing so I merged it as well (oxygen file had been assigned cast # 65). |

## PROCESSING NOTES

Cruise: 2003-07

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG

Party Chief: Cummins P.

Platform: Vector

Date: April 23, 2003 – April 28, 2003

Processed by: Germaine Gatien

Date of Processing: 22 August 2003 – 29 August 2003

Number of original CTD casts: 78 (including one file with only a little surface data and one aborted cast)

Number of casts processed: 74

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with Transmissometer 333DR, Seapoint Fluorometer (#2229) and SeaBird Model SBE43 Dissolved Oxygen sensor (#47). The deck unit was a SeaBird model 11 (#0619).

# SUMMARY OF QUALITY AND CONCERNS

One file had a non-standard name. It is recommended that a normal consecutive number be assigned even for repeat casts at a single site.

It is believed that the primary conductivity calibration changed during this mission, probably due to a cracked cell. The secondary conductivity is very noisy and there are some doubts about the calibration which showed a lot of scatter. The secondary salinity was chosen for the archive and is considered ±0.003.

The dissolved oxygen sensor continues to show poor time response but the calibration against bottles during bottle stops looked better than usual. Attempts were made to correct for time response problems but errors remain and are estimated to be of the order:

- ±0.5 ml/l in the top 100 (but most points are within 0.3 ml/l)

- ±0.3 ml/l from 100m to 200m

- ±0.15 ml/l from 200m to 400m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Salinity and oxygen calibration data was obtained. The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. An error in the serial number for the pressure sensor was corrected and the resulting file named 2003-07-CTD0585.con.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration file 2003-07-ctd0585.con. Rosette files were converted.

Cast 2003-07-0000.cnv contains only surface records and will not be processed further.

Cast 2003-07-0070.cnv was aborted and rerun as 2003-07-0071. No reason is given in the log but the data looks bad. The repeat cast looks ok. The first cast will not be processed further.

A few casts were checked to ensure that all expected channels contained reasonable data. The only problem noted is that the CTD had an initial partial downcast during cast 1 followed by a full cast.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.

Parameters used were: 	Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

For 2003-08 which preceded this cruise CELLTM was run using (0.01,7) was used for (alpha, 1/beta).

For 2003-12 and 2003-15 which followed this cruise (0.02,7) was used for (alpha, 1/beta).

For this cruise fewer external sensors were mounted on the CTD; checks were made to see what parameters would work best. Casts #26, 59 and 76 were checked using (0.01, 7),(0.01,9), (0.02,9), (0.03,9), (0.02,7) and (0.03,7). There was very little difference in the results for the primary sensors, and often doing nothing looked best. The secondary salinity was very noisy so no attempt was made to study it. CELLTM was not run on this data.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The secondary temperature is a little noisier than the primary. In other uses of this equipment it has been noted that the temperature sensors were farther apart than usual and that is the case for this cruise as well. In the report for 2003-12 it was noted that there was a large change in the salinity differences since 2003-08. This cruise occurred between those two missions and it looks like the change happened mid-cruise. The conductivity differences move from ones that look like those for 2003-08 to ones like 2003-12. The primary sensors were recently sent to SeaBird and they report that the conductivity cell was cracked. They also showed a large drift in temperature from the previous calibration.

The differences between sensors were generally noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 51 | 350 | -0.003 | -0.00015 | +0.0015 | Fairly high, steady |
| 59 | 350 | -0.003 | +0.0 | +0.0035 | High, steady |
| 75 | 350 | -0.003 | +0.00005 | +0.003 | High, steady |

Since there appears to have been a change to the conductivity mid-cruise an attempt was made to determine when it occurred. Salinity differences were plotted at many stations and the change appears to have occurred between casts #57 and 58. The CTD hit the bottom during cast #57.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

T0, T1, S0 and S1 were plotted for all casts. The primary and secondary are reasonably close and the upcasts look as much like the downcasts as expected for this region. The only problem noted is casts #1 and #25 have initial downcast sections. A text editor was used to remove records from the beginning of the IOS files for those casts to ensure that DELETE does not select the initial downcast data. A note was made in the headers about this editing.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

The rosette files were converted to IOS files, put through CLEAN to add event numbers and named BOT. All BOT files were plotted and no significant outliers were found for T0, T1, S0, S1 or DOX. The fluorescence was frequently very noisy even at depth so will be checked carefully later.

##### Checking Headers

A header summary and a header check were produced. These summaries were examined and no errors were found in the headers.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 1.4db.

##### COMPARE and BOTTLE FILE PREPARATION

The BOT files were examined for outliers in temperature and salinity and none were found.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Sample #220 was used for two different casts. The sample for cast # 60 was renamed #922 and that for cast #61 was left as #220. There was no rosette sheet for cast #60.

The ADDSAMP file was input to the ADD SAMPLE NUMBER routine to create SAM files from the BOT files. Those files were averaged on bottle number. (Output: SAMAVG)

The spreadsheet of salinity data was found and converted to SAL files using the IOS SHELL routine CONVERT SPREADSHEET.  A comment was added to one SAL file to explain the renaming of the sample number in cast #60.

The SAL and OXY files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

DISSOLVED OXYGEN

COMPARE was run. Points flagged by the analyst were excluded from the comparison. When data from above 10dbars were also excluded there was a fairly tight relationship with oxygen. There was no obvious drift with time but the scatter is large. Looking at differences versus pressure shows little change below 200db, something that was noted during 2003-15 but not previously.(See 2003-07-oxy-comp1.xls)

The best fit is against SBE dissolved oxygen so the following recalibration will be used:

Titrated oxygen = 1.1335*(SBE_DOX) – 0.0695

That is close to the results of 2003-08 and 2003-15 which bracketed this cruise:

- Titrated oxygen = 1.1520*(SBE_DOX) - 0.1390. (2003-08)

- Titrated oxygen = 1.1474*(SBE_DOX) - 0.1024. (2003-15)

- There were no severe outliers that were not already flagged by the analyst.

- After running SHIFT and DELETE a set of recalibrated, averaged and calibrated downcast files will be compared with the titrated values. At that point a fit of differences vs pressure may be chosen to correct the data further to remove time-response errors not corrected by SHIFT.

Salinity comparison

The primary sensors were sent to SeaBird in July 2003 and they reported that the conductivity cell was cracked. They report a drift of 0.004 PSU/year in the salinity, but they also report a drift in the primary temperature sensor of 0.00264 Cº/year. The previous recalibration took place just over a year before this cruise. The temperature error appears to have held steady since August of 2002, but the results of salinity comparisons suggests that most of the drift in salinity has occurred since March. See Report on sensors 2023-1763.doc for details on the calibration. From the differences in salinity described in section 7 it seems likely that the change occurred after cast #57 due to the CTD hitting bottom.

COMPARE was run. One value was excluded as an outlier since it was lower than both CTD salinities by about 0.52 units. That value was flagged and a note of explanation was added to the header in file 2003-07-0010.mrg. The differences are noisy as is usually the case for this region. The data was too noisy to confirm the change of calibration that apparently happened after cast #57. Averaged over all unflagged bottle data the primary salinity is low by 0.002 and the secondary is low by 0.0001 units but the standard deviation is 0.01 for both sensors.

A few casts were examined in detail to try to determine why there was such a large scatter in the comparison. There are a few cases where the stop was not a full 30s and where there was evidence that the shed wakes had not finished passing through. On the other hand the data is sampling waters with rapidly changing properties and/or the CTD was bobbing up and down which probably lead to real differences between what is in the Niskin bottles and what the CTD is sampling. Even allowing for these problems it seems strange that the salinity from the bottles was often higher than the CTD by an amount that does not fit observations for this equipment from other missions. It is a pity that there was no bottle cast in Saanich Inlet, but the mission was rushed and some regular stations dropped.

##### SHIFT

Conductivity

It is anticipated that the primary salinity will not be archived, but just in case it is wanted later SHIFT will be run on the primary conductivity using -0.5 records based on the results of 2003-08 and 2003-15 which both had similar equipment mounted.

There is no history on shifting the secondary since that channel has not been archived since the sensors were last calibrated and the conductivity is extremely noisy so that any improvements may be hard to detect. Tests were run on cast #63 cast using -1, -0.5, 0, +0.5 and +1 records to see if there is any noteworthy change. The best results were with +0.5 records although this was only slightly better than 0. The net advancement is +0.05s for the primary and +0.09s for the secondary since the deck unit advanced both by +0.073s.

All data was put through SHIFT using -0.5 records for the primary and +0.5 records for the secondary.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for 2 casts were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The range of values was from 1.4s to 2s. Since there tend to be more flow-rate problems on the upcast, a conservative choice of 1s was made; a shift of +24 records was applied. This is the same figure that has been applied to all recent data sets.

Oxygen

- The dissolved oxygen sensor continues to show time-response problems, but this data looks a little better than has been observed recently. During bottle stops the oxygen values reach equilibrium faster. In the past the setting that was chosen made up and downcasts look reasonably similar, but did not match the time taken to reach equilibrium during stops. This may mean that the flow rate to the sensor was better for this cruise and the delay is just that due to the response time of the sensor. Or perhaps the smaller range of values means that the sensor is better able to keep up.

To determine the best shift value to apply the data was studied in by applying a variety of shifts and seeing what shift makes the offset between downcast and upcast DOX equal that between downcast and upcast temperature. The optimal choice varies greatly from cast to cast and according to depth. Estimates of 150 to 220 records were made. For all previous cruises using this equipment a shift of +220 or +240 was deemed best. Using +220 records looks like a good choice for 2003-07. All casts were shifted by +220 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: None

##### DETAILED EDITING

The secondary sensors were chosen for editing. They have not been chosen in recent times because the conductivity signal is so noisy, resulting in noise at the level of ±0.002 units in the secondary salinity. However, that noise is two-sided and pretty random. A quick test of metre-averaged data suggests the errors from this noise are not significant, probably less than 0.001. It is clear that a change took place after cast #57 and it seems to affect the primary more than the secondary. It is possible that the calibration changed for both sensors, but for the secondary the change is probably small, less than 0.001 units. To ensure that is the case casts #56 and 57 were looked at for changes between downcast and upcast data. We expect that upcast data will have slightly higher T and S reflecting the entrainment of deeper water by the equipment and/or flow-rate problems on the upcast. But the differences between the pairs of sensors should not change. That is true for cast #56, but not for #57 when the upcast primary salinity appears to have dropped by at least 0.001 relative to the secondary. The differences between the secondary upcast and the secondary downcast look similar to cast #56. This sort of analysis is not terribly reliable given that the two casts were in Haro Strait where active mixing was taking place so that it is impossible to pick out portions of the two casts that are really alike. While some doubt remains about the secondary it should be good to ±0.003 allowing for noise level and calibration uncertainties.

Page plots were produced using T1,S1. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. In a few cases the descent rate was examined to determine if unstable features should be removed or not.

The following casts required no editing: 2, 12, 13, 18, 29, 71 and 73.

The following casts required fairly heavy editing: 3-8 and 44-45.

All other casts required only light editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

Casts #19 and 9941 were not edited since they were repeat casts done to test the effects of low descent rate and will not be archived.

##### RECALIBRATION

The SAM files were bin-averaged using Bottle Number as the bin channel. (Output: SAMAVG)

The EDT, MRG, SAM and SAMAVG files were recalibrated using file 2003-07-rcal1.ccf to recalibrate the dissolved oxygen using the equation given in section 10. COMPARE was rerun for oxygen (2003-07-oxy-comp2.xls) and the results were satisfactory. (Output: COR1; MRGCOR1, SAMCOR1, SAMAVGCOR1)

##### FINAL OXYGEN COMPARISON

A set of downcast files were created from data that had been shifted to align dissolved oxygen, edited to remove bad records, thinned to standard depths (the depths of bottles) and recalibrated using the first oxygen comparison. COMPARE was run comparing those files to the bottle channel in the MRG files. (See 2003-07-oxy-comp3.xls.) When the differences were examined compared to pressure or dissolved oxygen values, the results fall fairly close to a simple offset. This is probably due to the fact that the oxygen range is smaller than usual and the maximum pressure quite low. The CTD oxygen was high by an average of about 0.07 ml/l. The scatter is on the order of ±0.5 ml/l down to 100db, about ±0.3 ml/l from 100 to 200db and ±0.15 ml/l from 200 to 400db.

A final recalibration was applied to the SBE dissolved oxygen channel using 2003-07-rcal3.ccf (Corrected DOX = DOX - 0.0001*Pressure + 0.0856) based on the results of the fit of differences versus pressure. COMPARE was run again and the results were as expected. (See 2003-07-oxy-comp4.xls.) This final recalibration was applied to the CTD files only, not the bottle files since the lag is not an issue in the bottle files. (Output: COR3)

- At the end of the processing job it was discovered that one cast had been mislabeled and missed from the runs of COMPARE for oxygen, and a few other casts had not been included in the first comparison. A test run of COMPARE was done to see whether including those casts made a significant difference and the effects varied from 0.003 to 0.008ml/l depending on oxygen value. It was decided the effect was too small to justify rerunning the oxygen comparison. Since the downcast vs bottle comparison included all but one cast some of the error will be removed by the later calibration. The differences between the calibration based on the first run (cal 1) and the corrected run (cal 2) are as follows:

| DOX (ml/l) | 2003-07-rcal1.ccf | Calibration based on test run of COMPARE | Difference (ml/l) |
| --- | --- | --- | --- |
| 2 | 2.1975 | 2.2013 | -0.0038 |
| 3 | 3.3310 | 3.3354 | -0.0044 |
| 4 | 4.4645 | 4.4695 | -0.0050 |
| 5 | 5.5980 | 5.6036 | -0.0056 |
| 6 | 6.7315 | 6.7377 | -0.0062 |
| 7 | 7.8650 | 7.8718 | -0.0068 |
| 8 | 8.9985 | 9.0059 | -0.0074 |
| 9 | 10.1320 | 10.1400 | -0.0080 |

##### Special Fluorometer Processing

There was no chlorophyll sampling so no special bottle files were prepared.

The calibrated downcast files were clipped to 100db, unwanted channels removed, formats fixed as needed and header comments added. These files were saved on a CD-ROM for the use of Angelica Peña.

To prepare files for general use a median filter, fixed size=11, was applied to the fluorescence channel in the cor2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: BOX)

##### Study of Special casts testing effect of descent rate

Two repeat casts were run using different descent rates.

18/19 – Comparing these two casts with average descent rates 1m/s and 0.25m/s shows evidence of noisier data in the latter than in the former. The water is quite well mixed so the differences are small.

9941 and 42 – Comparing these two casts with average descent rates 0.5m/s and 1m/s there is clear evidence of a shed wake around 183db in cast 9941.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. As mentioned above the secondary has not been selected for the archive. During 2003-08 in March 2003 the salinity was found to be high by 0.003 and for 2003-12 in May it was low by 0.0013.

Historic ranges – There were no significant excursions from the historic ranges where local climatology was available.

Repeat casts – Cast 76 was at the same site as cast #46 but about 33 hours later. The two casts are very similar below 175db but quite different above that. The site is one at which rapid changes are expected so this is not indicative of instrumental problems.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying Temperature and Salinity profiles.

Profile plots were produced with Temperature, Dissolved Oxygen, Fluorescence and Transmissivity versus Pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Bottle_Number, Oxygen:Voltage, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment, fix formats and add a comment to the headers using file 2003-07-header.hdr. The final files were named CTD. The Standards Check routine was run and no problems were found.

##### Final Bottle Files

The SAMAVGCOR1 files were put through REMOVE routine to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage and Flag for all casts.

HEADEDIT was used to fix formats and correct errors in the headers; these files were named RAC.

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage and Flag.

HEADEDIT was used to add the standard comment about quality flag definitions and to fix formats. Standards check was run on all files and no problems were found. These files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

0	surface data only – not processed

6. 	weight hit bottom

18/19 	repeat station – test of descent speed

19. 	not archived since cast 18 is better

25. 	restarted cast – pumps not on first time

40	may have hit bottom

41a 	TEST cast to be compared with 42. Renamed 9941. Not archived.

42 	repeat station – test of bow wake

49	air vents on Niskin Bottles open on recovery

57	hit bottom

70	aborted cast – not processed

71	repeat of previous cast

76	test cast – bow wake study – full downcast/up 50m/down/full upcast

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-07

Dates: Start: April 23, 2003 End: April 28, 2003

Location: Strait of Georgia / Juan de Fuca Strait

Vessel: Vector

Party Chief: Cummins P.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2003-07

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 03/05/02 | Factory |  |  |
| Conductivity | 1763 | 23/04/02 | “ |  |  |
| Secondary Temp. | 2106 | 19/04/02 | “ |  |  |
| Secondary Cond. | 1764 | 23/04/02 | “ |  |  |
| Transmissometer | 333DR | 11/10/02 | IOS |  |  |
| SBE Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2229 | July01 | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/99 | Factory |  |  |