## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 26-Aug-2003 | Janet sent a revised nutrient file with changes in red. Changes made to the CHE files using a text editor. |
| 25-Aug-2003 | Added nutrient data to rosette files. |
| 12-Aug-2003 | Subtracted 0.0027 from the CTD primary temperature for all the CTD and Rosette files due to sensor drift suspected by Germaine and confirmed by Seabird. See file “Report on sensors 2023-1763.doc” in the DOC directory. |

## PROCESSING NOTES

Cruise: 2003-08

Agency: OSAP

Location: B.C. Inlets

Project: Paleoclimate and Oceanography

Party Chief: Thomson R.E.

Platform: Vector

Date: 31 March 2003 – 12 April 2002

Processed by: Germaine Gatien

Date of Processing: 30 June 2003 – 8 July 2003

Number of original CTD casts: 40

Number of casts processed: 39

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Transmissometer (333DR), SBE 43 Dissolved Oxygen Sensor (#0047), SeaBird pH/ORP (#160) and Seapoint Fluorometer (#2356). The DO sensor was mounted on the primary pump and the fluorometer on the secondary. The fluorometer had a 10X cable. The deck unit S/N was 0619.

A second SBE 911+ CTD (#0506) was mounted with a Transmissometer (197) and Seapoint Fluorometer (#2228). That instrument was used for only one test cast.

# SUMMARY OF QUALITY AND CONCERNS

The CTD data was generally in good order.

The Rosette log is incomplete and difficult to interpret – one cast is missing. There were a few errors in the dissolved oxygen sample files. There was some duplication of sample numbers.

There were no pressures in the dissolved oxygen files.

The difference between the temperature sensors is larger than usual. The problem is likely with the secondary sensor.

The Oxidation Reduction Potential channel was removed from all files due to very poor time response.

The dissolved oxygen sensor continues to show poor time response. Attempts have been made to correct for some of this error using titrated samples, but the errors are considered to be on the order of:

- ±1 ml/l  in the top 10m

- ±0.8 ml/l  from 10 to 125m

- ±0.3 ml/l  from 125 to 400m

- ±0.1 ml/l  below 400m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The files for cast #13 were named 2003-8-00013. These were changed to 2003-080013.

The Log Book was obtained and read.

Salinity and dissolved oxygen calibration data was obtained. There was no chlorophyll sampling.

The cruise summary sheet was completed.

The data entry in the configuration files was correct except for the serial number of the pressure sensor for CTD #0585.

File 2003-08-0550CTD.con for CTD #0585 was prepared with the corrected serial number for pressure. This file was name 2003-08-0550CTD.con by mistake, but the contents are correct for CTD #0585.

File 2003-08-0506CTD.con for CTD #0506 is a copy of 2003-08-0002.con and is to be used for cast #2.

The sensor history was found. This is the first use of the sensors on #506 since they were last calibrated.

##### Conversion of Raw Data

The raw data was converted using configuration file 2003-08-0506 for cast #2 and 2003-08-0585ctd.con for all other casts. Rosette files were converted separately with salinity channels included.

A preliminary check of a few casts shows all expected channels present.

The temperature and conductivity traces track reasonably well during downcasts and up and downcasts are similar.

The transmissivity looks ok and the oxygen looks like other data logged with this instrument – there is clear evidence of poor response time.

The fluorescence is fairly high with the usual level of spikiness.

The ORP channel shows evidence of very poor time response. The pH channel looks ok.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from pressure only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### CELLTM

Tests were run on casts #20 and 90 to find the optimal parameter choice for CELLTM. Runs were made with (0.005,5), (0.005,7), (0.01,7), (0.01,9), (0.02,7), (0.03,7), (0.02,9), (0.03,9) and (0.0245,9.5) for (alpha, 1/beta). None of these choices improved the data. Two casts with no bottle stops (#3 and 4) were examined and (0.01,7) and (0.01,9) produced very slight improvements. It is likely because of the small temperature range that no notable change results. Since the choice of (0.01,7) was used for similar equipment and condition in 2002 it was selected for this data as well. CELLTM was run on all casts for both channels using (0.01,7).

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. All expected channels were present. The differences were noisy and the values that follow are rough averages:

| Cast # | Max. Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 600 | -0.0028 | ~0.0015 | 0.001 | Fairly low but steady |
| 16 | 600 | -0.0025 | ~0.0012 | 0.0012 | 1m/s & steady |
| 48 | 320 | -0.0027 | ~0.0012 | 0.0012 | 1m/s & steady |

The temperature differences are quite high and differ between the upcast and downcast. The differences were also high in previous use of this equipment and some pressure-dependence in the differences has been noted before.

The secondary salinity is quite noisy. The salinity differences are pressure-dependent during the downcast; while noisy on the upcast the pressure dependence is not so obvious.

The conductivity differences are also a little high – this has not been noted before. The conductivity differences do not appear to be pressure dependent.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. CLEAN was used to add event numbers and to replace pad values in pressure with interpolated values.

The ROS files were converted to IOS files, CLEAN was run to add event numbers and the extensions were changed to BOT.

All casts were plotted and checked for evidence of problems with the processing or instruments. As usual there are notable differences between up and downcasts for the dissolved oxygen sensor which continues to show poor time response. Spikiness was noted in both salinity channels. There are significant differences between up and downcasts but this is expected in this region.

The ORP data shows evidence of a very poor time response. The upcasts look very different from the downcasts and the trend during bottle stops is to move closer to the downcast values.

##### Checking Headers

A header summary and a header check were produced and checked.

A number of station names were found to be in error and were fixed.

The average surface pressure is 1.6 db.

The cruise track was plotted and looks reasonable.

##### 10. COMPARE

The BOT files were examined for bad values and none were found. They were averaged and used to prepare the lookup table. The lookup table was then edited to add the bottle numbers.

The addsamp file was prepared. This contained the cast number and bottle number and a channel for sample numbers. The sample numbers were added to the file based on the rosette log. The sample numbers for cast #1 were entered as 901, 902, 903 instead of 1, 2, 3 to avoid confusion with cast #7 when the same sample numbers were used.

The sample numbers were then added to the BOT files with the output named SAM.

The SAL and OXY files lacked a quality flag channel. These were added using ADDCHANS. The output files were named SALQ and OXYQ. The oxygen files had all pressures as “0”. These were edited using the pressures from the rosette log sheets.

SALINITY

COMPARE was run. There were only 14 samples of which 6 were above 200m. Of the deep samples 1 was a severe outlier (difference greater than 0.04). One other point was a moderate outlier. With the outliers excluded the comparisons are quite flat with pressure and time. The primary salinity was found to be high by 0.0015 and the secondary by 0.0028; these are consistent with previous uses of these sensors. (See 2003-08-sal-comp.xls.)

The two outliers mentioned above plus one from shallower sampling were investigated:

| Cast | Sample # | Pressure | Comments |
| --- | --- | --- | --- |
| 1 | 3 | 80 | Bottle salinity >CTD salinity by >0.04psu |
| 48 | 185 | 322 | Bottle salinity <CTD salinity by ~0.05psu |
| 75 | 276 | 275 | Bottle salinity <CTD salinity by ~0.006psu and differs from a 2nd sample at the same pressure by 0.004psu. The sample #s in the analysis sheets for these two samples are not the same as in the rosette log. The station name is the same. |

The salinity quality flag was changed to “c” for these three samples and a comment was put in their headers to explain why.

Samples #1,2 and 3 for cast #1 were renamed 901, 902 and 903 because the same sample numbers were used for cast #7. There was no salinity sampling at cast #7 but there was oxygen sampling. There is no rosette sheet for cast #1. A note was put into the header of cast #1 and the quality flag set to “c” for the salinity samples - one was already flagged as an outlier.

DISSOLVED OXYGEN

COMPARE was run on the edited oxygen files. Warnings indicated problems with some casts. The OXY files for those casts were examined and a few errors were fixed. There were multiple entries out of order for a few samples - see the PARTICULARS section for details. The biggest problem was the lack of pressure values since 0.0 was entered for all. A text editor was used to enter the pressure readings given in the rosette log. The COMPARE output was checked to ensure that the pressures entered were close to the CTD pressures and a few further corrections made.

- A three-step approach was taken to determining the correction of the dissolved oxygen channel.

- First, COMPARE was run using the OXY files (titrated oxygen values) and the SBE_DOX channel from the CTD bottle files. This looks for errors in the calibration of the sensor. It is assumed that by the time the bottle is fired the sensor has come to equilibrium; in fact, in areas of high gradient of DOX this is far from true, but is the best we can do. As has been found for other data sets there was a lot of scatter with the best fit being differences versus dissolved oxygen rather than versus pressure or time. The results of this fit are as follows:

- Titrated oxygen = 1.152*(SBE_DOX) - 0.139

- For 2002-36 and -41 and 2002-39 the results of the same type of comparison gave:

- Titrated oxygen = 1.0996*(SBE_DOX) + 0.0213

- Titrated oxygen = 1.1054*(SBE_DOX) – 0.0565

- and	Titrated oxygen = 1.1170*(SBE_DOX) – 0.1051

- The average difference of points in the fit was -0.42ml/l with the sensor reading low. Reading low may be associated with not reaching equilibrium.

- The fit against file pair indicates little drift through the cruise. (See 2003-08-oxy-comp.xls)

- Outliers were identified (samples #100, 106, 108, 118, 176, 195, 340, 381, 383). In no case did the CTD have large standard deviations. The quality flag was set to “c” or “e” for these samples and a note put in the header. The choice of “e” was made for cases in which the rosette log indicated that the titration had failed. In cases where the value given for DOX was not zero, the value was given in the comment section and then replaced by -99 in the data itself.

- The second step in the analysis involved shifting the oxygen channel to improve the alignment relative to pressure. There is clearly a large delay in the oxygen – this step will attempt to align features in downcast and upcast data so the vertical offset is the same as for temperature. This is a fairly gross estimate since again it depends on descent rate and local gradients; moreover, the upcast temperature for this cruise was particularly noisy and the upcast contained many stops. Nonetheless, since the position of features is one of the more important uses of this data it seems important to try to get the features to a realistic pressure. For all previous cruises using this equipment a shift of +220 or +240 had the best results so tests were run using +200, +220 and +240 records. If we judge the shift by how well the up and downcast profiles of dissolved oxygen overlap then +240 records is best. If we judge by the match of depths of features then +200 looks better; this is based on the time it takes for DOX to react to the CTD moving upwards after a bottle stop. So it depends on what matters more – depth of features or absolute values. A compromise is +220 records.

- After running SHIFT and DELETE a set of recalibrated, averaged and calibrated downcast files will be compared with the titrated values. At that point a fit of differences vs pressure may be chosen to correct some of the time-response errors that remain after SHIFT.

###### FLUORESCENCE

- The offset between upcast and downcast fluorescence was compared to that of temperature, in the same way as was done for dissolved oxygen, to determine if alignment is needed. The 3 casts without stops for bottles show that the upcast is lower than the downcast by from 1.5 to 3m. Dividing this in half since the offset will apply to upcast and downcast in opposite directions and dividing by the average descent and ascent rates for the relevant depths give lags of from 1.1 to 1.7s. The fluorescence will be advanced by 24 records (1s) as has been done for all recent data sets.

##### SHIFT

###### Conductivity

The routine SHIFT was used to align the conductivity.  The deck unit was one of the newer versions that advance both channels by 0.073s. The setting for this step varies greatly from cruise to cruise. Tests were run on casts #3 and 90 using shifts between –1 records and +1 records. After shifting, the data was put through DELETE then displayed in T-S space to look for the choice that “just” removes unstable salinity spikes without oversmoothing. The differences were slight but shifts of –0.5 records did improve the primary salinity. No shifting will be done to the secondary conductivity since the primary looks better and has been used for all previous casts with this equipment.

###### Dissolved Oxygen (except for cast #2 which had no DOX channel)

Based on the analysis in the previous section the dissolved oxygen sensor was advanced by +220 records.

###### Fluorescence

The fluorescence was shifted by +24 records.

Oxidation Reduction Potential

An attempt was made to use SHIFT to make up and downcasts look alike. Shifting by as much as +1000 records brought the two closer together but not enough to be useful. Because of the long time constant the profiles are too smooth to enable an estimate of the lag by comparing distinctive features.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt (minimum salinity = 5 psu)

Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0 				Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates < 0.3m/s (calculated over 11 points) was deleted between 10db and 10db above the maximum pressure.

Sample interval =  .0412 seconds. (taken from the header)

COMMENTS ON WARNINGS: There were no warnings in the DELETE log.

All DEL files were copied to EDT files.

##### Final comparison of SBE DOX

The downcast data was averaged in 1m bins and thinned to depths corresponding to bottles. The DOX channel was recalibrated using the results of section 10 (2003-08rcal1.ccf) which corrects for the errors in the sensor itself. A comparison was then run between the titrated values and the thinned, recalibrated downcast data. (See 2003-08-oxy-comp2.xls.) There are two sources of error in what remains - time response errors not fixed by SHIFT and real differences between up and downcast. Since we cannot distinguish between the two we just hope the latter is not too significant. A number of fits were tried and a linear fit against pressure seems the best choice. A second recalibration was done using file 2003-08rcal2.ccf, based on the 2nd run of COMPARE, in an attempt to remove the time response error. There remain significant errors particularly in the top 200db. A final comparison shows errors on the order of ±1 at the surface, ±0.8 from 10 to 125db, ±0.3 from 125 to 400db and ±0.1 below that. (See 2003-08-oxy-comp3.xls)

##### DETAILED EDITING

The primary sensors were chosen for further processing because they are closer to the bottles than the secondary.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Salinity was cleaned where large spikes occurred. Small spikes were cleaned only if it was clear they were due to imperfect alignment of T and C. Records corrupted by shed wakes or ship effects were removed from many cast mostly near the surface and bottom.

Cast #52 required no editing and cast #97 required heavy editing. All other casts required light editing only. Most contained small spikes in salinity.

Note was made of the editing details in the relevant files.

The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### CALIBRATION

The dissolved oxygen channel was recalibrated in a two-step process using 2003-08-racl1.ccf and 2003-08-rcal2.ccf. The salinity was not recalibrated.

The first recalibration was applied to the oxygen channel in the SAM files. The output files were named SAC.

##### SPECIAL FLUOROMETER PROCESSING (FILTER AND BIN AVERAGE)

The files were clipped to 100db and stored in a separate directory for the use of Angelica Peña.

A median filter with fixed size 11 was applied to the SeaPoint fluorometer channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE

The following Bin Average values were applied to the DEL files for general use:

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

The BOT files were also averaged with Bin Channel = Bottle Number. The results were called SACAVG.

##### Other comparisons

Previous experience with these sensors – The conductivity sensors have been used five times since the last recalibration but there were no calibration samples available for the most recent cruise.

The primary salinity was found to be high by 0.0015, low by 0.0011, high by 0.0008 and high by 0.0004psu. The secondary was high by 0.005, 0.0019, 0.0037 and 0.0030psu.

- The oxygen sensor has been used 6 times since the last calibration. A 2- step recalibration was done plus a shift of the oxygen channel relative to pressure. For most of the casts an initial calibration based on the linear fit of differences to oxygen concentration was followed by one based on a linear fit of the remaining differences to pressure. For 2003-01 a 2nd order fit was used for the first step.

There is no history for the sensors on the 2nd CTD system which was used for cast #2 only.

Historic ranges – There were only 2 casts for which local climatology was available and the data fell within or very close to the historic ranges. The salinity looked slightly low around 100db for one cast.

Comparison of CTDs – There was a repeat cast in Saanich Inlet using the 2nd CTD system but one of the casts went to 100db only. The 2nd system has not been used since the latest calibration and there was no calibration sample so the comparison will be of no use in assessing the main system. It may be of interest for future use of the secondary system. The primary salinity was found to be low by about 0.028psu and the secondary by 0.030psu compared to the main system. The main system is probably high by up to 0.015psu and the secondary by 0.028psu, so the primary salinity for the secondary system is probably low by about 0.015 and the secondary is probably quite accurate. (See CTD_585_506_comp.xls)

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data.

##### REMOVE

The ORP (Oxidation reduction potential) channel will be removed because of the very long time constant and the lack of any means to calibrate it.

The following channels were removed from all CTD casts (Output – REM): Scan_Number, Temperature: Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, ORP, Descent Rate, Status:Pump and Flag.

The following channels were removed from all rosette files (Output – SACREM): Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, ORP, Status:Pump and Flag.

##### HEAD EDIT AND STANDARDS CHECK

EDIT HEADERS was used to fix the data format of CTD and bottle files and to add a warning to the headers of the CTD files as follows:

******************************************************************

The con file used to convert this file although named 2003-08-0550CTD.con,

contains the calibrations for CTD system #0585, the system used for this cast.

The time response of the dissolved oxygen sensor is poor. SHIFT was used and

corrections have been applied based on titrated samples, but large errors remain

and are considered to be of the order:

* ±1 ml/l  in the top 10m

* ±0.8 ml/l  from 10 to 125m

* ±0.3 ml/l  from 125 to 400m

* ±0.1 ml/l  below 400m

The pH, Fluorescence and Transmissivity data are nominal.

The pH, Dissolved Oxygen and Transmissivity are unedited except that some full

records were removed in the course of editing temperature and salinity. The

same is true for fluorescence except that it was put through a median filter.

******************************************************************

The final files were named CTD.

##### Merging of bottle files

##### The SACREM, OXYQ and SALQ files were put through SORT to arrange by sample_number.

##### SACREM was put through CLEAN to remove the SeaBird headers (SACCLN).

SACREM and OXYQ were merged and the results were merged with SALQ. The output was put through HEADEDIT twice, first to fix headers and channel names and formats, and the second to add a standard header comment explaining the codes and methods. The final files were named RAC.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### Particulars

1. Test cast for CTD #0585

1. There were 3 salinity samples from this cast but no rosette file. The samples were numbered from 1 to 3. Those numbers were also used for cast #7 for which there is a rosette sheet. The samples from cast #1 will be renumbered 991, 992 and 993 to avoid confusion. The quality flag was set to “c” for all three samples and a note put in the header to explain why.

2. Test cast for CTD #0506. The equipment has not been used since the latest sensor calibrations and there was no salinity sample taken for calibration. The salinity looks low compared to the other CTD. This cast will not be archived.

13. Original files named 2003-08-00013; changed to 2003-08-0013

50. Sample #205 appears twice in the OXY file with the second entry out of order– first entry with value 0, 2nd with 2.91 which agrees with the Rosette log entry. The value of 6.03 was used and the repeat removed.

75. Two bottles at 275db have quite different salinity and oxygen values. There is no such change in the CTD values at that depth.

79/80. The cast identified as #79 in the rosette log should be #80. The oxygen file was renamed.

94. Sample #356 appears twice in the OXY file with the second entry out of order– first entry with value 0, 2nd with 5.148 which agrees with the Rosette log entry. The value of 5.148 was used and the repeat removed.

97. The Thiosulfate Titre was blank for sample #362.

98. Station name in header QCS1 – should be QCS6. This was fixed in CTD and Bottle files.

99. Sample #382 appears three times in the OXY file with two entries out of order– 1st and 2nd entries with value 0, 3rd with 6.03 which agrees with the Rosette log entry. The value of 6.03 was used and repeats removed.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-08

Dates: Start: 31 March 2003 End: 12 April 2003

Location: B.C. Inlets

Vessel: Vector

Party Chief: Thomson R.E.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0585

Cruise ID#:		2003-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 03/05/02 | Factory |  |  |
| Conductivity | 1763 | 23/04/02 | “ |  |  |
| Secondary Temp. | 2106 | 19/04/02 | “ |  |  |
| Secondary Cond. | 1764 | 23/04/02 | “ |  |  |
| Fluorometer –pumped | 2229 | 07/01 | IOS |  |  |
| Oxygen SBE43 | 0047 | 21/08/02 | Factory |  |  |
| Transmissometer | 333DR | 11/10/02 | IOS |  |  |
| pH/ORP Sensor -SeaBird | 0160 |  |  |  |  |
| Pressure Sensor | 75636 | 30/09/1999 | Factory |  |  |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/	0506

Cruise ID#:		2003-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20Jun02 | Factory |  |  |
| Conductivity | 2128 | 14May02 | “ |  |  |
| Secondary Temp. | 2374 | 20Jun02 | “ |  |  |
| Secondary Cond. | 1766 | 20Jun02 | “ |  |  |
| Fluorometer –pumped | 2228 |  | IOS |  |  |
| Transmissometer | 197 | 22Jan03 | IOS |  |  |
| Pressure Sensor | 69698 | 11Dec97 | Factory |  |  |