## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-Nov-2017 | Corrected pad values in data section in Number_of_bin_records of 2003-11-54.che to match that in header. Roy Hourston. |
| 13-Jan-2014 | Added underway pCO2 data from Sophia Johannessen’s Excel files prepared for The Carbon Dioxide Information Analysis Center (CDIAC). The file is located in the cruise .DOC directory. |
| 18-Dec-2013 | Changed units for the Flag:Oxygen:Dissolved channel. J.L. |
| 28-Oct-2013 | Merged DMS data to bottle casts from Mike Arychuk’s spreadsheet file located in the \DOC directory. For information on methods see file "Method post 1999 for Dimethylsulfide Analysis.doc” in directory \\OSD_Data_Archive\Cruise_Data\Documents\Analysis Reference Papers\. |
| 20-Feb-2012 | Added one productivity cast, (9001), containing Primary Productivity, Chlorophyll, POC and PON, from Frank Whitney’s Productivity spreadsheet. |
| 7-Oct-2005 | CTD dissolved oxygen was recalibrated based on the 20-Sep-2005 corrections to titrated DO. All SBE DO values were multiplied by 0.98. |
| 20-Sep-2005 | Sheila Toews recalculated oxygen with new standards so the old oxygen values were replaced with the new ones for most of the rosette casts. |
| 25-Feb-2004 | Nina Nemcek found some errors in the chlorophyll and reported them to Marie Robert. See Note at end of document. |
| 8-Oct-2003 | Merged nutrient data with rosette files. Further comments in RED in section 10. |

## PROCESSING NOTES

Cruise: 2003-11

Agency: OSAP

Location: North-West Pacific

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: May 26, 2003 – June 16, 2003

Processed by: Germaine Gatien

Date of Processing: 12 August 2003 – 23 September 2003

Number of original CTD casts: 74 (including 4 upcasts)

Number of casts processed: 70

# INSTRUMENT SUMMARY

Two SeaBird Model SBE 911+ CTDs (#0585 and 0550) and two different deck units were used during this cruise as follows:

Casts 1-9: SeaBird Model SBE 911+ CTD (s/n #0585) and SeaBird model 11 deck unit (#0508)

Casts 10-11: SeaBird Model SBE 911+ CTD (s/n #0550) and SeaBird model 11 deck unit (#0508)

Casts 12-end: SeaBird Model SBE 911+ CTD (s/n #0550) and SeaBird model 11 deck unit (#0619)

The CTDs were mounted with Transmissometer 333DR, Seapoint Fluorometer (#2229), SeaBird Model and SBE43 Dissolved Oxygen sensor (#47). The PAR sensor (#4565) was used only for casts 1-6.

# SUMMARY OF QUALITY AND CONCERNS

This was a complex cruise with changes of equipment and problems with cables, computer crashes, and bottle firings. The notes provided to the processor by the chief scientist were a great help. While not all problems were noted, they were sufficient to give a “heads up”, preventing errors and speeding up the job.

There were many spikes in pressure, temperature and conductivity; associated with these spikes are jumps in pressure as though not all data was recorded.

At times there is noise in temperature and salinity that suggests poor flow rate. This occurs mostly in sections near the surface with fairly low descent rate. Once the average descent rate gets to about 1m/s this problem largely disappears.

The processing of this cruise was complicated by problems beyond the equipment malfunctions. There were a number of errors in the bottle calibration data; there are discrepancies among rosette sheets, analysis sheets and sample #s leading to difficulties in determining the level from which some samples were taken and a few errors were made in transferring data from analysis sheets to spreadsheets. There is scatter in the comparison of bottle salinity to CTD salinity that appears to be related to analysis problems or bottle leakage, but there remains some doubt about the CTD data itself. The result is less confidence in this data than usual.

The dissolved oxygen sensor continues to show poor time response but the calibration against bottles during bottle stops looked reasonably good. Attempts were made to correct for time response problems but errors are considered to be on the order of:

- ±0.45ml/l in the top 250 (most are within 0.25ml/l, with the largest differences between 150 &250db)

- ±0.15 ml/l from 250m-1500m

- unreliable below 1500m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Salinity, oxygen and chlorophyll calibration data were obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. Errors in the serial number of the pressure sensor for CTD 0585 and the date of calibration for CTD 0550 were corrected and the resulting files named 2003-11-0585.con and 2003-11-0550.con.

The sensor history was found.

##### Conversion of Raw Data

The raw data was converted using configuration files 2003-11-0585.con and 2003-11-0550.con. Rosette files were converted.

The PAR channel was converted for casts 1-6 only.

A few casts were checked to ensure that all expected channels contained reasonable data. As mentioned in the log there are many spikes in some casts. Otherwise the data looked ok.

##### WILDEDIT

There were many problems with spiky casts during this cruise. An example of a terrible cast was #12. A test was done to see if there was a need to run WILDEDIT on just pressure or if more channels should be included. Including the other channels did help. PAR will not be put through the routine since it removes too much data.

Program WILDEDIT was used to remove spikes from all channels except PAR.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

CTD #0585:

For 2003-15 which immediately preceded this cruise (0.02,7) was used for (alpha, 1/beta). For this cruise the same equipment was used. A check was made to see if the same parameter was effective for this mission. Casts #1 and #5 were checked using (0.02,9), (0.03,9), (0.02,7) and (0.03,7). The answer was not clear for the primary with different settings looking better in different areas, but for the secondary (0.03,9) looks best. The results of 2003-12 will be used for the primary conductivity.

CELLTM was run on all casts using (0.02,7) for the primary and (0.03,9) for the secondary.

CTD #0550:

Casts # 31, 78 and 87 were studied and (0.03,9) was found best for both channels.

CELLTM was run on all casts using (0.03,9) for both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

In previous use of CTD #0585 it was noted that the temperature sensors were farther apart than usual and that continues to be the case. The conductivity and salinity differences are similar to those of 2003-15 and the secondary conductivity continues to be very noisy. The secondary salinity is naturally very noisy too. These sensors were among some recently sent to SeaBird and they report that one of the conductivity cells was cracked. And the drift in temperature reported since the previous calibration is fairly high.

For CTD #0550 the secondary temperature is slightly noisier than the primary. The up and down traces and up versus down look reasonably close. The differences between conductivity and salinity show a gradual drift during the cruise. This is the first use of these sensors since their last calibration.

The differences between sensors were generally noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate | CTD |
| --- | --- | --- | --- | --- | --- | --- |
| 5 | 800 | -0.0045 | +0.0002 | +0.006 | Noisy, Moderate | 0585 |
| 17 | 1500 | +0.0003 | -0.0001 | -0.0018 | OK, high | 0550 |
| 23 | 1500 | +0.0002 | -0.00005 | -0.0007 | Noisy, High | 0550 |
| 47 | 1500 | +0.0004 | +0.0002 | +0.0018 | Noisy, High | 0550 |
| 64 | 1500 | +0.0004 | +0.00025 | +0.0025 | OK, very high | 0550 |
| 74 | 1500 | +0.0004 | +0.00025 | +0.003 | OK, high | 0550 |
| 75 | 1500 | +0.0003 | +0.0003 | +0.003 | OK, very high | 0550 |

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

The rosette files were converted to IOS files and renamed BOT.

All BOT files were plotted and outliers were found in casts #20, 49 and 55. CTDEDIT was used to remove records that appeared to be bad for all channels and salinity was cleaned where there was noise in a single salinity channel only. The edited BOT files were copied to the original BOT files.

##### Checking Headers

In plotting the cruise track it was noted that the positions were converted incorrectly in 3 files: 4, 39 and 54. These were fixed in both the CTD and BOT files and the track re-plotted. It looked reasonable.

A header summary and a header check were produced.

Errors were found in the instrument summary for casts #8 and 9. Those two casts were reconverted with the instrument number corrected. While the instrument summary lists PAR for these two casts, in fact that channel was not converted.

The average surface pressure is 1.6db.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

There were some inconsistencies between the rosette log, CTD daily log and the bottle files.

- For cast #10 there are two entries in the daily log. There is no rosette sheet. The first entry in the daily log indicates that there was a bottle, sample #54. However, this cast was aborted and restarted using the same cast number. The second entry of this cast in the daily log indicates that during this restart there were 3 bottles, sample #s 54, 54a and 54b. In the bottle file there are 2 bottles, from ~2000db and ~9.5db. In the salinity analysis sheets there are 2 bottles #s 54a and 54b, both said to be from 2000m. On the analysis sheet there is a note beside these two samples saying: “called 5401 and 5402 in file”. The values are similar and do look like they are from the bottom of the cast. Sample #54 is included in the loop sampling spreadsheet. The sample numbers were changed to 5401 for the deep bottle and 54 for the shallow bottle in both the addsamp.csv file and the salinity file. A note was added to the header of the salinity file giving the original file name. A text editor was used to reorder the SAL file so that the sample numbers are increasing, in order to allow the merge process to operate properly.

- For cast #33 there were 23 bottles fired. Both the BOT files and notes in the rosette log agree on this. However, there were 24 sample numbers assigned and 24 salinity samples analyzed. The sample that is said to be from bottle #9, sample # 175, has salinity that makes sense for 1000db, the depth from which it was intended to be taken, but the confirmed pressure for that bottle is 800db. Is it possible that the bottle did close but no confirm signal was received and no entry went into the BOT file. The samples from 800db to the surface all look closer to the CTD at the planned depth rather than the confirmed depth. A new addsamp file, addsamp1.csv, was prepared to reflect what seems to be correct by skipping sample #175 assuming it is from 1000db and assigning #176 to 800db etc. The last bottle will be sample #190, pressure 10db, bottle number 23. This agrees with the salinity analysis sheet but not the rosette sheet. Sample #175 was reinserted by the merge program. Bottle salinity and oxygen were entered from the analysis sheets. The bottle salinities seem to indicate that there were 24 bottle tripped.

- For casts #36 and 37 there were single bottles which the daily log notes as having sample numbers #191 and 192; however, those numbers had already been assigned during cast #35.  The sampling for #36 and #37 was loop sampling only. Janet Barwell-Clarke was notified of the potential confusion in sample numbers. She confirms that the numbers were repeated, but there appear to be no problems in the nutrient samples.

- Cast 40 was deleted since there were no chemical samples.

- For cast # 46 there is a single bottle in the BOT file, but there was no sampling. Sample numbers are mentioned in the daily log, but the cast was aborted and the numbers reused for the next cast; this BOT file will not be processed further.

- Cast #47 was interrupted by a computer crash and restarted as cast #48. The daily log gives all the bottles as being from cast #47. In fact, there are only 7 in cast #47 and the rosette log confirms this. The chlorophyll samples are all from #48 and are so named. The sal and oxy files are merged so were split to create two separate files as input to COMPARE.

- Cast #48 has 24 bottles in the BOT file, but the rosette log shows only 17. The first 8 bottles in the file were fired at the bottom. Presumably this is an error associated with the computer crash.  The first 7 trips were deleted since there were no samples taken and the rest of the records were merged with cast 47.

- Cast #54 shows bottles # 1-4 and 6-9 being fired with sample numbers 248-251 and 253-256. The BOT file contains numbers 1-8 but the pressures correspond to those on the rosette sheet. The sample numbers assigned to bottles 6-9 on the rosette sheet were assigned to bottles 5-8 in the BOT file. This was another cast with a crash at the bottom which probably accounts for the mix-up.

- Cast #55 has 14 bottles in the rosette log, but there are only 13 bottles in the BOT file. A correction to the rosette log indicates that the bottle shown as being fired at 35db did not fire there. All the bottles from bottle #6 up correspond to the pressures of the next bottle up as given on the rosette sheet.

- For cast #64 the rosette log has a cryptic note beside bottle numbers 3 and 4 (samples #356 & 357) that may be saying the samples are reversed. In fact, they do not appear to be reversed, but both bottle salinity samples were flagged because of problems with the caps. There was no dissolved oxygen or chlorophyll for those sample numbers.

- Cast #92 shows a sample for bottle #19 but the BOT file has only 17 bottles. It is assumed the extra bottle was a loop sample, or some other sort of sample not intended for the bottle files. The data in the BOT file corresponds to the pressures given for bottles #1-17 on the rosette sheet.

Those BOT files for which there was no rosette sheet were not processed further since they will be included in the loop sampling analysis.

The ADDSAMP file was input to the ADD SAMPLE NUMBER routine to create SAM files from the BOT files. Those files were averaged on bottle number. (Output: SAMAVG)

The salinity and chlorophyll analysis spreadsheets were converted to individual files: SAL and CHL.

The following general comment was added to all *.CHL headers using HEADER EDIT:

All samples were analyzed onboard by Philippe Tortell and Nina Nemcek from UBC, using the Chemistry fluorometer calibrated August '02.

There were no flagged values in the CHL files.

For the flagged values in the SAL files, individual explanatory comments were added to the headers.

The SAL, OXY and CHL files were merged with SAMAVG in three steps. (Output: MRG1, MRG2, MRG)

After the final merge cast #10 was again reordered so that pressure is decreasing.

##### 11. COMPARE

DISSOLVED OXYGEN

COMPARE was run. Points flagged by the analyst were excluded from the comparison. Points were excluded if the pressure was not in the range of 10 to 1500db. If a point was identified as an outlier in one fit, it was excluded from all three fits. The fit versus pressure suggests that the dissolved oxygen sensor did not perform well below 1250db. There appears to be some drift with time but this is probably a reflection of the different oxygen ranges in different parts of the mission. As usual, the best fit is versus oxygen concentration as follows:

Titrated oxygen = 1.1386*(SBE-DOX) – 0.0374 (See 2003-11-oxy-comp1.xls)

That compares reasonably well with the results of cruise 2003-015 which preceded this mission

- Titrated oxygen = 1.1474*(SBE_DOX) - 0.1024.

- The only outliers that were not already flagged by the analyst were not sufficiently out of line to be flagged or were at depths were the sensor is likely to perform poorly.

- After running SHIFT and DELETE a set of recalibrated, averaged and calibrated downcast files will be compared with the titrated values. At that point a fit of differences vs pressure may be chosen to correct the data further to remove time-response errors not corrected by SHIFT.

Salinity comparison

COMPARE was run separately on casts 1-9 and 10-96 since different CTDs were used. A first glance at the comparisons showed more scatter in the differences than usual. Since cast #62 included 24 bottles at a single depth, it was analyzed separately to see if it led to some understanding of why the differences are so noisy for this cruise.

COMPARE was run on cast #62 only using sample numbers as the reference channel to study the consistency of the salinity analysis. There were 24 salinity samples taken at ~1500db. Of these, two were flagged. The differences between the CTD values and the bottles shows a drift for both salinity channels even when the flagged values and one outlier are excluded. The CTD values themselves show little variation with the range of values being 34.5122 – 34.5128 units for the primary 34.5151 to 34.5156 for the secondary. Looking at the salinity analysis sheets there is no note of instability. The samples were analyzed in order of sample number. The bottle salinities vary from 34.5129 to 34.5149 when the outlier (34.5165) is ignored. This range of 0.002 would appear to be due to problems with the salinity analysis. The analysis was done at sea and it is possible that temperature was not controlled as well as usual. The variations look random for the first 11 samples; the variations are smaller after that and are probably varying systematically. Another possibility is that some bottles were leaking; there is evidence that this was the case during the autumn Line P cruise. (See 2003-11-0062-sal-comp.xls)

Repeat sampling at cast #71 shows more serious problems with the two bottles having salinities differing by 0.011 units whereas the two CTD readings are within 0.0001 of each other.

CTD #0585

The primary sensors were sent to SeaBird for recalibration and repair before this cruise. They reported damage to the conductivity cell and a drift of 0.004 PSU/year in the salinity, but they also report a drift in the primary temperature sensor of 0.00264 Cº/year. For details on the ramifications of SeaBird’s observations see Report on sensors 2023-1763.doc.

There were only 2 bottle casts with CTD #0585. Looking at data from pressures greater than 200db the primary salinity is low by 0.0064 and the secondary high by 0.0001. This is close to the results of cruise 2003-15 before recalibration of the primary temperature. Since there are very few deep bottles for casts 1-9, and the casts were very spiky, the same recalibration scheme that was applied to the primary temperature and salinity for 2003-15 will be applied to this data. That involved recalibrating primary temperature and recalculating salinity followed by a recalibration of the primary salinity. There were no significant outliers and no changes made to the flags.

(See 2003-11-sal-comp1.xls and file 2003-11-0585-racl.ccf and 2003-11-0585rcal2.ccf)

CTD #0550

The first run of COMPARE led to the discovery of 2 typos in the salinity spreadsheet and errors in the depths assigned to some samples due to confusion in the rosette sheet and sample numbering. The greatest confusion was over cast #33; the details are described in section 10. After these corrections COMPARE was run again.

For cast #62 there were 24 samples from ~1500db. Two of those values were flagged and one is a clear outlier when compared to the others. A single representative sample was chosen (closest to the average of the reasonable values). When values from above 200db and a few other outliers are removed from the comparison, the primary salinity was found to be low by 0.001 and the secondary is high by 0.001 units. The primary differences are noisy but the trendline is very flat. The secondary differences are also noisy but appear to have much more pressure dependence and time-dependence. Since the earlier casts were generally much deeper than the later ones it is difficult to determine whether this is significant. However, the results of section 7 do show a drift in differences between the sensors over time. Combining these results it seems likely that the calibration of the secondary sensor was drifting significantly.

The primary salinity will be used for the archive and will not be recalibrated. While the post-cruise calibration suggests that the salinity is low this is not confirmed by the salinity bottles and may indicate that significant drift occurred between May and September 2003.

The quality flags were set to “c” for 3 outliers that were not flagged by the analyst (casts #10, 68 and 71). Comments were added to the headers to indicate why they were flagged. (See 2003-11-salcomp1.xls.)

Fluorescence-chlorophyll comparison

The comparison of all fluorometer and chlorophyll data showed a lot of scatter. When samples at pressure<10db and outliers were excluded, the following relationship was found:

CHL = 0.3046* FL - 0.0927

Looking in more detail there appears to be a change between cast #74 and 79.

For the offshore casts (1 – 74) the relationship is

CHL = 0.3061* FL - 0.0846

or FLUOR ~3 * CHL.

For the inshore casts in the region of Queen Charlotte Sound, Queen Charolotte Strait and in inlets the fluorometer reads lower than the chlorophyll. Many of these records were removed from the analysis for the whole mission since they are outliers. The fluorometer was occasionally near the top of the scale where we can expect problems. But even when the chlorophyll is fairly low the fluorescence is lower than the chlorophyll.

A quick check was made to see if using metre-averaged, shifted, downcast fluorescence values gave better results and they appear to be a little closer to the bottles with CHL ~ 0.46 * FL using all casts, FL<14, Std.Dev.<0.1. This was not done very carefully, but does suggest that using downcast data might give different results. (See 2003-11-chl-comp-dn.xls)

For two recent cruises the results were CHL = 1.116 * FL + 0.09 (2003-15) and CHL = 0.9 * FL - 0.03 (2003-12). For 2003-15 it was noted that the fluorescence was significantly lower than chlorophyll in Queen Charlotte Sound and Knight Inlet.

(See 2003-11_chl_comp.xls for details.)

##### 12. SHIFT

Conductivity

CTD #0585

Based on the results of 2003-08 (-0.5 records), 2003-12 (-0.6 records) and 2003-15 (-0.5 records) , tests were run using advancements of -0.4, -0.5, -0.6 and -0.7 records. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best result was with a shift of -0.5 records for a net advancement of about 0.05s (since the deck unit had advanced it by +0.073s). No study was made of the secondary conductivity since the data looks extremely noisy and is unlikely to be used.

All data was put through SHIFT using -0.5 records for the primary sensor only.

CTD #0550

Tests were run on a wide variety of settings and the best setting appeared to be 0 and +.5 for the primary and 0 and -0.5 for the secondary. Overall using 0 seemed best for both channels so no shift will be applied.

Because only a few files were put through this step all IOS files were copied to SHF1 first; the shifted files were then overwritten to SHF1 so there is a complete set of SHF1 files.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for 1 cast using CTD #0585 and 3 using CTD #0550 were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The range of values was from 1s to 2s. Since there are likely to be more flow-rate problems in the upcast than the downcast the conservative choice of 1s was made; a shift of +24 records was applied. This is the same figure that has been applied to all recent data sets.

Dissolved Oxygen

- The dissolved oxygen sensor continues to show transit-time and/or time-response problems, but this data continues a trend to somewhat better results in 2003 than in 2002. During bottle stops the oxygen values reach equilibrium faster. In the past the setting that was chosen made up and downcasts look reasonably similar, but did not match the time taken to reach equilibrium during stops. This may mean that the flow rate to the sensor was better for this cruise and the delay is just that due to the response time of the sensor. The flow-rate problem may have been more severe during stops in the past.

- To determine the best shift value to apply, the data was studied in three different ways.

- During stops for bottles the difference was found between the time taken for the temperature and salinity to reach equilibrium and that for the DOX to do it. The conditions were rough so the CTD was moving a lot making this a difficult estimate, but about 11s or 260 records is a reasonable guess. At times it took much longer but this was during stops that were especially rough so it is possible there were flow-rate problems. Another way of using information from bottle stops is to see how close the end point of the oxygen data is to the downcast data; we don’t expect them to be equal but to close the gap by about 50%. Using this standard a shift of no more than +200 records is a better choice.

- Features in upcast and downcast traces were compared for depth. This is difficult with DOX data because the poor time response reduces the sharpness of features and the errors are a function of gradient. A setting that looks good for one depth is not so good at another. Nonetheless an estimate of from 200 to 240 records was made.

- Studying what shift makes the offset between downcast and upcast DOX equal that between downcast and upcast temperature. Estimates of at least +260 records were made looking at small sections of profiles away from bottle stops.

For all previous cruises using this equipment a shift of +220 or +240 was deemed best, with +220 used for all recent cruises. It is likely that the discrepancies in the different methods of analysis is due to differences in the transit times during upcast and downcast, and the instrument response in moving from high gradient to low being quite different than moving from low to high. Since no clear answer has emerged from this analysis +220 records will be used.

All casts were shifted by +220 records.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values. On-screen plots T0, T1, S0 and S1 vs pressure were examined to check for any processing errors. None were found.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min   	Maximum Surface Pressure (relative): 10.00

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: There were warnings for 8 casts. Of these, 4 refer to casts that contain only upcast data. Those casts (#13, 25, 48 and 54) will not be processed further. The problems with cast #55 refer only to the upcast. Casts #10, 11 and 53 have sudden jumps in pressure that can not be believed.

##### DETAILED EDITING

CTD #0585: Based on experience from other missions the primary sensors were selected; while the secondary salinity was closer to the bottles the signal was very noisy. The primary sensors will need recalibration of both temperature and salinity. CTDEDIT was used to edit these casts. There were some odd unstable features that were not related to shed wakes; they were generally small enough and fairly random so that metre-averaging will remove most of these effects. In some cases it looks like the flow-rate was unsteady, but in others there is no obvious explanation. This equipment has been used often over the past year and this behaviour was not observed before or after this cruise. This CTD was not used after cast #9 because of fuse problems but similar problems were noted with the other CTD which led to a change of deck unit. The first cast after that also had spikes. The pump status looks ok.

CTD #0550: The primary sensors were chosen because the secondary salinity looks pressure dependent and may be drifting with time.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case. In some cases the descent rate was examined to determine if unstable features should be removed or not.

All casts required a little editing except for #82 and 83.

Most casts required only light editing but the following casts were edited more extensively: #5-15 ,18, 22-29, 32, 41, 53, 56, 82-83.

Jumps in pressure were noted for casts #10, 12, 15, 24, 41, 53, 55 and 56. There may be others that are affected as well. Problems were noted at sea and were thought to be cable-related. After cast #59 the CTD was moved to a different winch and no problems were noted after that time. The jumps are not associated with DELETE having removed records. They occur in sections in which there are some records with spikes that were set to PAD values by WILDEDIT. There are often spikes in T and C as well. It looks as though the CTD did not record all records at these times, as the pressure changes are impossibly fast.

The descent rate was kept high for the most part. There are a number of casts for which the descent rate was very noisy, but because the average descent rate was kept very high there is very little evidence of shed wakes.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### Other Comparisons

Previous experience with these sensors –

- The primary sensors on CTD #0550 have not been used since the last calibration. The secondary sensors have been used twice and gave values low by 0.004 to 0.006 units with significant pressure dependence.

- The sensors on CTD #-0585 have been used frequently. As mentioned in section 11 the sensors were recalibrated after this cruise and significant problems were found with the primary temperature.  For 2003-12 both sensors were found to be low, by 0.0056 and 0.0013, respectively. For 2003-15 the primary was found to be low by 0.0072 before recalibration of the temperature and low by 0.0044 after that recalibration.

Post-cruise calibration sensors –

- The primary sensor for CTD 0585 was recalibrated two months after the cruise and was found to have drifted by -0.0004 units per month since the previous calibration. If the drift was steady then it would be low by about -0.0027 at the time of this cruise. This is reasonably consistent with the COMPARE results of -0.0064 since the error in temperature accounts for about -0.003 of the difference.

- The secondary sensor for CTD 0585 was recalibrated in Sept. 2003 and if the drift is assumed to be constant the salinity during this cruise would be low by 0.0013 units. COMPARE gives differences close to zero, but there was little data and none deeper than 1300db.

- The primary sensor for CTD 0550 was recalibrated in Sept. 2003 and if the drift is assumed to be constant the salinity would have been low by 0.0088 during this cruise whereas COMPARE suggests it is low by only 0.001.

- The secondary sensor has not been recalibrated since this cruise.

Historic ranges – There were no significant excursions from the historic ranges where local climatology was available.

##### RECALIBRATION

- CTD #0585 – casts 1 – 9

The EDT, MRG and SAM were recalibrated using file 2003-11-0585-rcal1.ccf to subtract 0.0027 from the primary temperature and recalculate salinity, and then by 2003-11-0585-rcal2.ccf to add 0.0044 units to the primary salinity and to recalibrate the dissolved oxygen using the equation given in section 11.

- CTD #0550 – casts 10-96

No temperature or salinity recalibration was applied. The oxygen was recalibrated using 2003-11-0550-rcal2.ccf to apply the equation given in section 11. (There is no 2003-11-0550-rcal1.ccf – the ccf file name was chosen for consistency with casts 1-9 that required a 2-step initial recalibration and the output extensions were named *.cor2, again for consistency)

COMPARE was rerun for salinity for CTD #0585 after the recalibration (2003-11-sal-comp2.xls) and for oxygen for all casts (2003-11-comp-oxy3.xls) and both results were satisfactory. (Output: COR2; MRGCOR2; SAMCOR2)

##### FINAL OXYGEN COMPARISON

A set of downcast files were created (from COR2 files, so after SHIFT and the 1st recalibration of the oxygen channel) thinned to pressures close to those used for the bottles during this cruise. COMPARE was run comparing those files to the bottle channel in the MRG files. (See 2003-11-comp-oxy3.xls.) A final recalibration was applied to the thinned, downcast SBE dissolved oxygen channel using 2003-11-cal3.ccf based on the results of the fit of differences versus pressure. COMPARE was run again and the results were reasonable. (See 2003-11-comp-oxy4.xls.) This final recalibration was applied to the downcast COR2 files only, not the bottle files since the lag is not an issue in the bottle files. (Output: COR3)

##### Special Fluorometer Processing

The COR3 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

The full bottle files were also saved after being put through REMOVE and HEADEDIT to produce files named BOF which were also saved to a CD-ROM. Copies of the processing report and file document were put on the CD-ROM as well.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR3 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: BOX)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S and Transmissivity profiles.

Profile plots were produced with Dissolved Oxygen, Fluorescence and PAR versus Pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Bottle_Number, Oxygen:Voltage, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment, fix formats and channel names using file 2003-11-header.hdr. The Standards Check routine was run and no problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRGCOR2 files were put through CLEAN to remove the SeaBird headers and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, and Flag.

HEADER EDIT was run to add a standard comments about quality flag definitions and to fix formats. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Quality Control Plots

A few nearby casts were plotted together to check for consistency. In particular, casts at which spiking was a problem were examined. There were no variations large enough to be ascribed to instrumental problems. Most repeat casts include one shallow cast and one deep; the only exception is for P26 where a comparison as deep as 800db is possible. No significant differences were found.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1-6. PAR on and CTD #0585

8-9. PAR off and CTD #0585

10-96. PAR off and CTD #0550

4. Cruise number wrong in header. Corrected

10. Lot of spikes and sudden jumps in pressure

11. Lot of spikes and sudden jumps in pressure; no bottle file – bottles were fired but there was no confirmation.

12/13. Crash at bottom, new file started for upcast. Lot of spikes in #12, not in #13

24/25. Crash at bottom, new file started for upcast

33. 23 bottles closed but 24 samples taken. Apparent mislabelling of samples was corrected in the bottle files.

35. Bottle number 13 did not trip.

39. Wrong station name in file. Corrected

42. Error light turned on during cast. Spikes

47. Lot of spikes

48. Lot of spikes

49. Cable re-terminated before this cast.

53/54. Crash at bottom, new file started for upcast. Spikes and jumps in pressure. Errors in headers corrected

55. Spikes and jumps in pressure during upcast

59. Cable re-terminated before this cast.

95. Two bottles closed. No samples taken and no bottle file was created.

96. Two bottles closed. No samples taken and no bottle file was created.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-11

Dates: Start: May 26, 2003 End: June 16, 2003

Location: North-West Pacific

Vessel: Tully

Party Chief: Robert M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |
| 2 | SEABIRD | 911+ | 0550 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2003-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 03/05/02 | Factory | 16/07/03 | Factory |
| Conductivity | 1763 | 23/04/02 | “ | 24/07/03 | “ |
| Secondary Temp. | 2106 | 19/04/02 | “ | 05/09/03 | “ |
| Secondary Cond. | 1764 | 23/04/02 | “ | 05/09/03 | “ |
| Transmissometer | 333DR | 11/10/02 | IOS |  |  |
| SBE Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2229 | July01 | IOS |  |  |
| PAR | 4565 | 22/07/02 | ? |  |  |
| Pressure Sensor | 75636 | 30/09/99 | Factory |  |  |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550

Cruise ID#:		2003-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 23/01/03 | Factory | 05/09/03 |  |
| Conductivity | 2102 | 20/06/02 | “ | 05/09/03 |  |
| Secondary Temp. | 2710 | 23/01/03 | “ |  |  |
| Secondary Cond. | 2278 | 23/01/03 | “ |  |  |
| Transmissometer | 333DR | 11/10/02 | IOS |  |  |
| SBE Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2229 | July01 | IOS |  |  |
| Pressure Sensor | 75636 | 06/04/99 | Factory |  |  |

February 25, 2004

Nina Nemcek found some errors in the chlorophyll data and reported them to Marie Robert. I made the corrections to cast numbers 71 and 74 in the CHE files and to the loop file. Previous versions were copied to the history directory.

Nina’s email to Marie:

From: Nina Nemcek [la_nina_@hotmail.com]

Sent: January 9, 2004 12:14

To: RobertM@pac.dfo-mpo.gc.ca

Subject: Tully 2003-11 chl a errors!

Hi again Marie,

I just wanted to let you know that I found some errors in the chlorophyll a data that Phil and I did. They were all cut and paste formula errors but have been fixed now.

Affected samples with revised numbers are as follows:

Argo 5 75 m chl a = 0.175 phae= 0.179

Argo 6 100 m chl a = 0.061 phae= 0.114

75 m chl a= 0.168 phae= 0.176

10 m chl a= 0.591  phae= 0.188

0 m chl a = 0.835  phae= 0.243

New 7 chl a= 0.751  phae= 0.226

New 8 chl a = 0.338  phae=0.151

New 9 chl a= 0.702  phae=0.275

Sorry for the mistake, hope it doesn't cause any serious problems.

Nina

October 7, 2005

Based on the recalculation of titrated DO by Sheila Toews, the bottle DO was compared with the previously processed SBE DO channel in the bottle files. When outliers are ignored and the trendline is forced through the origin, the results indicate that the SBE DO values are high by about 2%. A recalibration was applied to the SAMCOR2 files multiplying the SBE DO channel by 0.98 and the comparison rerun. The results were satisfactory, so the same recalibration was applied to the CTD and CHE files using file 2003-11-recal4.ccf.