## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 2 Aug 2005 | Added loop data to the archive. J.L. |
| 22 Dec. 2004 | Recalibrated salinity based on post-cruise calibration; see note below. |

## PROCESSING NOTES

Cruise: 2003-13

Agency: OSAP

Location: N.W.Pacific

Project: High Seas Salmon

Party Chief: Morris J.

Platform: W. E. RICKER

Date: June 8, 2003 – June 12, 2003

Processed by: Germaine Gatien

Date of Processing: 29 September 2003 – 1 October 2003

Number of original CTD casts: 21

Number of casts processed: 21

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit S/N was a SeaBird model 11 (#0471).

# SUMMARY OF QUALITY AND CONCERNS

The files had non-standard names.

The salinity should be considered ±0.002psu. There was only one deep bottle calibration sample.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Salinity and titrated chlorophyll data were obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. A small error in the calibration constants for the pressure sensor was corrected and the resulting file named 2003-13CTD.con.

The sensor history was found. The conductivity sensors were last used for cruise 2003-38 but with different temperature sensors.

##### Conversion of Raw Data

The raw data was converted using configuration file 2003-13-ctd.con.

A few casts were checked to ensure that all expected channels contained reasonable data.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### CELLTM

Tests were run on two casts running CELLTM with choices of (0.02,9), (0.03,9), (0.02,7), (0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The best choice was found to be (0.02,7) for the primary and (0.03,9) for the secondary channel; CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The primary sensors are generally noisier than the secondary. It looks as though the flow may have been irregular, perhaps due to plumbing problems. The casts were so shallow that looking at differences between sensors was not useful – the variability was so great that little could be deduced except that the primary salinity is lower than the secondary by about 0.006 to 0.008 units at about 250db. The transmissivity looks good with maxima around 90% and up and downcast similar for the deepest cast. The fluorescence looks reasonable with a dark value of about 0.03ml/l and a vertical offset of about 2db.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced.

The times in the log differ from those in the files by about 2 hours. Work on the RICKER follows a 7am to 7pm pattern, so a check can be made by looking at the times at which work started and stopped. The log and UPLOAD are both supposedly in UTC, so 7 hours was subtracted from all times to see how they fit with the normal activities of this ship.

- On June 9th the first task of the day was at 7:04 PDT according to the log. The UPLOAD time is 8:57 PDT. The log time makes sense if the ship was on station before the start of the work day, but it is unknown if that is true.

- On June 9th the final task was a trawl that started at about 7pm PDT according to the log. This was about a half hour after a CTD cast which according to the UPLOAD time started at 8:19pm PDT. The UPLOAD time is after work on the RICKER would have ceased.

- On June 10th the first activity of the day was a CTD cast at 7:07 PDT according to the log and 9:00 PDT according to the UPLOAD entry. The 7am start makes sense; this cast was in the middle of a line of stations and the ship would surely have been on station in good time.

- On June 10th the last activity was a trawl at about 17:22 PDT according to the log, which is a little early, but this was the end of a sampling line that was followed by a long steam.

- On June 11th work started with a CTD cast at 6:56 PDT according to the log and 8:50 PDT according to the UPLOAD entry. The UPLOAD time is too late.

It is clear that the UPLOAD time was out by about 1 hour and 53 minutes. A closer look at the files shows that while the UPLOAD time is wrong, the NMEA time is ok. A change was made to the conversion program so that it chooses the NMEA time when it is available instead of the UPLOAD time. Hugh MacLean was notified that the computer time should be checked on the RICKER CTD.

The data was reconverted using the new routine. This time no errors were found in the headers.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 4db.

T0, T1, S0 and S1 were plotted for all casts. The primary and secondary are reasonably close. The secondary upcasts look like the downcasts, but there are significant differences in the primary, especially in the high gradient regions. The transmissivity and fluorescence are fairly noisy at depth. The fluorometer cable was changed after cast #34 because the data was particularly noisy at the bottom of that cast; however, the transmissivity also looks bad for that cast so the problem may have been of a more general nature. The transmissivity moves suddenly to low values at about 35db on the upcast of cast #52, suggesting a close encounter with something.

##### SHIFT

Conductivity

Tests were run using advancements of +0.5, 0 and -0.5 records. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the secondary conductivity of +0.5 records. Further tests were then tried for settings between 0 and 1.5, narrowing down the best choice to about +0.8 records. This implies an advancement of about +0.019s. This setting would make sense if the secondary conductivity was already advanced by the deck unit. Normally deck units of the 400 series do not advance the secondary sensor so we expect a higher advancement than this. Perhaps the deck unit serial number is wrong in the log book or the unit has been adapted to advance both channels. Whatever the reason the shift of +0.8 records does produce the best results. No study was made of the primary conductivity since the data looks very noisy and is unlikely to be used.

All data was put through SHIFT using +0.008 records for the secondary conductivity.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for 2 casts were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. As has been found in the past a choice of 1s looks reasonable although the descent rate was very noisy and the offsets highly variable; a shift of +24 records was applied.

##### BOTTLE FILE PREPARATION

The SAL and CHL spreadsheets were converted to separate SAL and CHL files. Two comments from the CHL spreadsheet were copied as comments in the headers of the relevant files.

The cast # for the deep salinity sample from cast #34A was changed from 934 to 34 so that the two samples taken at that station are entered in one file. In fact the shallow sample was probably taken from a second shallow cast, but it was not given a separate event number.

Salinity comparison

There was only one deep salinity sample taken during cast #46 while the CTD was at the bottom. The bottle was about 5db above the CTD. The CTD moved up and down in the water during the stop so the estimated depth of the bottle is very rough – between 245 and 247db.

Looking at the differences between the bottle salinity and the upcast and downcast salinity from the two sensors indicates that the secondary sensors are high by 0.0017 or 0.0019psu according to the downcast data. They are closer to the bottles if we use the upcast data, but that is likely to be less reliable. The primary sensors give salinity that is low by from 0.001 to 0.008psu.

There were samples for most casts from around 10db, but for the most part the water was not sufficiently well-mixed for these to be useful for calibration purposes. Cast #22 is quite well mixed at that depth. Using upcast data from about 9.5db the primary salinity appears to be low by 0.008psu and the secondary high by 0.0011psu.

Fluorescence-chlorophyll comparison

For comparison with the extracted chlorophyll the shifted files were put through REVERSE and DELETE. These upcast files were then averaged in 0.5db-bins and thinned to values from 9db to 16db. The number of points in the bins was used as a guide to determine the pressure at which the CTD stopped and the FL value was taken from the appropriate bin 5 db above. These values were compared with the extracted chlorophyll and the results overall show a relationship of

CHL = 0.925 * FL + 1.185

For 2003-12 the result was CHL = 0.9 * FL - 0.03.

However, the data from 2003-13 has a lot of scatter and the trendline is probably not significant. When the ratio of FL to CHL is compared with distance from shore a pattern emerges. Those casts close to shore, say within 25km roughly, have a lower ratio than those well away from the shore. Since all these samples are from about 9.5db the gradient at that depth was examined to see where it was largest and most likely to lead to errors associated with uncertainties about the actual depth of sampling. Those errors are largest for the casts farthest offshore and there are 2 records that look like outliers far offshore. A quick check was made to see if using downcast fluorescence data made a difference. The downcast results are slightly different with the trendline:

CHL = 1.0903 * FL + 0.6687

The offset is much smaller, but the observation of a dependence on distance offshore remains. For the inshore casts the low ratios are not related to errors in depth. All sampling was in daylight hours as the RICKER works between 7am and 7pm. (See 2003-13-chl_comp.xls.)

Similar observations have been made for other cruises, notably 2003-11 and 2003-11, in which the casts in Queen Charlotte Sound and Knight Inlet showed fluorescence significantly lower than chlorophyll.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values and to fix the event numbers.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Surface Salinity:  5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warning referred to upcast data.

##### DETAILED EDITING

The secondary sensors appear to be closer to the bottles and are less noisy than the primary sensors.

Page plots were produced using T1,S1. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data was removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case. In a few cases the descent rate was examined to determine if unstable features should be removed or not.

Cast #37 required no editing. All other casts required light editing.

The descent rate was kept high for the most, so that while it was very unsteady at times, corruption by shed wakes was minimized.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that does not require editing.

##### Recalibration

The salinity is believed to be within ±0.002psu. The calibration information is insufficient to justify recalibration.

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the EDT files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. In all cases both sensors gave salinity that was too high. The best comparison was in October 2002 when there were surface bottles from well-mixed water. The primary salinity was found to be high by 0.004psu and the secondary salinity high by 0.0017psu; however, both temperature sensors were different for those cruises so this information is of limited use.

Historic ranges – The only excursions from the historic ranges were for salinity between 30 and 60db for 3 casts well offshore. The low salinity fits other observations from this area that salinity was lower than usual at those depths.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment and fix formats and channel names using file 2003-13-header.hdr. The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

34. After this cast the fluorometer cable was changed due to noise at depth. The gain is the same.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-15

Dates: Start: June 8, 2003 End: June 12, 2003

Location: North-West Pacific

Vessel: Ricker

Party Chief: Morris J.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2003-15

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 22/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: Recalibrated CTD files using file 2003-13-recal2.ccf to raise salinity by 0.0083 based on post-cruise calibration. It was assumed that the drift was linear with time. G. Gatien