Report on sensors #2023 (Temperature) and 1763 (Conductivity)

SeaBird report drifts of:		 +0.00264 Cº/year in sensor #2023

-0.0004 PSU/month in sensor # 1763 at conductivity = 3 S/m

The drift in temperature is relatively large and I think we should consider recalibrating the data in the archive. We could assume the drift is linear and base the recalibration on time since the previous calibration, or base it on use (number of casts since previous calibration / total casts between calibrations * drift). But there is other information available that may allow a better method of determining when the changes occurred.

I have kept a record of differences between sensors at a few test casts for each of the cruises and a record of comparisons with bottle salinity. Here is a table showing the results of all cruises that have been processed using these sensors since the previous calibration:

| Cruise | Depth | Date | T1-T0 | C1-C0 | S1-S0 | S0-Sbot | Comments |
| --- | --- | --- | --- | --- | --- | --- | --- |
| 2002-16 | Deep | July 2002 | -0.0027 | +0.0001 | +0.005 | +0.0015 |  |
| 2002-28 | Shallow | Aug. 2002 | -0.004 | ~0 very noisy | +0.0045 | -0.0011 |  |
| 2002-31 | Shallow | Sept. 2002 | -0.004 | -0.0001 | +0.003 | +0.0008 |  |
| 2002-39 | Shallow | Oct. 2002 | -0.003 | ~0 | +0.0025 | +0.0004 |  |
| 2002-40 | Shallow | Nov. 2002 | -0.003 | ~-0.00005 | +0.0025 | ~-0.001 | Only 4 bottles |
| 2003-08 | Shallow | Mar. 2003 | -0.0027 | -0.0012 | -0.0012 | +0.0015 | Very few bottles |
| 2003-12 | Some deep | May 2003 | -0.004 | ~0.0001 | +0.005 | -0.0056 |  |
| 2003-15 | Some deep | May 2003 | -0.004 | +0.0002 | +0.006 | -0.0072 | Before temp. recal |

Note that the differences between pairs of sensors are very noisy, especially for conductivity and salinity and especially for shallow casts. Minor differences in flow rates cause a lot of noise. We also should not assume that there was no change in the secondary sensors with which we are comparing. At this time there is no report on those.

Looking at the table it looks like most of the drift in temperature happened very early. For salinity it looks like there was little drift until after the March 2003 cruise.

As of the 16th of July, 2003, the drift in salinity due to the conductivity sensor was -0.0058 PSU at conductivity = 3 s/m. Given the higher conductivities measured the observed errors will be larger with values on the order of -0.007 PSU near the surface and about -0.0065 PSU at depth. The temperature error would have an effect in the opposite direction, with salinity errors of from +0.002 near the surface to +0.003 at depth.  For 2003-15 I found errors before correcting the temperature of -0.0072 and after correcting temperature they were -0.0044 PSU using bottles below 500db. Given the scatter in the bottle data this is remarkably good agreement.

For 2003-12 I found an error in salinity of -0.0056 with uncorrected temperature. Again there is a lot of scatter in the bottle data. The recalibration done already looks adequate.

There is no need to recalibrate the salinity in the archive since the bulk of the drift seems only to have affected 2003-12 and 2003-15 which have been (or will be) recalibrated appropriately.

I would recommend that we subtract 0.0027 from all the temperature data from these cruises. Because the sensor has been used so extensively, the systematic error might cause some problems in future. If the temperature data is recalibrated it is important to NOT recalculate the salinity since the bottle comparisons were the basis of recalibration decisions for those.