## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2003-17

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Project: Sea Lice Sampling

Area: Broughton Archipelago

Chief Scientist: Galbraith, M.

Platform: Manyberries

Date: 29 April 2003– 1 May 2003

Processed by: Germaine Gatien

Date of Processing: 8 August 2003 – 11 August 2003

Number of original CTD casts: 12 hex files

Number of casts processed: 11

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD  (S/N#1294) was used.

SUMMARY OF QUALITY AND CONCERNS

The file names were non-standard.

There was no salinity calibration sampling and no history for this instrument since its last calibration.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. Salinity for this cruise has been edited extensively, but errors as high as 0.05units of salinity are likely.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The Daily Log sheets were obtained.

There was no salinity calibration data.

The data files have non-standard names but contain the station ID from the log sheets so were easily understood. The names were changed to standard format.

Con file 1294-Nov22.con contains the most up-to-date calibration coefficients for this instrument.

- Conversion of Raw Data

- The raw data was converted using conversion file 1294-Nov22.con.

- Cast #1 could not be converted because the line “END” was missing from the header. Attempts were made to repair the file, but were unsuccessful.

- The positions and station names will be added to the headers when the files have been converted to IOS Headers.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. DERIVE

Program DERIVE was run to calculate salinity.

6.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

CLEAN was run to add event numbers and a text editor was used to add latitude, longitude and bottom depths to the headers.

7. Checking Headers

A Cruise Track Plot was produced and a Header Check and Header Summary were run and details compared with the log book. The times in the headers are in PST and in the log in PDT.

The surface check was run. The average surface pressure is –0.2db but the CTD was turned on in air. Examining files closely suggests that the conductivity went to non-zero values at about 0db. No adjustment will be made.

8. Test Plots

Profiles were plotted for all casts and up and down traces compared. There were no differences so great as to suggest errors in processing.

9. SHIFT

In order to align the temperature with pressure SHIFT will be used. In previous use of this equipment the temperature was advanced by 0.7s using the SeaBird routine ALIGNCTD. That is equivalent to 1.4 records. Attempts were made to test various settings, but the results are confusing. Using +1.4 records looks reasonably good on a T-S surface. But when comparing up and downcast temperature traces versus up and downcast conductivity traces settings on the order of -1.5 look better. However, the conductivity looks suspiciously different between upcast and downcast, so that may be the problem. The casts for this cruise are all very shallow so it is hard to find a section with interesting features and a steady descent rate. It seems wisest to go with the results of past cruises. The temperature will be shifted by +1.4 records for all cast.

11. DELETE

DELETE was run twice. Because many of the downcasts looked suspicious, the SHF casts were put through DELETE and then through REVERSE followed by DELETE with outputs DEL and DELREV.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 5

Swells deleted. Warning message if pressure difference of 2.00

Drop rates feature was turned OFF

There were no warnings.

11. CTDEDIT

Both upcast and downcast page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT. Each cast was examined to decide whether the upcast or downcast should be used. In general the descent rate was low but steady for the upcasts but very noisy for the downcast. So the upcasts were selected for all casts. The top 0.5db and some bottom records were removed from all casts. There were some unstable features in areas where the temperature changed smoothly but the salinity did not. Most of the instabilities appear to be due to the mismatch of temperature and conductivity time response; Sea-Bird advise that this problem is most serious when the descent rate is uneven. The resulting errors are seen as salinity spikes particularly in areas of high T and C gradients. CTDEDIT was used to clean salinity in such cases.

Note was made of the editing details in the relevant files

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value =  .000

Average value will be used.

Interpolated values are NOT used for empty bins

Page plots were examined. Even after graphical editing and metre-averaging there remain instabilities. Given the region of the study is one of active mixing it is impossible to establish if these are real or not. No further editing will be done.

- Intercomparisons

COMPARE - There was no salinity sampling.

Previous Use of CTD – This is the first processing of data using this instrument since its latest calibration. Before that calibration it was used for 2000-11 and 2001-27. There were no useful bottles from 2000-11. During 2001-27 the CTD salinity was found to be low by about 0.001units which was within the noise level.

Historic ranges – There is no local climatology available.

# Comparison of nearby casts

Plots were overlaid of T-S curves for casts #7 and #8, which were at the same site but 5 hours apart, to check repeatability. The differences along lines of constant σt are minimal below 5db.

14. Final Plots

THIN and DERIVED QUANTITIES routines were used to produce information for tables. Page plots were prepared using the edited data and tables.

15. REMOVE, FIX TIME and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity and Flag.

The ADD TIME CHANNEL routine was used to add 8 hours to the times in the headers.

The HEADEDIT routine was used to fix formats and channel names and to add the following comment to the headers:

There was no salinity calibration sampling and no history for this instrument since its last calibration, so no estimate can be made of those errors.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. Salinity for this cruise has been edited lightly, but errors as high as 0.05units of salinity are likely in areas of high gradient and unsteady descent rate.

16. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Particulars

1. Unable to convert file.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2003-17

Dates: 	Start:	29 April 2003	    End:	1 May 2003

Location: Broughton Archipelago

Vessel:	Manyberries

Party Chief:       Galbraith, M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294

Cruise ID#:		2003-17

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 22 Nov. 2003 | Factory |  |  |
| Conductivity | 1294 | 22 Nov. 2003 | Factory |  |  |
| Pressure Sensor | 163223 | 17 June 1999 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.