## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

PROCESSING NOTES

Cruise: 2003-23

Agency: IOS, Ocean Science and Productivity, Sidney, B.C.

Project: Sea Lice Sampling

Area: Broughton Archipelago

Chief Scientist: Galbraith, M.

Platform: Manyberries

Date: 13 May 2003– 14 May 2003

Processed by: Germaine Gatien

Date of Processing: 8 August 2003 – 12 August 2003

Number of original CTD casts: 12 hex files

Number of casts processed: 12

INSTRUMENT SUMMARY

A Sea Bird Model SBE 19 SEACAT CTD (S/N#1294) was used.

SUMMARY OF QUALITY AND CONCERNS

There was no salinity calibration sampling and no history for this instrument since its last calibration.

All the casts are very shallow with a low descent rate and the downcast traces look suspicious. Since the ascent rate was generally higher and smoother and the data more believable, the upcast data was selected for the archive.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature response times especially when the descent rate is non-uniform. These errors are likely to be much larger than calibration errors. Salinity for this cruise has been edited, but errors as high as 0.05units of salinity likely remain.

PROCESSING SUMMARY

1. Seasave - This step was completed at sea; the raw data files are *.hex.

2. Preliminary Steps

The Daily Log sheets were obtained.

There was no salinity calibration data.

Con file 1294-Nov22.con contains the most up-to-date calibration coefficients for this instrument.

- Conversion of Raw Data

- The raw data was converted using conversion file 1294-Nov22.con.

- The positions and station names will be added to the headers when the files have been converted to IOS Headers.

- FILTER

The conductivity was low-pass filtered with a time constant of 0.5 seconds to force it to have the same response as the temperature.

The pressure was filtered with a time constant of 2 seconds to increase the pressure resolution.

5. DERIVE

Program DERIVE was run to calculate salinity.

6.  Conversion to IOS Headers

The IOSSHELL routine for Sea Bird ASCII files was used to convert the Sea-Bird data to IOS Headers.

CLEAN was run to add event numbers and a text editor was used to add latitude, longitude and bottom depths to the headers.

7. Checking Headers

A Cruise Track Plot was produced and a Header Check and Header Summary were run and details compared with the log book. The times in the headers are in PST and in the log in PDT.

The surface check was run. The average surface pressure is –0.3db but the CTD was turned on in air. Examining files closely suggests that the conductivity went to non-zero values at about 0db. No adjustment will be made.

8. Test Plots

Profiles were plotted for all casts and up and down traces compared. There were no differences so great as to suggest errors in processing.

9. SHIFT

The temperature was shifted by +1.4 records for all casts as was done for 2003-17 using the same equipment.

11. DELETE

DELETE was run twice.

During 2003-17 the upcasts were found to be more useful than the downcasts. For this cruise the descent rate was very noisy and very low. The upcasts were more regular. The SHF casts were put through DELETE and then through REVERSE followed by DELETE with outputs DEL and DELREV.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Swell Pressure Tolerance: 1.0

Pressure filtered over width: 5

Swells deleted. Warning message if pressure difference of 2.00

Drop rates feature was turned OFF

There were no warnings.

11. CTDEDIT

Both upcast and downcast page plots were produced and were examined for spikes and instabilities to guide the use of CTDEDIT. Each cast was examined to decide whether the upcast or downcast should be used. In general the descent rate was low but steady for the upcasts but very noisy for the downcast. Just as was done for 2003-17, the upcasts were selected for all casts. The top 0.5db and some bottom records were removed from all casts. A few other records were removed from some. There were some unstable features in areas where the temperature changed smoothly but the salinity did not. Most of the instabilities appear to be due to the mismatch of temperature and conductivity time response; Sea-Bird advise that this problem is most serious when the descent rate is uneven. The resulting errors are seen as small salinity spikes (or lumps) particularly in areas of high T and C gradients. CTDEDIT was used to clean salinity in such cases.

Note was made of the editing details in the relevant files.

12. BIN AVERAGE

The following Bin Average values were used:

Bin channel = pressure

Averaging interval = 1.000            Minimum bin value =  .000

Average value will be used.

Interpolated values are NOT used for empty bins

Page plots were examined. Even after graphical editing and metre-averaging there remain instabilities. Given the region of the study is one of active mixing and the casts extremely shallow, it is quite possible that some instabilities are real. No further editing will be done.

- Intercomparisons

COMPARE - There was no salinity sampling.

Previous Use of CTD – The only data processed using this instrument since its latest calibration was from 2003-17 for which there was no salinity calibration sampling. Before the last calibration it was used for 2000-11 and 2001-27. There were no useful bottles from 2000-11. During 2001-27 the CTD salinity was found to be low by about 0.001units which was within the noise level.

Historic ranges – There is no local climatology available.

# Comparison of nearby casts

Plots were overlaid of T-S curves for nearby casts but they are so shallow it is difficult to conclude anything about the instrument. For 2003-17 using the same equipment differences along lines of constant σt were minimal below 5db for two casts at the same site.

14. Final Plots

THIN and DERIVED QUANTITIES routines were used to produce information for tables. Page plots were prepared using the edited data and tables.

15. REMOVE, FIX TIME and HEADEDIT

The following channels were removed from all casts: Scan_Number, Conductivity and Flag.

The ADD TIME CHANNEL routine was used to add 8 hours to the times in the headers.

The HEADEDIT routine was used to fix formats and channel names and to add the following comment to the headers:

There was no salinity calibration sampling and no history for this instrument

since its last calibration, so no estimate can be made of those errors.

Salinity for this cruise has been edited lightly but large errors likely remain.

SeaCat salinity is prone to error due to mismatch of conductivity and temperature

response times, especially when the descent rate is non-uniform. These errors are

likely to be much larger than calibration errors. SeaCat salinity errors are

expected to be as high as 0.05units in areas of high gradients. Since all casts

for this cruise are from upcast data and the ascent rate was very low the error

is probably significantly higher than that.

16. Producing final files

a.) The final files were renamed *.ctd.

b.) A cross-reference listing was produced.

Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#:	2003-23

Dates: 	Start:	12 May 2003	    End:	14 May 2003

Location: Broughton Archipelago

Vessel:	Manyberries

Party Chief:       Galbraith, M.

| CTD# | Make | Model | Serial # | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | SBE19 SEACAT | 1294 | No | Yes |

Institute of Ocean Sciences

CTD Calibration Information

Make/Model/Serial#:	SEABIRD/19 SEACAT / 1294

Cruise ID#:		2003-23

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 1294 | 22 Nov. 2003 | Factory |  |  |
| Conductivity | 1294 | 22 Nov. 2003 | Factory |  |  |
| Pressure Sensor | 163223 | 17 June 1999 | Factory |  |  |

Sensor Calibration Notes:

The configuration file used is attached; this includes the sensor calibrations.