## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| Dec. 22, 2004 | Recalibrated CTD files based on post-cruise calibration; see note below. |

## PROCESSING NOTES

Cruise: 2003-25

Agency: Stock Assessment

Location: Hecate Strait/Queen Charlotte Sound

Project: Hecate Strait Multi-species Survey

Party Chief: Workman, G.

Platform: W. E. RICKER

Date: May 22, 2003 – June 4, 2003

Processed by: Germaine Gatien

Date of Processing: 8 March 2004 – 22 March 2004

Number of original CTD casts: 42

Number of casts processed: 42

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit is unknown, but was probably the one used on the next Ricker cruise, a SeaBird model 11 (S/N 0471).

# SUMMARY OF QUALITY AND CONCERNS

The files had non-standard names.

There was no salinity calibration sampling during this cruise and the recalibration is based on other cruises using the same equipment; salinity should be considered ±0.005psu.

There was no list of equipment in the CTD Daily log book.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The CTD Log Book and a spreadsheet of positions and times was found.

The cruise summary sheet was completed.

The CTD # is presumed to have been #0506 since that is the usual one on Ricker cruises.

The configuration files were obtained and the calibration constants were checked. A small error in the calibration constants for the pressure sensor was corrected and the resulting file named 2003-25CTD.con.

The sensor history was found. All temperature and conductivity sensors were last used for cruise 2002-38 and were used immediately afterwards during 2003-13.

##### Conversion of Raw Data

The raw data were converted using configuration file 2003-25-ctd.con.

A few casts were checked and all channels contained reasonable data.

There were no rosette casts.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

Since the deck unit was one of the older versions that does not advance the secondary conductivity, all casts were put through ALIGNCTD to advance the secondary conductivity by +0.073 so that it matches the primary. Fine-tuning of the alignment will be done later using SHIFT.

##### CELLTM

During 2003-13 using the same equipment the best choice was found to be (0.02, 7) for the primary and (0.03, 9) for the secondary channel and these were also the choices made for the last cruise before this one. Tests were done on cast #1 to ensure that these settings were appropriate and they are.

CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The primary sensors are a little noisier than the secondary. None of these casts were very deep and most extremely shallow, so the differences are expected to be highly variable. The results for #38 and #39, for which the descent rate was fairly steady, suggest no great differences between sensors, although there is quite a lot of depth dependence and a probable time-dependence.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 5 | 250 | -0.0006 | +0.00019 | +0.0026 | Very noisy |
| 6 | 215 | -0.0008 | +0.00016 | +0.0025 | Very noisy |
| 38 | 220 | -0.0006 | -0.00005 | +0.0005 | Steady |
| 39 | 230 | -0.0007 | -0.00001 | +0.0008 | Steady |

The transmissivity and fluorescence look reasonable.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Two casts had errors in the headers that had to be corrected before they could be converted.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. There are small discrepancies between the log and header positions and times. The worst is for cast #1 with a 17 minute difference. Since, according to G. Workman, the log reflects the time of arrival on station rather than the start of the CTD cast, the GPS time is considered more reliable and no changes were made. There was one error in station name which was corrected.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 2.7db. There are a few casts with very low conductivity in the top 1db, but the pumps weren’t on yet and the temperature does not change much when the conductivity does, so these are probably not out-of-water values; even after the pump came on, the two conductivity channels differ markedly for some time. Perhaps there was a problem with the pump, or the cell had dried out.

T0, T1, S0 and S1 were plotted for all casts. There are severe problems with the secondary salinity data, mostly during the downcasts. The secondary temperature is occasionally bad as well (ex. cast #8). The upcast for cast #8 was much worse than the downcast. The primary sensors will have to be used.

##### SHIFT

Fluorescence

There was no note in the CTD log about whether the fluorometer was pumped or not. A check of a few casts during which the descent rate was fairly noisy shows no sign of the data reflecting the descent rate; that suggests that it was pumped. That seems likely since the CTD was probably set up in the usual way.

To find what shift is needed for the fluorescence, upcast and downcast profiles for a few casts were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The results were highly variable, with shifts of from 0 to 10s. The area is one of active mixing with great variability in T and Fl, so the much of the difference between up and downcast is likely due to sampling different water. The usual 1s (24 records) shift will be applied.

Conductivity

Initial tests were run using advancements of +0.5, 0 and -0.5 records on cast #40. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the secondary conductivity of -0.5 records. Further tests were then tried for settings between -0.3 and -1.4 records, narrowing down the best choice to about -1.1 records, or -0.026s. This implies a net advancement of about +0.047s. SHIFT was run on several other casts using -1.1 records and that setting proved satisfactory. The same value was used during 2003-36 when the same equipment was used.

All data were put through SHIFT -1.1 for the primary; no shift was applied to the secondary since the data is too poor to use.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Surface Salinity:  5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The were no warnings.

##### DETAILED EDITING

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case. The noise in the primary salinity data was much greater than usual, and is particularly striking because the range of salinity is low, so the effects of the noise are large. Hence, detailed editing was frequently necessary. Problems have been noted with the equipment before, leading to the choice of secondary salinity from other cruises, but that was not possible this time. Attempts were made to filter the salinity then rerun Delete, but the results were poor; careful editing works better. For cast #35 the upcast was considered, but was just as bad as the downcast.

All casts required some editing. The following casts required heavy editing of salinity: 35, 40, 42

The descent rate was generally high and steady with a few exceptions.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Comparisons

Previous experience with these sensors – The conductivity sensors have been used many times since the last recalibration although not always with the same temperature sensors. There has never been much deep salinity sampling and generally there was a lot of scatter in the comparisons. The primary salinity was found to be low by 0.006psu and 0.007psu in June and October of 2003. However the calibration sampling was very limited and little confidence is placed in those results. For the preceding cruise, 2002-38 the salinity was high by about 0.004psu, but the temperature sensor was a different one. There were some hints of a drift with time during 2002-38, and reviewing all cruises on which this sensor has been used there certainly seems to have been a drift.

Historic ranges – All casts were examined with the local climatology superimposed. All salinity data fell within the ranges except for the top 35db of cast 20 which has low salinity; that cast was very close to shore. Much of the temperature data fell above the historic maxima near the bottom of casts; most of these were casts near shore and may reflect that such areas were not sampled in the cruises that went into the climatology. There is no evidence of instrumental problems. The deeper casts all fell within the historic ranges.

##### Recalibration

For the cruises that preceded and followed this one the secondary sensor was used because it produced salinity closer to the bottles. The data will be recalibrated by adding 0.005psu to the primary salinity, based primarily on the results of 2003-13 but allowing for a little drift with time as suggested by 2003-38. There is some suggestion from the differences described in section 8 that the calibration may have changed between casts #6 and #38 but there is certainly not enough information to justify appling a time-dependent recalibration.

The results should be considered ±0.005psu.

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and two header entries and to add the following comment:

There were no salinity calibration samples for this cruise. The

recalibration of salinity is based on the results of other cruises

using the same equipment; the calibration sampling for those cruises

was limited and there is evidence of drift in the calibration.

There was a lot of fine-scale noise in the salinity channel.

Salinity should be considered +/- 0.005psu.

Transmissivity - The data are nominal and unedited.

Fluorescence - The data are nominal and unedited.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-25

Dates: Start: May 22, 2003 End: June 4, 2003

Location: Hecate Strait

Vessel: W.E. Ricker

Party Chief: Workman G.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2003-25

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 22/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: All CTD files were recalibrated using file 2003-25-recal2.ccf to raise the primary salinity by 0.0015psu based on a post-cruise calibration of the conductivity sensor. It was assumed that the drift was linear with time. G. Gatien