## File List
Number of Files= |  | 20
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2003-26-0008 | 2003-26-0008 | 11 | 21 | 2003-26-0008 & 2003-26-0008
2 | 2003-26-0010 | 2003-26-0010 | 23 | 28 | 2003-26-0010 & 2003-26-0010
3 | 2003-26-0012 | 2003-26-0012 | 30 | 38 | 2003-26-0012 & 2003-26-0012
4 | 2003-26-0014 | 2003-26-0014 | 40 | 45 | 2003-26-0014 & 2003-26-0014
5 | 2003-26-0018 | 2003-26-0018 | 47 | 51 | 2003-26-0018 & 2003-26-0018
6 | 2003-26-0019 | 2003-26-0019 | 53 | 58 | 2003-26-0019 & 2003-26-0019
7 | 2003-26-0020 | 2003-26-0020 | 60 | 65 | 2003-26-0020 & 2003-26-0020
8 | 2003-26-0021 | 2003-26-0021 | 67 | 71 | 2003-26-0021 & 2003-26-0021
9 | 2003-26-0022 | 2003-26-0022 | 73 | 77 | 2003-26-0022 & 2003-26-0022
10 | 2003-26-0023 | 2003-26-0023 | 79 | 84 | 2003-26-0023 & 2003-26-0023
11 | 2003-26-0024 | 2003-26-0024 | 86 | 93 | 2003-26-0024 & 2003-26-0024
12 | 2003-26-0025 | 2003-26-0025 | 95 | 102 | 2003-26-0025 & 2003-26-0025
13 | 2003-26-0026 | 2003-26-0026 | 104 | 111 | 2003-26-0026 & 2003-26-0026
14 | 2003-26-0027 | 2003-26-0027 | 113 | 119 | 2003-26-0027 & 2003-26-0027
15 | 2003-26-0028 | 2003-26-0028 | 121 | 126 | 2003-26-0028 & 2003-26-0028
16 | 2003-26-0029 | 2003-26-0029 | 128 | 132 | 2003-26-0029 & 2003-26-0029
17 | 2003-26-0030 | 2003-26-0030 | 134 | 140 | 2003-26-0030 & 2003-26-0030
18 | 2003-26-0031 | 2003-26-0031 | 142 | 148 | 2003-26-0031 & 2003-26-0031
19 | 2003-26-0032 | 2003-26-0032 | 150 | 156 | 2003-26-0032 & 2003-26-0032
20 | 2003-26-0033 | 2003-26-0033 | 158 | 164 | 2003-26-0033 & 2003-26-0033

## (DOX-CTD - DOX-BOT)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.17835 |  | X | CALC Y
m= | -0.0014344 | 0.00030336 |  | Std Dev of Resid |  | 0.16109 |  | 5.44 | 0.36355
b= | 0.37135 | 0.02477 |  | Resid Sum of Squares |  | 2.6728 |  | 302.405 | -0.062415
# of Values Extracted = |  |  | 135
# of Values Used in Fit = |  |  | 105
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 302.405 | 1.391 | 0.035355 | -0.003 | -0.062415 | -0.059415 |  | 302.405 | -0.003
1 | YES | 250.01 | 1.721 | 0 | 0.164 | 0.012739 | -0.15126 |  | 250.01 | 0.164
1 | YES | 199.98 | 2.328 | 0 | 0.144 | 0.084501 | -0.059499 |  | 199.98 | 0.144
1 | YES | 150 | 2.993 | 0 | 0.186 | 0.15619 | -0.029808 |  | 150 | 0.186
1 | YES | 99.98 | 4.374 | 0 | 0.091 | 0.22794 | 0.13694 |  | 99.98 | 0.091
1 | YES | 74.99 | 4.969 | 0 | 0.053 | 0.26378 | 0.21078 |  | 74.99 | 0.053
1 | YES | 50.01 | 5.614 | 0 | -0.104 | 0.29962 | 0.40362 |  | 50.01 | -0.104
1 | YES | 32.485 | 5.6165 | 0.1096 | 0.1785 | 0.32475 | 0.14625 |  | 32.485 | 0.1785
1 | NO | 22.505 | 5.654 | 0.35638 | -0.595 | 0.33907 | 0.93407 |  | 12.475 | -0.049
1 | YES | 12.475 | 6.286 | 0.045255 | -0.049 | 0.35345 | 0.40245 |  | 134.8 | 0.124
1 | NO | 4.97 | 6.198 | 0 | -0.129 | 0.36422 | 0.49322 |  | 100 | 0.31
2 | YES | 134.8 | 1.62 | 0 | 0.124 | 0.17799 | 0.053994 |  | 75.01 | 0.083
2 | YES | 100 | 2.131 | 0 | 0.31 | 0.22791 | -0.082089 |  | 49.99 | 0.108
2 | YES | 75.01 | 3.306 | 0 | 0.083 | 0.26376 | 0.18076 |  | 32.465 | 0.0375
2 | YES | 49.99 | 4.048 | 0 | 0.108 | 0.29964 | 0.19164 |  | 176.375 | 0.032
2 | YES | 32.465 | 4.1255 | 0.030405 | 0.0375 | 0.32478 | 0.28728 |  | 149.99 | 0.118
2 | NO | 12.565 | 5.83 | 0.66468 | 1.362 | 0.35333 | -1.0087 |  | 99.98 | 0.237
3 | YES | 176.375 | 1.502 | 0.0042427 | 0.032 | 0.11836 | 0.08636 |  | 75.03 | 0.328
3 | YES | 149.99 | 1.365 | 0 | 0.118 | 0.15621 | 0.038206 |  | 49.98 | 0.345
3 | YES | 99.98 | 1.492 | 0 | 0.237 | 0.22794 | -0.0090608 |  | 32.48 | 0.297
3 | YES | 75.03 | 1.792 | 0 | 0.328 | 0.26373 | -0.064273 |  | 22.49 | 0.348
3 | YES | 49.98 | 2.329 | 0 | 0.345 | 0.29966 | -0.045342 |  | 5.44 | 0.08
3 | YES | 32.48 | 3.051 | 0.26163 | 0.297 | 0.32476 | 0.02776 |  | 58.76 | 0.066
3 | YES | 22.49 | 3.596 | 0.17112 | 0.348 | 0.33909 | -0.0089107 |  | 49.98 | 0.322
3 | NO | 12.48 | 4.646 | 0.62084 | -1.035 | 0.35345 | 1.3884 |  | 32.475 | 0.328
3 | YES | 5.44 | 7.244 | 0 | 0.08 | 0.36355 | 0.28355 |  | 22.485 | 0.3035
4 | YES | 58.76 | 1.659 | 0 | 0.066 | 0.28706 | 0.22106 |  | 12.49 | 0.506
4 | YES | 49.98 | 1.948 | 0 | 0.322 | 0.29966 | -0.022342 |  | 36.28 | 0.1285
4 | YES | 32.475 | 2.84 | 0.32386 | 0.328 | 0.32477 | -0.003233 |  | 32.49 | 0.1815
4 | YES | 22.485 | 3.6355 | 0.23688 | 0.3035 | 0.3391 | 0.035596 |  | 22.495 | 0.3605
4 | YES | 12.49 | 4.755 | 0.7764 | 0.506 | 0.35343 | -0.15257 |  | 64.75 | 0.094
4 | NO | 4.95 | 6.562 | 0 | -0.733 | 0.36425 | 1.0972 |  | 50 | 0.396
5 | YES | 36.28 | 1.5835 | 0.040305 | 0.1285 | 0.31931 | 0.19081 |  | 32.5 | 0.3455
5 | YES | 32.49 | 1.6745 | 0.088388 | 0.1815 | 0.32475 | 0.14325 |  | 22.495 | 0.514
5 | YES | 22.495 | 1.7665 | 0.047376 | 0.3605 | 0.33908 | -0.021418 |  | 97.56 | 0.358
5 | NO | 12.485 | 2.095 | 0.41436 | 0.924 | 0.35344 | -0.57056 |  | 50 | 0.621
5 | NO | 4.99 | 4.727 | 0 | -0.913 | 0.36419 | 1.2772 |  | 32.5 | 0.4095
6 | YES | 64.75 | 1.555 | 0 | 0.094 | 0.27847 | 0.18447 |  | 22.495 | 0.395
6 | YES | 50 | 1.774 | 0 | 0.396 | 0.29963 | -0.096371 |  | 12.49 | 0.6705
6 | YES | 32.5 | 2.4725 | 0.21425 | 0.3455 | 0.32473 | -0.020769 |  | 50.01 | 0.375
6 | YES | 22.495 | 2.921 | 0.14991 | 0.514 | 0.33908 | -0.17492 |  | 32.495 | 0.4665
6 | NO | 12.505 | 3.404 | 0.44548 | 1.259 | 0.35341 | -0.90559 |  | 22.49 | 0.7295
6 | NO | 5 | 6.662 | 0 | -0.535 | 0.36418 | 0.89918 |  | 50.01 | 0.375
7 | YES | 97.56 | 0.358 | 0 | 0.358 | 0.23141 | -0.12659 |  | 32.505 | 0.4095
7 | YES | 50 | 2.176 | 0 | 0.621 | 0.29963 | -0.32137 |  | 22.495 | 0.611
7 | YES | 32.5 | 2.7875 | 0.15344 | 0.4095 | 0.32473 | -0.084769 |  | 64.88 | 0.399
7 | YES | 22.495 | 3.264 | 0.16405 | 0.395 | 0.33908 | -0.055918 |  | 49.98 | 0.478
7 | YES | 12.49 | 3.7515 | 0.4377 | 0.6705 | 0.35343 | -0.31707 |  | 32.495 | 0.5265
7 | NO | 5 | 6.118 | 0 | -0.644 | 0.36418 | 1.0082 |  | 22.49 | 0.611
8 | YES | 50.01 | 2.283 | 0 | 0.375 | 0.29962 | -0.075385 |  | 174.99 | 0.076
8 | YES | 32.495 | 2.8605 | 0.15061 | 0.4665 | 0.32474 | -0.14176 |  | 150 | 0.274
8 | YES | 22.49 | 3.4495 | 0.21991 | 0.7295 | 0.33909 | -0.39041 |  | 125.01 | 0.471
8 | NO | 12.49 | 4.439 | 0.67175 | 1.094 | 0.35343 | -0.74057 |  | 99.99 | 0.215
8 | NO | 4.98 | 7.285 | 0 | -0.038 | 0.36421 | 0.40221 |  | 75 | 0.258
9 | YES | 50.01 | 2.568 | 0 | 0.375 | 0.29962 | -0.075385 |  | 50 | 0.36
9 | YES | 32.505 | 3.0695 | 0.15203 | 0.4095 | 0.32472 | -0.084776 |  | 22.495 | 0.471
9 | YES | 22.495 | 3.553 | 0.20223 | 0.611 | 0.33908 | -0.27192 |  | 149.99 | 0.25
9 | NO | 12.49 | 4.4255 | 0.59043 | 0.9795 | 0.35343 | -0.62607 |  | 125 | 0.523
9 | NO | 4.98 | 6.71 | 0 | -0.837 | 0.36421 | 1.2012 |  | 100 | 0.199
10 | YES | 64.88 | 2.211 | 0 | 0.399 | 0.27829 | -0.12071 |  | 49.99 | 0.388
10 | YES | 49.98 | 2.435 | 0 | 0.478 | 0.29966 | -0.17834 |  | 32.49 | 0.39
10 | YES | 32.495 | 3.2325 | 0.15627 | 0.5265 | 0.32474 | -0.20176 |  | 22.475 | 0.5135
10 | YES | 22.49 | 3.763 | 0.21779 | 0.611 | 0.33909 | -0.27191 |  | 147.96 | 0.084
10 | NO | 12.49 | 4.725 | 0.59397 | 1.062 | 0.35343 | -0.70857 |  | 125 | 0.119
10 | NO | 4.98 | 7.668 | 0 | -0.174 | 0.36421 | 0.53821 |  | 100 | 0.151
11 | YES | 174.99 | 0.164 | 0 | 0.076 | 0.12035 | 0.044346 |  | 74.98 | 0.226
11 | YES | 150 | 0.336 | 0 | 0.274 | 0.15619 | -0.11781 |  | 50 | 0.409
11 | YES | 125.01 | 0.801 | 0 | 0.471 | 0.19204 | -0.27896 |  | 27.5 | 0.559
11 | YES | 99.99 | 1.777 | 0 | 0.215 | 0.22792 | 0.012925 |  | 112.51 | 0.123
11 | YES | 75 | 1.921 | 0 | 0.258 | 0.26377 | 0.00577 |  | 99.98 | 0.092
11 | YES | 50 | 2.425 | 0 | 0.36 | 0.29963 | -0.060371 |  | 75 | 0.202
11 | YES | 22.495 | 3.712 | 0.28567 | 0.471 | 0.33908 | -0.13192 |  | 50 | 0.487
11 | NO | 4.99 | 7.832 | 0 | -0.285 | 0.36419 | 0.64919 |  | 32.495 | 0.381
12 | NO | 0 | 0 | 0 | -0.049 | 0.37135 | 0.42035 |  | 22.5 | 0.4605
12 | YES | 149.99 | 0.3 | 0 | 0.25 | 0.15621 | -0.093794 |  | 76.38 | 0.0825
12 | YES | 125 | 0.743 | 0 | 0.523 | 0.19205 | -0.33095 |  | 50 | 0.499
12 | YES | 100 | 1.537 | 0 | 0.199 | 0.22791 | 0.028911 |  | 32.51 | 0.4015
12 | YES | 49.99 | 2.471 | 0 | 0.388 | 0.29964 | -0.088356 |  | 22.49 | 0.4825
12 | YES | 32.49 | 3.242 | 0.16405 | 0.39 | 0.32475 | -0.065255 |  | 14.99 | 0.36
12 | YES | 22.475 | 3.7435 | 0.24112 | 0.5135 | 0.33911 | -0.17439 |  | 51.305 | 0.197
12 | NO | 0 | 0 | 0 | -7.767 | 0.37135 | 8.1383 |  | 32.505 | 0.307
13 | YES | 147.96 | 0.143 | 0 | 0.084 | 0.15912 | 0.075118 |  | 22.485 | 0.2455
13 | NO | 0 | 0 | 0 | -0.105 | 0.37135 | 0.47635 |  | 12.49 | 0.142
13 | YES | 125 | 0.943 | 0 | 0.119 | 0.19205 | 0.073051 |  | 92.57 | -0.043
13 | YES | 100 | 2.004 | 0 | 0.151 | 0.22791 | 0.076911 |  | 75 | 0.232
13 | YES | 74.98 | 2.082 | 0 | 0.226 | 0.2638 | 0.037799 |  | 50 | 0.386
13 | YES | 50 | 2.524 | 0 | 0.409 | 0.29963 | -0.10937 |  | 32.495 | 0.298
13 | YES | 27.5 | 3.676 | 0.20082 | 0.559 | 0.3319 | -0.2271 |  | 22.495 | 0.4035
13 | NO | 0 | 0 | 0 | -7.865 | 0.37135 | 8.2363 |  | 75.91 | 0.239
14 | YES | 112.51 | 2.014 | 0 | 0.123 | 0.20997 | 0.086966 |  | 49.99 | 0.426
14 | YES | 99.98 | 2.06 | 0 | 0.092 | 0.22794 | 0.13594 |  | 42.495 | 0.183
14 | YES | 75 | 2.052 | 0 | 0.202 | 0.26377 | 0.06177 |  | 32.49 | 0.279
14 | YES | 50 | 2.533 | 0 | 0.487 | 0.29963 | -0.18737 |  | 22.485 | 0.3395
14 | YES | 32.495 | 3.426 | 0.14566 | 0.381 | 0.32474 | -0.056262 |  | 12.485 | 0.5415
14 | YES | 22.5 | 3.9745 | 0.25809 | 0.4605 | 0.33907 | -0.12143 |  | 86.92 | -0.01
14 | NO | 4.96 | 7.586 | 0 | -0.076 | 0.36423 | 0.44023 |  | 49.99 | 0.366
15 | YES | 76.38 | 1.9875 | 0.00070714 | 0.0825 | 0.26179 | 0.17929 |  | 32.49 | 0.159
15 | YES | 50 | 2.608 | 0 | 0.499 | 0.29963 | -0.19937 |  | 22.495 | 0.292
15 | YES | 32.51 | 3.4195 | 0.11526 | 0.4015 | 0.32472 | -0.076783 |  | 12.57 | 0.288
15 | YES | 22.49 | 3.8985 | 0.2185 | 0.4825 | 0.33909 | -0.14341 |  | 101.255 | 0.0185
15 | YES | 14.99 | 4.462 | 0 | 0.36 | 0.34985 | -0.010153 |  | 75 | 0.075
15 | NO | 4.99 | 7.255 | 0 | 0.771 | 0.36419 | -0.40681 |  | 49.99 | 0.117
16 | YES | 51.305 | 2.569 | 0.041012 | 0.197 | 0.29776 | 0.10076 |  | 32.495 | 0.084
16 | YES | 32.505 | 3.35 | 0.13011 | 0.307 | 0.32472 | 0.017724 |  | 22.495 | 0.2055
16 | YES | 22.485 | 3.8125 | 0.23547 | 0.2455 | 0.3391 | 0.093596 |  | 12.49 | 0.121
16 | YES | 12.49 | 4.748 | 0.55861 | 0.142 | 0.35343 | 0.21143
16 | NO | 4.98 | 6.771 | 0 | -0.326 | 0.36421 | 0.69021
17 | YES | 92.57 | 1.655 | 0 | -0.043 | 0.23857 | 0.28157
17 | YES | 75 | 2.023 | 0 | 0.232 | 0.26377 | 0.03177
17 | YES | 50 | 2.651 | 0 | 0.386 | 0.29963 | -0.086371
17 | YES | 32.495 | 3.429 | 0.10748 | 0.298 | 0.32474 | 0.026738
17 | YES | 22.495 | 3.8325 | 0.18031 | 0.4035 | 0.33908 | -0.064418
17 | NO | 12.49 | 4.9165 | 0.63852 | 0.8645 | 0.35343 | -0.51107
17 | NO | 4.99 | 6.811 | 0 | 0.913 | 0.36419 | -0.54881
18 | YES | 75.91 | 2.169 | 0.019799 | 0.239 | 0.26246 | 0.023465
18 | YES | 49.99 | 2.772 | 0 | 0.426 | 0.29964 | -0.12636
18 | YES | 42.495 | 3.131 | 0.19092 | 0.183 | 0.31039 | 0.12739
18 | YES | 32.49 | 3.528 | 0.057983 | 0.279 | 0.32475 | 0.045745
18 | YES | 22.485 | 3.8375 | 0.17183 | 0.3395 | 0.3391 | -0.00040356
18 | YES | 12.485 | 4.5465 | 0.37689 | 0.5415 | 0.35344 | -0.18806
18 | NO | 4.99 | 6.246 | 0 | -0.144 | 0.36419 | 0.50819
19 | YES | 86.92 | 2.25 | 0 | -0.01 | 0.24667 | 0.25667
19 | NO | 0 | 0 | 0 | -2.132 | 0.37135 | 2.5033
19 | YES | 49.99 | 2.912 | 0 | 0.366 | 0.29964 | -0.066356
19 | YES | 32.49 | 3.585 | 0.041012 | 0.159 | 0.32475 | 0.16575
19 | YES | 22.495 | 3.941 | 0.22062 | 0.292 | 0.33908 | 0.047082
19 | YES | 12.57 | 4.737 | 0.37618 | 0.288 | 0.35332 | 0.065318
19 | NO | 4.98 | 6.354 | 0 | -0.38 | 0.36421 | 0.74421
20 | YES | 101.255 | 2.3035 | 0.016264 | 0.0185 | 0.22611 | 0.20761
20 | YES | 75 | 2.463 | 0 | 0.075 | 0.26377 | 0.18877
20 | YES | 49.99 | 3.102 | 0 | 0.117 | 0.29964 | 0.18264
20 | YES | 32.495 | 3.882 | 0.13576 | 0.084 | 0.32474 | 0.24074
20 | YES | 22.495 | 4.1985 | 0.1096 | 0.2055 | 0.33908 | 0.13358
20 | YES | 12.49 | 4.961 | 0.27294 | 0.121 | 0.35343 | 0.23243
20 | NO | 4.97 | 6.191 | 0 | -0.892 | 0.36422 | 1.2562
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
NO
NO
YES
YES
YES
NO
NO
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
NO
YES
NO
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
NO
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX-CTD - DOX-BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=DOX-CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.10188 |  | X | CALC Y
m= | -0.058094 | 0.015427 |  | Std Dev of Resid |  | 0.31816 |  | 0 | 0.4057
b= | 0.4057 | 0.058381 |  | Resid Sum of Squares |  | 12.6534 |  | 7.832 | -0.049292
# of Values Extracted = |  |  | 135
# of Values Used in Fit = |  |  | 127
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | DOX-CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 302.405 | 1.391 | 0.035355 | -0.003 | 0.32489 | 0.32789 |  | 1.391 | -0.003
1 | YES | 250.01 | 1.721 | 0 | 0.164 | 0.30572 | 0.14172 |  | 1.721 | 0.164
1 | YES | 199.98 | 2.328 | 0 | 0.144 | 0.27046 | 0.12646 |  | 2.328 | 0.144
1 | YES | 150 | 2.993 | 0 | 0.186 | 0.23182 | 0.045823 |  | 2.993 | 0.186
1 | YES | 99.98 | 4.374 | 0 | 0.091 | 0.1516 | 0.060596 |  | 4.374 | 0.091
1 | YES | 74.99 | 4.969 | 0 | 0.053 | 0.11703 | 0.06403 |  | 4.969 | 0.053
1 | YES | 50.01 | 5.614 | 0 | -0.104 | 0.07956 | 0.18356 |  | 5.614 | -0.104
1 | YES | 32.485 | 5.6165 | 0.1096 | 0.1785 | 0.079415 | -0.099085 |  | 5.6165 | 0.1785
1 | YES | 22.505 | 5.654 | 0.35638 | -0.595 | 0.077236 | 0.67224 |  | 5.654 | -0.595
1 | YES | 12.475 | 6.286 | 0.045255 | -0.049 | 0.040521 | 0.089521 |  | 6.286 | -0.049
1 | YES | 4.97 | 6.198 | 0 | -0.129 | 0.045633 | 0.17463 |  | 6.198 | -0.129
2 | YES | 134.8 | 1.62 | 0 | 0.124 | 0.31159 | 0.18759 |  | 1.62 | 0.124
2 | YES | 100 | 2.131 | 0 | 0.31 | 0.2819 | -0.0281 |  | 2.131 | 0.31
2 | YES | 75.01 | 3.306 | 0 | 0.083 | 0.21364 | 0.13064 |  | 3.306 | 0.083
2 | YES | 49.99 | 4.048 | 0 | 0.108 | 0.17053 | 0.062535 |  | 4.048 | 0.108
2 | YES | 32.465 | 4.1255 | 0.030405 | 0.0375 | 0.16603 | 0.12853 |  | 4.1255 | 0.0375
2 | NO | 12.565 | 5.83 | 0.66468 | 1.362 | 0.067012 | -1.295 |  | 1.502 | 0.032
3 | YES | 176.375 | 1.502 | 0.0042427 | 0.032 | 0.31844 | 0.28644 |  | 1.365 | 0.118
3 | YES | 149.99 | 1.365 | 0 | 0.118 | 0.3264 | 0.2084 |  | 1.492 | 0.237
3 | YES | 99.98 | 1.492 | 0 | 0.237 | 0.31902 | 0.082022 |  | 1.792 | 0.328
3 | YES | 75.03 | 1.792 | 0 | 0.328 | 0.30159 | -0.026406 |  | 2.329 | 0.345
3 | YES | 49.98 | 2.329 | 0 | 0.345 | 0.2704 | -0.074602 |  | 3.051 | 0.297
3 | YES | 32.48 | 3.051 | 0.26163 | 0.297 | 0.22845 | -0.068546 |  | 3.596 | 0.348
3 | YES | 22.49 | 3.596 | 0.17112 | 0.348 | 0.19679 | -0.15121 |  | 7.244 | 0.08
3 | NO | 12.48 | 4.646 | 0.62084 | -1.035 | 0.13579 | 1.1708 |  | 1.659 | 0.066
3 | YES | 5.44 | 7.244 | 0 | 0.08 | -0.015133 | -0.095133 |  | 1.948 | 0.322
4 | YES | 58.76 | 1.659 | 0 | 0.066 | 0.30932 | 0.24332 |  | 2.84 | 0.328
4 | YES | 49.98 | 1.948 | 0 | 0.322 | 0.29253 | -0.029469 |  | 3.6355 | 0.3035
4 | YES | 32.475 | 2.84 | 0.32386 | 0.328 | 0.24071 | -0.087288 |  | 4.755 | 0.506
4 | YES | 22.485 | 3.6355 | 0.23688 | 0.3035 | 0.1945 | -0.109 |  | 6.562 | -0.733
4 | YES | 12.49 | 4.755 | 0.7764 | 0.506 | 0.12946 | -0.37654 |  | 1.5835 | 0.1285
4 | YES | 4.95 | 6.562 | 0 | -0.733 | 0.024487 | 0.75749 |  | 1.6745 | 0.1815
5 | YES | 36.28 | 1.5835 | 0.040305 | 0.1285 | 0.31371 | 0.18521 |  | 1.7665 | 0.3605
5 | YES | 32.49 | 1.6745 | 0.088388 | 0.1815 | 0.30842 | 0.12692 |  | 2.095 | 0.924
5 | YES | 22.495 | 1.7665 | 0.047376 | 0.3605 | 0.30308 | -0.057425 |  | 4.727 | -0.913
5 | YES | 12.485 | 2.095 | 0.41436 | 0.924 | 0.28399 | -0.64001 |  | 1.555 | 0.094
5 | YES | 4.99 | 4.727 | 0 | -0.913 | 0.13109 | 1.0441 |  | 1.774 | 0.396
6 | YES | 64.75 | 1.555 | 0 | 0.094 | 0.31536 | 0.22136 |  | 2.4725 | 0.3455
6 | YES | 50 | 1.774 | 0 | 0.396 | 0.30264 | -0.09336 |  | 2.921 | 0.514
6 | YES | 32.5 | 2.4725 | 0.21425 | 0.3455 | 0.26206 | -0.083439 |  | 6.662 | -0.535
6 | YES | 22.495 | 2.921 | 0.14991 | 0.514 | 0.23601 | -0.27799 |  | 0.358 | 0.358
6 | NO | 12.505 | 3.404 | 0.44548 | 1.259 | 0.20795 | -1.0511 |  | 2.176 | 0.621
6 | YES | 5 | 6.662 | 0 | -0.535 | 0.018678 | 0.55368 |  | 2.7875 | 0.4095
7 | YES | 97.56 | 0.358 | 0 | 0.358 | 0.3849 | 0.0269 |  | 3.264 | 0.395
7 | YES | 50 | 2.176 | 0 | 0.621 | 0.27929 | -0.34171 |  | 3.7515 | 0.6705
7 | YES | 32.5 | 2.7875 | 0.15344 | 0.4095 | 0.24376 | -0.16574 |  | 6.118 | -0.644
7 | YES | 22.495 | 3.264 | 0.16405 | 0.395 | 0.21608 | -0.17892 |  | 2.283 | 0.375
7 | YES | 12.49 | 3.7515 | 0.4377 | 0.6705 | 0.18776 | -0.48274 |  | 2.8605 | 0.4665
7 | YES | 5 | 6.118 | 0 | -0.644 | 0.050281 | 0.69428 |  | 3.4495 | 0.7295
8 | YES | 50.01 | 2.283 | 0 | 0.375 | 0.27307 | -0.10193 |  | 7.285 | -0.038
8 | YES | 32.495 | 2.8605 | 0.15061 | 0.4665 | 0.23952 | -0.22698 |  | 2.568 | 0.375
8 | YES | 22.49 | 3.4495 | 0.21991 | 0.7295 | 0.2053 | -0.5242 |  | 3.0695 | 0.4095
8 | NO | 12.49 | 4.439 | 0.67175 | 1.094 | 0.14782 | -0.94618 |  | 3.553 | 0.611
8 | YES | 4.98 | 7.285 | 0 | -0.038 | -0.017515 | 0.020485 |  | 4.4255 | 0.9795
9 | YES | 50.01 | 2.568 | 0 | 0.375 | 0.25651 | -0.11849 |  | 6.71 | -0.837
9 | YES | 32.505 | 3.0695 | 0.15203 | 0.4095 | 0.22738 | -0.18212 |  | 2.211 | 0.399
9 | YES | 22.495 | 3.553 | 0.20223 | 0.611 | 0.19929 | -0.41171 |  | 2.435 | 0.478
9 | YES | 12.49 | 4.4255 | 0.59043 | 0.9795 | 0.1486 | -0.8309 |  | 3.2325 | 0.5265
9 | YES | 4.98 | 6.71 | 0 | -0.837 | 0.015889 | 0.85289 |  | 3.763 | 0.611
10 | YES | 64.88 | 2.211 | 0 | 0.399 | 0.27725 | -0.12175 |  | 7.668 | -0.174
10 | YES | 49.98 | 2.435 | 0 | 0.478 | 0.26424 | -0.21376 |  | 0.164 | 0.076
10 | YES | 32.495 | 3.2325 | 0.15627 | 0.5265 | 0.21791 | -0.30859 |  | 0.336 | 0.274
10 | YES | 22.49 | 3.763 | 0.21779 | 0.611 | 0.18709 | -0.42391 |  | 0.801 | 0.471
10 | NO | 12.49 | 4.725 | 0.59397 | 1.062 | 0.13121 | -0.93079 |  | 1.777 | 0.215
10 | YES | 4.98 | 7.668 | 0 | -0.174 | -0.039765 | 0.13424 |  | 1.921 | 0.258
11 | YES | 174.99 | 0.164 | 0 | 0.076 | 0.39617 | 0.32017 |  | 2.425 | 0.36
11 | YES | 150 | 0.336 | 0 | 0.274 | 0.38618 | 0.11218 |  | 3.712 | 0.471
11 | YES | 125.01 | 0.801 | 0 | 0.471 | 0.35916 | -0.11184 |  | 7.832 | -0.285
11 | YES | 99.99 | 1.777 | 0 | 0.215 | 0.30247 | 0.087465 |  | 0 | -0.049
11 | YES | 75 | 1.921 | 0 | 0.258 | 0.2941 | 0.0361 |  | 0.3 | 0.25
11 | YES | 50 | 2.425 | 0 | 0.36 | 0.26482 | -0.095179 |  | 0.743 | 0.523
11 | YES | 22.495 | 3.712 | 0.28567 | 0.471 | 0.19005 | -0.28095 |  | 1.537 | 0.199
11 | YES | 4.99 | 7.832 | 0 | -0.285 | -0.049292 | 0.23571 |  | 2.471 | 0.388
12 | YES | 0 | 0 | 0 | -0.049 | 0.4057 | 0.4547 |  | 3.242 | 0.39
12 | YES | 149.99 | 0.3 | 0 | 0.25 | 0.38827 | 0.13827 |  | 3.7435 | 0.5135
12 | YES | 125 | 0.743 | 0 | 0.523 | 0.36253 | -0.16047 |  | 0.143 | 0.084
12 | YES | 100 | 1.537 | 0 | 0.199 | 0.31641 | 0.11741 |  | 0 | -0.105
12 | YES | 49.99 | 2.471 | 0 | 0.388 | 0.26215 | -0.12585 |  | 0.943 | 0.119
12 | YES | 32.49 | 3.242 | 0.16405 | 0.39 | 0.21736 | -0.17264 |  | 2.004 | 0.151
12 | YES | 22.475 | 3.7435 | 0.24112 | 0.5135 | 0.18822 | -0.32528 |  | 2.082 | 0.226
12 | NO | 0 | 0 | 0 | -7.767 | 0.4057 | 8.1727 |  | 2.524 | 0.409
13 | YES | 147.96 | 0.143 | 0 | 0.084 | 0.39739 | 0.31339 |  | 3.676 | 0.559
13 | YES | 0 | 0 | 0 | -0.105 | 0.4057 | 0.5107 |  | 2.014 | 0.123
13 | YES | 125 | 0.943 | 0 | 0.119 | 0.35092 | 0.23192 |  | 2.06 | 0.092
13 | YES | 100 | 2.004 | 0 | 0.151 | 0.28928 | 0.13828 |  | 2.052 | 0.202
13 | YES | 74.98 | 2.082 | 0 | 0.226 | 0.28475 | 0.058747 |  | 2.533 | 0.487
13 | YES | 50 | 2.524 | 0 | 0.409 | 0.25907 | -0.14993 |  | 3.426 | 0.381
13 | YES | 27.5 | 3.676 | 0.20082 | 0.559 | 0.19215 | -0.36685 |  | 3.9745 | 0.4605
13 | NO | 0 | 0 | 0 | -7.865 | 0.4057 | 8.2707 |  | 7.586 | -0.076
14 | YES | 112.51 | 2.014 | 0 | 0.123 | 0.2887 | 0.1657 |  | 1.9875 | 0.0825
14 | YES | 99.98 | 2.06 | 0 | 0.092 | 0.28602 | 0.19402 |  | 2.608 | 0.499
14 | YES | 75 | 2.052 | 0 | 0.202 | 0.28649 | 0.08449 |  | 3.4195 | 0.4015
14 | YES | 50 | 2.533 | 0 | 0.487 | 0.25855 | -0.22845 |  | 3.8985 | 0.4825
14 | YES | 32.495 | 3.426 | 0.14566 | 0.381 | 0.20667 | -0.17433 |  | 4.462 | 0.36
14 | YES | 22.5 | 3.9745 | 0.25809 | 0.4605 | 0.1748 | -0.2857 |  | 7.255 | 0.771
14 | YES | 4.96 | 7.586 | 0 | -0.076 | -0.035001 | 0.040999 |  | 2.569 | 0.197
15 | YES | 76.38 | 1.9875 | 0.00070714 | 0.0825 | 0.29024 | 0.20774 |  | 3.35 | 0.307
15 | YES | 50 | 2.608 | 0 | 0.499 | 0.25419 | -0.24481 |  | 3.8125 | 0.2455
15 | YES | 32.51 | 3.4195 | 0.11526 | 0.4015 | 0.20705 | -0.19445 |  | 4.748 | 0.142
15 | YES | 22.49 | 3.8985 | 0.2185 | 0.4825 | 0.17922 | -0.30328 |  | 6.771 | -0.326
15 | YES | 14.99 | 4.462 | 0 | 0.36 | 0.14648 | -0.21352 |  | 1.655 | -0.043
15 | YES | 4.99 | 7.255 | 0 | 0.771 | -0.015772 | -0.78677 |  | 2.023 | 0.232
16 | YES | 51.305 | 2.569 | 0.041012 | 0.197 | 0.25646 | 0.059455 |  | 2.651 | 0.386
16 | YES | 32.505 | 3.35 | 0.13011 | 0.307 | 0.21108 | -0.095916 |  | 3.429 | 0.298
16 | YES | 22.485 | 3.8125 | 0.23547 | 0.2455 | 0.18422 | -0.061284 |  | 3.8325 | 0.4035
16 | YES | 12.49 | 4.748 | 0.55861 | 0.142 | 0.12987 | -0.012131 |  | 4.9165 | 0.8645
16 | YES | 4.98 | 6.771 | 0 | -0.326 | 0.012345 | 0.33835 |  | 6.811 | 0.913
17 | YES | 92.57 | 1.655 | 0 | -0.043 | 0.30955 | 0.35255 |  | 2.169 | 0.239
17 | YES | 75 | 2.023 | 0 | 0.232 | 0.28817 | 0.056174 |  | 2.772 | 0.426
17 | YES | 50 | 2.651 | 0 | 0.386 | 0.25169 | -0.13431 |  | 3.131 | 0.183
17 | YES | 32.495 | 3.429 | 0.10748 | 0.298 | 0.20649 | -0.091505 |  | 3.528 | 0.279
17 | YES | 22.495 | 3.8325 | 0.18031 | 0.4035 | 0.18305 | -0.22045 |  | 3.8375 | 0.3395
17 | YES | 12.49 | 4.9165 | 0.63852 | 0.8645 | 0.12008 | -0.74442 |  | 4.5465 | 0.5415
17 | YES | 4.99 | 6.811 | 0 | 0.913 | 0.010022 | -0.90298 |  | 6.246 | -0.144
18 | YES | 75.91 | 2.169 | 0.019799 | 0.239 | 0.27969 | 0.040693 |  | 2.25 | -0.01
18 | YES | 49.99 | 2.772 | 0 | 0.426 | 0.24466 | -0.18134 |  | 2.912 | 0.366
18 | YES | 42.495 | 3.131 | 0.19092 | 0.183 | 0.22381 | 0.040806 |  | 3.585 | 0.159
18 | YES | 32.49 | 3.528 | 0.057983 | 0.279 | 0.20074 | -0.078257 |  | 3.941 | 0.292
18 | YES | 22.485 | 3.8375 | 0.17183 | 0.3395 | 0.18276 | -0.15674 |  | 4.737 | 0.288
18 | YES | 12.485 | 4.5465 | 0.37689 | 0.5415 | 0.14157 | -0.39993 |  | 6.354 | -0.38
18 | YES | 4.99 | 6.246 | 0 | -0.144 | 0.042845 | 0.18684 |  | 2.3035 | 0.0185
19 | YES | 86.92 | 2.25 | 0 | -0.01 | 0.27499 | 0.28499 |  | 2.463 | 0.075
19 | NO | 0 | 0 | 0 | -2.132 | 0.4057 | 2.5377 |  | 3.102 | 0.117
19 | YES | 49.99 | 2.912 | 0 | 0.366 | 0.23653 | -0.12947 |  | 3.882 | 0.084
19 | YES | 32.49 | 3.585 | 0.041012 | 0.159 | 0.19743 | 0.038432 |  | 4.1985 | 0.2055
19 | YES | 22.495 | 3.941 | 0.22062 | 0.292 | 0.17675 | -0.11525 |  | 4.961 | 0.121
19 | YES | 12.57 | 4.737 | 0.37618 | 0.288 | 0.13051 | -0.15749 |  | 6.191 | -0.892
19 | YES | 4.98 | 6.354 | 0 | -0.38 | 0.03657 | 0.41657
20 | YES | 101.255 | 2.3035 | 0.016264 | 0.0185 | 0.27188 | 0.25338
20 | YES | 75 | 2.463 | 0 | 0.075 | 0.26261 | 0.18761
20 | YES | 49.99 | 3.102 | 0 | 0.117 | 0.22549 | 0.10849
20 | YES | 32.495 | 3.882 | 0.13576 | 0.084 | 0.18018 | 0.096178
20 | YES | 22.495 | 4.1985 | 0.1096 | 0.2055 | 0.16179 | -0.043709
20 | YES | 12.49 | 4.961 | 0.27294 | 0.121 | 0.11749 | -0.003505
20 | YES | 4.97 | 6.191 | 0 | -0.892 | 0.04604 | 0.93804
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX-BOT | DOX-CTD |  | DOX-CTD - DOX-BOT
Average Size: | 4
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX-BOT | DOX-CTD | DOX-CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2003-26-0008 | 2003-26-0008 | 2 | 303.71 | 302.405 | 3.3588 | -1.305 | 1.394 | 1.391 | 0.035355 | -0.003
2003-26-0008 | 2003-26-0008 | 1 | 252.69 | 250.01 | 0 | -2.68 | 1.557 | 1.721 | 0 | 0.164
2003-26-0008 | 2003-26-0008 | 1 | 202.92 | 199.98 | 0 | -2.94 | 2.184 | 2.328 | 0 | 0.144
2003-26-0008 | 2003-26-0008 | 1 | 151.98 | 150 | 0 | -1.98 | 2.807 | 2.993 | 0 | 0.186
2003-26-0008 | 2003-26-0008 | 1 | 102.32 | 99.98 | 0 | -2.34 | 4.283 | 4.374 | 0 | 0.091
2003-26-0008 | 2003-26-0008 | 1 | 76.87 | 74.99 | 0 | -1.88 | 4.916 | 4.969 | 0 | 0.053
2003-26-0008 | 2003-26-0008 | 1 | 52.25 | 50.01 | 0 | -2.24 | 5.718 | 5.614 | 0 | -0.104
2003-26-0008 | 2003-26-0008 | 2 | 32.19 | 32.485 | 3.5285 | 0.295 | 5.438 | 5.6165 | 0.1096 | 0.1785
2003-26-0008 | 2003-26-0008 | 2 | 22.01 | 22.505 | 3.5285 | 0.495 | 6.249 | 5.654 | 0.35638 | -0.595
2003-26-0008 | 2003-26-0008 | 2 | 12.13 | 12.475 | 3.585 | 0.345 | 6.335 | 6.286 | 0.045255 | -0.049
2003-26-0008 | 2003-26-0008 | 1 | 2.15 | 4.97 | 0 | 2.82 | 6.327 | 6.198 | 0 | -0.129
2003-26-0010 | 2003-26-0010 | 1 | 134.76 | 134.8 | 0 | 0.040009 | 1.496 | 1.62 | 0 | 0.124
2003-26-0010 | 2003-26-0010 | 1 | 101.85 | 100 | 0 | -1.85 | 1.821 | 2.131 | 0 | 0.31
2003-26-0010 | 2003-26-0010 | 1 | 76.68 | 75.01 | 0 | -1.67 | 3.223 | 3.306 | 0 | 0.083
2003-26-0010 | 2003-26-0010 | 1 | 50.88 | 49.99 | 0 | -0.89 | 3.94 | 4.048 | 0 | 0.108
2003-26-0010 | 2003-26-0010 | 2 | 31.75 | 32.465 | 3.5709 | 0.715 | 4.088 | 4.1255 | 0.030405 | 0.0375
2003-26-0010 | 2003-26-0010 | 2 | 12.34 | 12.565 | 3.6133 | 0.225 | 4.468 | 5.83 | 0.66468 | 1.362
2003-26-0012 | 2003-26-0012 | 2 | 177.48 | 176.375 | 2.185 | -1.105 | 1.47 | 1.502 | 0.0042427 | 0.032
2003-26-0012 | 2003-26-0012 | 1 | 152.99 | 149.99 | 0 | -3 | 1.247 | 1.365 | 0 | 0.118
2003-26-0012 | 2003-26-0012 | 1 | 102.9 | 99.98 | 0 | -2.92 | 1.255 | 1.492 | 0 | 0.237
2003-26-0012 | 2003-26-0012 | 1 | 77.8 | 75.03 | 0 | -2.77 | 1.464 | 1.792 | 0 | 0.328
2003-26-0012 | 2003-26-0012 | 1 | 53.35 | 49.98 | 0 | -3.37 | 1.984 | 2.329 | 0 | 0.345
2003-26-0012 | 2003-26-0012 | 2 | 32.38 | 32.48 | 3.5214 | 0.099999 | 2.754 | 3.051 | 0.26163 | 0.297
2003-26-0012 | 2003-26-0012 | 2 | 22.89 | 22.49 | 3.5355 | -0.4 | 3.248 | 3.596 | 0.17112 | 0.348
2003-26-0012 | 2003-26-0012 | 2 | 12.23 | 12.48 | 3.5214 | 0.25 | 5.681 | 4.646 | 0.62084 | -1.035
2003-26-0012 | 2003-26-0012 | 1 | 2.68 | 5.44 | 0 | 2.76 | 7.164 | 7.244 | 0 | 0.08
2003-26-0014 | 2003-26-0014 | 1 | 58.89 | 58.76 | 0 | -0.13 | 1.593 | 1.659 | 0 | 0.066
2003-26-0014 | 2003-26-0014 | 1 | 53.48 | 49.98 | 0 | -3.5 | 1.626 | 1.948 | 0 | 0.322
2003-26-0014 | 2003-26-0014 | 2 | 32.18 | 32.475 | 3.5285 | 0.295 | 2.512 | 2.84 | 0.32386 | 0.328
2003-26-0014 | 2003-26-0014 | 2 | 22.27 | 22.485 | 3.5002 | 0.215 | 3.332 | 3.6355 | 0.23688 | 0.3035
2003-26-0014 | 2003-26-0014 | 2 | 12.22 | 12.49 | 3.4931 | 0.27 | 4.249 | 4.755 | 0.7764 | 0.506
2003-26-0014 | 2003-26-0014 | 1 | 3.5 | 4.95 | 0 | 1.45 | 7.295 | 6.562 | 0 | -0.733
2003-26-0018 | 2003-26-0018 | 2 | 37.52 | 36.28 | 1.8243 | -1.24 | 1.455 | 1.5835 | 0.040305 | 0.1285
2003-26-0018 | 2003-26-0018 | 2 | 32.23 | 32.49 | 3.5355 | 0.26 | 1.493 | 1.6745 | 0.088388 | 0.1815
2003-26-0018 | 2003-26-0018 | 2 | 22.48 | 22.495 | 3.5285 | 0.014999 | 1.406 | 1.7665 | 0.047376 | 0.3605
2003-26-0018 | 2003-26-0018 | 2 | 12.56 | 12.485 | 3.5426 | -0.075001 | 1.171 | 2.095 | 0.41436 | 0.924
2003-26-0018 | 2003-26-0018 | 1 | 2.99 | 4.99 | 0 | 2 | 5.64 | 4.727 | 0 | -0.913
2003-26-0019 | 2003-26-0019 | 1 | 64.89 | 64.75 | 0 | -0.14 | 1.461 | 1.555 | 0 | 0.094
2003-26-0019 | 2003-26-0019 | 1 | 52.28 | 50 | 0 | -2.28 | 1.378 | 1.774 | 0 | 0.396
2003-26-0019 | 2003-26-0019 | 2 | 32.68 | 32.5 | 3.5355 | -0.18 | 2.127 | 2.4725 | 0.21425 | 0.3455
2003-26-0019 | 2003-26-0019 | 2 | 22.22 | 22.495 | 3.5426 | 0.275 | 2.407 | 2.921 | 0.14991 | 0.514
2003-26-0019 | 2003-26-0019 | 2 | 12.77 | 12.505 | 3.5426 | -0.265 | 2.145 | 3.404 | 0.44548 | 1.259
2003-26-0019 | 2003-26-0019 | 1 | 3.14 | 5 | 0 | 1.86 | 7.197 | 6.662 | 0 | -0.535
2003-26-0020 | 2003-26-0020 | 1 | 97.77 | 97.56 | 0 | -0.21 | 0 | 0.358 | 0 | 0.358
2003-26-0020 | 2003-26-0020 | 1 | 52.52 | 50 | 0 | -2.52 | 1.555 | 2.176 | 0 | 0.621
2003-26-0020 | 2003-26-0020 | 2 | 32.9 | 32.5 | 3.5214 | -0.4 | 2.378 | 2.7875 | 0.15344 | 0.4095
2003-26-0020 | 2003-26-0020 | 2 | 22.95 | 22.495 | 3.5426 | -0.455 | 2.869 | 3.264 | 0.16405 | 0.395
2003-26-0020 | 2003-26-0020 | 2 | 12.46 | 12.49 | 3.5355 | 0.03 | 3.081 | 3.7515 | 0.4377 | 0.6705
2003-26-0020 | 2003-26-0020 | 1 | 3.09 | 5 | 0 | 1.91 | 6.762 | 6.118 | 0 | -0.644
2003-26-0021 | 2003-26-0021 | 1 | 52.31 | 50.01 | 0 | -2.3 | 1.908 | 2.283 | 0 | 0.375
2003-26-0021 | 2003-26-0021 | 2 | 32.49 | 32.495 | 3.5426 | 0.0049972 | 2.394 | 2.8605 | 0.15061 | 0.4665
2003-26-0021 | 2003-26-0021 | 2 | 22.84 | 22.49 | 3.5214 | -0.35 | 2.72 | 3.4495 | 0.21991 | 0.7295
2003-26-0021 | 2003-26-0021 | 2 | 12.51 | 12.49 | 3.5355 | -0.020001 | 3.345 | 4.439 | 0.67175 | 1.094
2003-26-0021 | 2003-26-0021 | 1 | 2.64 | 4.98 | 0 | 2.34 | 7.323 | 7.285 | 0 | -0.038
2003-26-0022 | 2003-26-0022 | 1 | 52.8 | 50.01 | 0 | -2.79 | 2.193 | 2.568 | 0 | 0.375
2003-26-0022 | 2003-26-0022 | 2 | 32.86 | 32.505 | 3.5285 | -0.355 | 2.66 | 3.0695 | 0.15203 | 0.4095
2003-26-0022 | 2003-26-0022 | 2 | 22.84 | 22.495 | 3.5285 | -0.345 | 2.942 | 3.553 | 0.20223 | 0.611
2003-26-0022 | 2003-26-0022 | 2 | 12.86 | 12.49 | 3.5355 | -0.37 | 3.446 | 4.4255 | 0.59043 | 0.9795
2003-26-0022 | 2003-26-0022 | 1 | 3.12 | 4.98 | 0 | 1.86 | 7.547 | 6.71 | 0 | -0.837
2003-26-0023 | 2003-26-0023 | 1 | 63.62 | 64.88 | 0 | 1.26 | 1.812 | 2.211 | 0 | 0.399
2003-26-0023 | 2003-26-0023 | 1 | 52.56 | 49.98 | 0 | -2.58 | 1.957 | 2.435 | 0 | 0.478
2003-26-0023 | 2003-26-0023 | 2 | 32.61 | 32.495 | 3.5568 | -0.115 | 2.706 | 3.2325 | 0.15627 | 0.5265
2003-26-0023 | 2003-26-0023 | 2 | 22.86 | 22.49 | 3.5497 | -0.37 | 3.152 | 3.763 | 0.21779 | 0.611
2003-26-0023 | 2003-26-0023 | 2 | 12.52 | 12.49 | 3.5355 | -0.030001 | 3.663 | 4.725 | 0.59397 | 1.062
2003-26-0023 | 2003-26-0023 | 1 | 3.26 | 4.98 | 0 | 1.72 | 7.842 | 7.668 | 0 | -0.174
2003-26-0024 | 2003-26-0024 | 1 | 177.27 | 174.99 | 0 | -2.28 | 0.088 | 0.164 | 0 | 0.076
2003-26-0024 | 2003-26-0024 | 1 | 152.71 | 150 | 0 | -2.71 | 0.062 | 0.336 | 0 | 0.274
2003-26-0024 | 2003-26-0024 | 1 | 127.06 | 125.01 | 0 | -2.05 | 0.33 | 0.801 | 0 | 0.471
2003-26-0024 | 2003-26-0024 | 1 | 102.43 | 99.99 | 0 | -2.44 | 1.562 | 1.777 | 0 | 0.215
2003-26-0024 | 2003-26-0024 | 1 | 76.9 | 75 | 0 | -1.9 | 1.663 | 1.921 | 0 | 0.258
2003-26-0024 | 2003-26-0024 | 1 | 52.62 | 50 | 0 | -2.62 | 2.065 | 2.425 | 0 | 0.36
2003-26-0024 | 2003-26-0024 | 2 | 22.2 | 22.495 | 3.5426 | 0.295 | 3.241 | 3.712 | 0.28567 | 0.471
2003-26-0024 | 2003-26-0024 | 1 | 2.57 | 4.99 | 0 | 2.42 | 8.117 | 7.832 | 0 | -0.285
2003-26-0025 | 2003-26-0025 | 0 | 182.69 | 0 | 0 | -182.69 | 0.049 | 0 | 0 | -0.049
2003-26-0025 | 2003-26-0025 | 1 | 151.43 | 149.99 | 0 | -1.44 | 0.05 | 0.3 | 0 | 0.25
2003-26-0025 | 2003-26-0025 | 1 | 127.07 | 125 | 0 | -2.07 | 0.22 | 0.743 | 0 | 0.523
2003-26-0025 | 2003-26-0025 | 1 | 102.32 | 100 | 0 | -2.32 | 1.338 | 1.537 | 0 | 0.199
2003-26-0025 | 2003-26-0025 | 1 | 52.42 | 49.99 | 0 | -2.43 | 2.083 | 2.471 | 0 | 0.388
2003-26-0025 | 2003-26-0025 | 2 | 32.58 | 32.49 | 3.5497 | -0.090004 | 2.852 | 3.242 | 0.16405 | 0.39
2003-26-0025 | 2003-26-0025 | 2 | 22.45 | 22.475 | 3.5285 | 0.024998 | 3.23 | 3.7435 | 0.24112 | 0.5135
2003-26-0025 | 2003-26-0025 | 0 | 3.12 | 0 | 0 | -3.12 | 7.767 | 0 | 0 | -7.767
2003-26-0026 | 2003-26-0026 | 1 | 148.13 | 147.96 | 0 | -0.17 | 0.059 | 0.143 | 0 | 0.084
2003-26-0026 | 2003-26-0026 | 0 | 137.5 | 0 | 0 | -137.5 | 0.105 | 0 | 0 | -0.105
2003-26-0026 | 2003-26-0026 | 1 | 122.35 | 125 | 0 | 2.65 | 0.824 | 0.943 | 0 | 0.119
2003-26-0026 | 2003-26-0026 | 1 | 102.47 | 100 | 0 | -2.47 | 1.853 | 2.004 | 0 | 0.151
2003-26-0026 | 2003-26-0026 | 1 | 77.4 | 74.98 | 0 | -2.42 | 1.856 | 2.082 | 0 | 0.226
2003-26-0026 | 2003-26-0026 | 1 | 52.7 | 50 | 0 | -2.7 | 2.115 | 2.524 | 0 | 0.409
2003-26-0026 | 2003-26-0026 | 2 | 27.74 | 27.5 | 3.5355 | -0.24 | 3.117 | 3.676 | 0.20082 | 0.559
2003-26-0026 | 2003-26-0026 | 0 | 2.5 | 0 | 0 | -2.5 | 7.865 | 0 | 0 | -7.865
2003-26-0027 | 2003-26-0027 | 1 | 112.58 | 112.51 | 0 | -0.07 | 1.891 | 2.014 | 0 | 0.123
2003-26-0027 | 2003-26-0027 | 1 | 102.44 | 99.98 | 0 | -2.46 | 1.968 | 2.06 | 0 | 0.092
2003-26-0027 | 2003-26-0027 | 1 | 77.54 | 75 | 0 | -2.54 | 1.85 | 2.052 | 0 | 0.202
2003-26-0027 | 2003-26-0027 | 1 | 52.44 | 50 | 0 | -2.44 | 2.046 | 2.533 | 0 | 0.487
2003-26-0027 | 2003-26-0027 | 2 | 32.42 | 32.495 | 3.5285 | 0.075005 | 3.045 | 3.426 | 0.14566 | 0.381
2003-26-0027 | 2003-26-0027 | 2 | 22.49 | 22.5 | 3.5355 | 0.01 | 3.514 | 3.9745 | 0.25809 | 0.4605
2003-26-0027 | 2003-26-0027 | 1 | 2.51 | 4.96 | 0 | 2.45 | 7.662 | 7.586 | 0 | -0.076
2003-26-0028 | 2003-26-0028 | 2 | 77.88 | 76.38 | 1.9516 | -1.5 | 1.905 | 1.9875 | 0.00070714 | 0.0825
2003-26-0028 | 2003-26-0028 | 1 | 52.41 | 50 | 0 | -2.41 | 2.109 | 2.608 | 0 | 0.499
2003-26-0028 | 2003-26-0028 | 2 | 32.18 | 32.51 | 3.5355 | 0.33 | 3.018 | 3.4195 | 0.11526 | 0.4015
2003-26-0028 | 2003-26-0028 | 2 | 22.26 | 22.49 | 3.5214 | 0.23 | 3.416 | 3.8985 | 0.2185 | 0.4825
2003-26-0028 | 2003-26-0028 | 1 | 12.14 | 14.99 | 0 | 2.85 | 4.102 | 4.462 | 0 | 0.36
2003-26-0028 | 2003-26-0028 | 1 | 2.62 | 4.99 | 0 | 2.37 | 6.484 | 7.255 | 0 | 0.771
2003-26-0029 | 2003-26-0029 | 2 | 52.66 | 51.305 | 1.8456 | -1.355 | 2.372 | 2.569 | 0.041012 | 0.197
2003-26-0029 | 2003-26-0029 | 2 | 32.3 | 32.505 | 3.5426 | 0.205 | 3.043 | 3.35 | 0.13011 | 0.307
2003-26-0029 | 2003-26-0029 | 2 | 22.22 | 22.485 | 3.5285 | 0.265 | 3.567 | 3.8125 | 0.23547 | 0.2455
2003-26-0029 | 2003-26-0029 | 2 | 12.42 | 12.49 | 3.5497 | 0.07 | 4.606 | 4.748 | 0.55861 | 0.142
2003-26-0029 | 2003-26-0029 | 1 | 2.8 | 4.98 | 0 | 2.18 | 7.097 | 6.771 | 0 | -0.326
2003-26-0030 | 2003-26-0030 | 1 | 92.62 | 92.57 | 0 | -0.050003 | 1.698 | 1.655 | 0 | -0.043
2003-26-0030 | 2003-26-0030 | 1 | 76.92 | 75 | 0 | -1.92 | 1.791 | 2.023 | 0 | 0.232
2003-26-0030 | 2003-26-0030 | 1 | 52.06 | 50 | 0 | -2.06 | 2.265 | 2.651 | 0 | 0.386
2003-26-0030 | 2003-26-0030 | 2 | 32.06 | 32.495 | 3.5426 | 0.435 | 3.131 | 3.429 | 0.10748 | 0.298
2003-26-0030 | 2003-26-0030 | 2 | 22.18 | 22.495 | 3.5285 | 0.315 | 3.429 | 3.8325 | 0.18031 | 0.4035
2003-26-0030 | 2003-26-0030 | 2 | 12.29 | 12.49 | 3.5214 | 0.2 | 4.052 | 4.9165 | 0.63852 | 0.8645
2003-26-0030 | 2003-26-0030 | 1 | 2.96 | 4.99 | 0 | 2.03 | 5.898 | 6.811 | 0 | 0.913
2003-26-0031 | 2003-26-0031 | 2 | 77.12 | 75.91 | 1.2728 | -1.21 | 1.93 | 2.169 | 0.019799 | 0.239
2003-26-0031 | 2003-26-0031 | 1 | 52.7 | 49.99 | 0 | -2.71 | 2.346 | 2.772 | 0 | 0.426
2003-26-0031 | 2003-26-0031 | 2 | 42.44 | 42.495 | 3.5285 | 0.055004 | 2.948 | 3.131 | 0.19092 | 0.183
2003-26-0031 | 2003-26-0031 | 2 | 32.84 | 32.49 | 3.5214 | -0.35 | 3.249 | 3.528 | 0.057983 | 0.279
2003-26-0031 | 2003-26-0031 | 2 | 22.56 | 22.485 | 3.5285 | -0.074999 | 3.498 | 3.8375 | 0.17183 | 0.3395
2003-26-0031 | 2003-26-0031 | 2 | 12.93 | 12.485 | 3.5426 | -0.445 | 4.005 | 4.5465 | 0.37689 | 0.5415
2003-26-0031 | 2003-26-0031 | 1 | 2.58 | 4.99 | 0 | 2.41 | 6.39 | 6.246 | 0 | -0.144
2003-26-0032 | 2003-26-0032 | 1 | 87.26 | 86.92 | 0 | -0.34 | 2.26 | 2.25 | 0 | -0.01
2003-26-0032 | 2003-26-0032 | 0 | 67.51 | 0 | 0 | -67.51 | 2.132 | 0 | 0 | -2.132
2003-26-0032 | 2003-26-0032 | 1 | 52.42 | 49.99 | 0 | -2.43 | 2.546 | 2.912 | 0 | 0.366
2003-26-0032 | 2003-26-0032 | 2 | 32.64 | 32.49 | 3.5355 | -0.15 | 3.426 | 3.585 | 0.041012 | 0.159
2003-26-0032 | 2003-26-0032 | 2 | 22.72 | 22.495 | 3.5285 | -0.225 | 3.649 | 3.941 | 0.22062 | 0.292
2003-26-0032 | 2003-26-0032 | 2 | 12.76 | 12.57 | 3.4224 | -0.19 | 4.449 | 4.737 | 0.37618 | 0.288
2003-26-0032 | 2003-26-0032 | 1 | 3.2 | 4.98 | 0 | 1.78 | 6.734 | 6.354 | 0 | -0.38
2003-26-0033 | 2003-26-0033 | 2 | 102.49 | 101.255 | 1.7607 | -1.235 | 2.285 | 2.3035 | 0.016264 | 0.0185
2003-26-0033 | 2003-26-0033 | 1 | 77.46 | 75 | 0 | -2.46 | 2.388 | 2.463 | 0 | 0.075
2003-26-0033 | 2003-26-0033 | 1 | 51.85 | 49.99 | 0 | -1.86 | 2.985 | 3.102 | 0 | 0.117
2003-26-0033 | 2003-26-0033 | 2 | 32.16 | 32.495 | 3.5285 | 0.335 | 3.798 | 3.882 | 0.13576 | 0.084
2003-26-0033 | 2003-26-0033 | 2 | 22.58 | 22.495 | 3.5285 | -0.085001 | 3.993 | 4.1985 | 0.1096 | 0.2055
2003-26-0033 | 2003-26-0033 | 2 | 12.49 | 12.49 | 3.5355 | 0 | 4.84 | 4.961 | 0.27294 | 0.121
2003-26-0033 | 2003-26-0033 | 1 | 3.1 | 4.97 | 0 | 1.87 | 7.083 | 6.191 | 0 | -0.892
Create Fit Sheet