## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2003-31

Agency: OSAP

Location: West Coast Vancouver Island / Queen Charlotte Sound

Project: LaPerouse/ Nestucca /Covariability

Party Chief: Mackas D.

Platform: John P. Tully

Date: September 22, 2003 – September 30, 2003

Processed by: Germaine Gatien

Date of Processing: 17 January 2004 – 29 January 2004

Number of original CTD casts: 55 (including 2 surface only files and 1 upcast)

Number of casts processed: 53 (52 downcasts plus one upcast)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with Transmissometer 498DR, Seapoint Fluorometer (#2229 on primary pump) and SeaBird Model SBE43 Dissolved Oxygen sensor (#47 on secondary pump). Two SeaBird model 11 deck units were used (#0619 and #0425). The data logging computer was FS01 until 26 September when it was changed to FS02. Salinity samples were analyzed by Doug Anderson using salinometer model 8410, s/n 58879.

# SUMMARY OF QUALITY AND CONCERNS

The data were generally in good order although problems were experienced at sea with erratic bottle firing and spiking in the data due to power shorts. Misfiring during cast #66 led to uncertainties in matching samples with rosette data. There were a few errors in the assignment of sample numbers for chlorophyll and nutrients. The dissolved oxygen files had a few errors. The notes in the logs were extensive and helpful in tracking down errors.

When power shorts occur there are small gaps in the data (1-5db.) But the salinity data sometimes spike to very large values and only slowly come back to correct values; this can mean the loss of 50 to 100db of data. When such problems are occurring it would be helpful is someone could watch the screen; then the CTD descent could be halted if the salinity looks bad, and restarted when the system has recovered.

Because of problems with the salinity due to power shorts, the selection of primary or secondary sensors varied from cast to cast, and in the case of cast #66 an amalgam of the two was created to minimize the data loss. The upcast of #66 will also be archived and the primary sensors were selected for that.

The dissolved oxygen sensor continues to show poor time response. Attempts were made to correct for this but errors are considered to be on the order of:

- ±0.5ml/l in the top 200

- ±0.2 ml/l below 200m – 1500db

- the instrument is considered unreliable below 1500db.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Many problems with data spikes and bottle firing were mentioned in the log and changes in the logging computer and deck unit were made late in the cruise. Changing the logging computer did not solve the problem, but after the deck unit change there is no note of further trouble.

Salinity and oxygen calibration data were obtained. Titrated chlorophyll data were obtained.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The con file was changed after the first cast to remove a -1db offset to pressure. File 2003-31CTD.con was created with the pressure offset set to 0.

The sensor history was found.

##### Conversion of Raw Data

The raw data were converted using configuration file 2003-31ctd.con. Rosette files were converted. A few casts were checked to ensure that all expected channels contained reasonable data.

The last two casts had errors in the naming format; those were corrected in the converted files.

##### WILDEDIT

There appear to be quite a few spikes in conductivity but most are in the upcasts.

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

This routine was used to advance the secondary conductivity by 0.073s for casts 60 to the end. The deck unit used for the earlier casts was one of the newer ones that do the advancement automatically. Later in the processing SHIFT will be used to fine-tune the alignment of both sensors.

##### CELLTM

For 2003-27, which preceded this cruise, (0.03,9) was used for (alpha, 1/beta) for the primary conductivity and (0.02,9) for the secondary.

Tests were run on casts 26, 68 and 79 using (0.02,9), (0.03,9), (0.02,7) and (0.03,7). The overall best choice was found to be the same as for 2003-27 although for the secondary the results varied somewhat with depth and cast; CELLTM was run on all casts using (0.03,9) for the primary and (0.02,9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

The differences between sensors were generally noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 22 | 1800 | +0.001 | +0.00006 | -0.0001 | Noisy, high |
| 56 | 1800 | +0.0006 | +0.00004 | -0.0003 | Noisy, very high |
| 82 | 1800 | +0.0008 | +0.00005 | -0.0001 | Extremely noisy, high |
| 86 | 1800 | +0.0005 | +0.00003 | -0.0001 | Fairly noisy, High |

The differences in temperature are similar to those during 2003-27, but for conductivity and salinity the sensors are much closer than during the earlier cruise. The differences are generally similar to those of 2003-26.

The transmissivity looks good with maxima around 88.8% and up and downcast similar for the deepest cast. The fluorescence looks reasonable. The dissolved oxygen looks too smooth, as has been observed in all cruises using that sensor.

As was found during 2003-27 there are frequent spikes in the data.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

The rosette files were converted to IOS files, put through CLEAN to add event numbers and named BOT.

##### Checking Headers

A header summary and a header check were produced.

- Cast #66 was found to have errors in the lat and long; these were corrected in the CNV file and Conversion and CLEAN were rerun.

- According to the CTD log event #30 was a net cast, but the file contains a full upcast. There was a computer crash at the bottom of cast #29 and a restart as cast #30. The rosette log also identifies the bottle cast as #29. The upcast CTD data will not be archived. The bottle files will be renamed as event #29, file 2003-31-0029.BOT.

- The station name was wrong for cast #104. This was corrected in the headers.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 2.2db. This is reasonable but there is one cast with pressure -0.04db and a few casts with surface values <1db. These were investigated. The values at these low pressures look like they are from very close to the surface, so the pressure are probably about right.

T0, T1, S0 and S1 were plotted for all casts. As noted in the log there are a lot of spikes. Severe problems occurred with salinity in casts 19, 21, 53, 56 and 66. For #19 the worst problems are in the upcast. For the others the downcasts are corrupted. Upcasts should be considered for the archive, but given they were all rosette casts with many bottle stops, the data may not be good enough.

Casts # 89 and 102 contain only surface data; these were Net casts and should not be processed further.

Cast #30 contains only upcast data and only the bottle file will be processed further.

##### BOTTLE FILE PREPARATION

The BOT files were examined for bad values in temperature and salinity and a few outliers were found in casts #57, 67, 90, 104, 110.  CTDEDIT was used to clean outliers in the secondary salinity in casts #57 and 90. The output was copied to the BOT files.

The problems in cast #67 turned out to be due to the fact that the CTD was moving when the bottles fired; the firing was unintentional. For casts #104 and #110 the noise appears to be due to real conditions; the CTD was relatively still and had equilibrated properly, but large changes occurred around the firing time.

The BOT files were averaged to enable an addsamp file to be created. A few problems were encountered:

- Cast #29 was identified as a rosette cast in the logs, but this cast was aborted during the downcast and restarted as cast #30. The BOT files are from cast #30 only.

- Cast #60 includes many bottles fired as a test. Only bottle #24 was fired for sampling.

- Cast #66 – some bottles did not fire (125,150,175,200,250,300,400db). The bottle numbers in the BOT files do not correspond to the bottle numbers in the rosette frame. A cursory look at the CHL and SAL data suggests that the samples did come from the bottle positions given in the original entries in the rosette log, so the bottle numbers in the SAMAVG file were changed to correspond to the rosette log. There is some confusion as extra notes in the rosette log are contradictory, so extra care will be exercised in running COMPARE to ensure the match was done correctly. NOTE: the bottle numbers were later changed for the shallow bottles.

- Cast #67 – 23 bottles were fired, but only one sample taken from Niskin #23. The addsamp file was edited to reflect this.

The addsamp file was edited to add sample numbers taken from the rosette sheets. The ADDSAMP file was input to the ADD SAMPLE NUMBER routine to create SAM files from the BOT files. Those files were averaged on bottle number. (Output: SAMAVG).

(a) The DOXX files contain dissolved oxygen data with flags and comments.

There were errors in the DOXX files for the following casts:

- Cast #25 was a concatenation of 25 and 35. They were separated with a text editor and run through CLEAN to fix the headers.

- Cast #30 – one line was repeated and was out-of order. This was fixed using a text editor.

(b) The file 2003-31chlarc.xls contains the titrated chlorophyll data. The data lacked a cast number, so that was added to the spreadsheet. In doing that job it became apparent that there were errors in some of the data. An edited version of the original spreadsheet 2003-31chlarc-edited.xls was prepared with the following changes:

- Station LD02, sample #38, chl=2.89 – According to the log sample #38 is from cast 21, station LC11 and is too deep for a CHL value this high. This is presumed to be from cast #38, station LD02 from a Niskin bottle at 5m which had no sample number assigned. The sample number was changed to #938 on the edited spreadsheet.

- Station LBP1, sample #238, chl=5.19 – According to the log sample #238 was from cast 68, station LCP3, 150m. For station LBP1, station 73, there was a sample that was originally called #238, but that was renamed #249. That sample is from the surface so the CHL value is reasonable. The sample number was changed to 249 in the edited spreadsheet.

- Sample 341s from station CPE1. There is no note about this sample in either the CTD log or the Rosette log. It may be a duplicate. The value is close to that of sample #341. This value was dropped from the edited spreadsheet since it will not have a significant effect on the comparison with the fluorometer.

The edited file was used to prepare 2003-31chlarc.csv which was converted to CHL1 files. HEADER EDIT was used to add the following header comment to all CHL1 files with output named CHL:

Average of two samples is reported. Variability is assessed as the % (std dev/mean*100).			Note that average variability is usually around 4.5% for total chlorophyll measurements, and   		some of the chlorophyll duplicates in this cruise vary by more than this.

Variabilities greater than 10% have been flagged “c”.

Specific comments were added to those files with flags where the comment is anything other than “variability>10%. (casts 10 and 90)

(c) The file 2003-31sal.csv was converted to SAL files; the only comments were from cast #66 and these were copied into the headers of the relevant SAL file. Other problems found in the data were:

- Cast #10 has errors in the sample numbers. A sample number 0 with Salinity:Bottle ~16 was entered for the sample from Niskin bottle #3. It should be sample #15. All the sample numbers and salinity bottle values thereafter are shifted one position from those given in the rosette sheet. The SAL file was edited to match the information on the analysis sheets and rosette sheet, but no sample #22 was found in the salinity analysis sheet, so Niskin #10 was removed from the sal file.

- Cast #21 sample #38 – This should be sample #938, cast #38. See CHL section above.

(d) The nutrient data in file QFNuts2003-31.xls have errors in sample numbers:

- Station LD02, sample #38 – As noted above for CHL this is believed to come from cast #38, station LD02 from a Niskin bottle at 5m which had no sample number assigned. The sample number was changed to #938 in the csv spreadsheet.

- Station LBP1, sample #238– As noted above for CHL the sample that was originally called #238 was renamed #249 in the log. The sample number was changed to 249 in the csv spreadsheet.

- Station LG03, sample #47 – According to the rosette log and CTD log, sample #47 is from cast #21, station LC11 and there was no nutrient sampling noted for that sample on the rosette log. The cast at station LG03 is cast #47 and the only sample was #138 from a Niskin bottle fired at 4m. The sample number was changed to 138 in the csv spreadsheet.

The csv spreadsheet was converted to NUT1 files. HEADEDIT was used to add the following comment to all headers (output NUT):

Nutrients analyzed fresh onboard JP Tully by Melissa Hennekes Sept 24 - 29, 2003

Any Phosphate samples flagged “d” had a noisy peak.

The SAL, DOXX, CHL and NUT files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3,MRG)

All bottle files for cast #30 were renamed as cast #29 to match the records in the logs. Those files named #30 were from the upcast portion of event #29 and all the bottles were fired during that portion.

##### 11. COMPARE

DISSOLVED OXYGEN

The analyst noted concern about the low dissolved oxygen values. It was later discovered that the concentration of the thorium was different from usual. The files were corrected to reflect this.

COMPARE was run. The best fit was against SBE dissolved oxygen; when outliers were ignored the following trendline was found:

- Titrated oxygen = 1.1553*(SBE_DOX) – 0.0269. (See 2003-31-dox-comp1.xls)

which lies between the results of 2003-27 (Aug. 03) and 2003-41 (Dec 03) when the fits were:

Titrated oxygen = 1.1362*(SBE_DOX) - 0.0143. (See 2003-27-oxy-comp.xls)

- Titrated oxygen = 1.1731*(SBE_DOX) - 0.0289. (See 2003-41-oxy-comp.xls)

- In late December the sensor was recalibrated before being repaired. At that time the relationship was roughly (estimated from a graph, no data available):

Titrated oxygen = 1.16*(SBE_DOX) - 0.07. (See 2003-27-oxy-comp.xls)

- From March to August the sensor did not appear to change calibration much, but from August to December it appears to have been changing more rapidly.

- The only outliers sufficiently far off to need flagging were for cast #1 (0.3db and 10db). Flags will be entered for the bottles from 0 to 100db for cast #66 since there remains some doubt about the depth of firing.

- After running SHIFT and DELETE a set of recalibrated, averaged and edited downcast files will be compared with the titrated values. At that point a fit of differences vs pressure may be chosen to correct the data further to minimize time-response errors not corrected by SHIFT.

Salinity comparison

The first run of COMPARE turned up 2 severe outliers and 4 minor ones after data from above 200db were excluded. Investigation of these outliers led to the discovery that one sample was mislabelled as sample #38 from cast #21. There was no such salinity sample according to the rosette log. The same problem occurred with CHL and NUT data and was relabelled #938. It is a surface sample. This was fixed in the SAL file and the merging process was rerun. Nothing could be found to explain the problem in cast #1, bottle #1. The other differences were not severe and from examining the bottle files it appears that for two cases the wait time before firing was a little short, and for two others there was a lot of motion during the wait time with variations in salinity that could easily explain the differences. The salinity for cast #1 Bottle #1 was flagged as “c” and a comment added to the header.

When the flagged samples, data from above 200dbars and 4 other outliers are excluded there is a reasonable fit of differences versus pressure. There is a slight hint of time dependence, but excluding one cast removes that. Given most of the deep samples were in the middle of the cruise the time-dependence is considered insignificant. Both sensors gave salinity that was high, the primary by 0.0035 units and the secondary by 0.0032 units. (See 2003-31-sal-comp1.xls for details.)

For cast #66 the analyst flagged all samples because of problems identifying the depth at which the bottles fired. In COMPARE the samples from 500db to the bottom look ok. From the surface to 100db the differences are very large and confirm the concerns of the analyst. Realigning the values improves the picture. If we set the bottle numbers so that samples #228 to 234 are from Niskin #7- 13 the comparison is reasonable. The improved differences do not affect the fit since they are from the top 100m only. All salinity values will be left flagged “d” because there is still some doubt about them. (See 2003-31-sal-comp1-cast66.xls for details.)

Fluorescence-chlorophyll comparison

The changes made to bottle numbers for cast #66 had a good effect on the chlorophyll comparison.

For the whole cruise the differences between chlorophyll-a and fluorescence showed a lot of scatter but, on average, there is a reasonable correspondence between CHL and FL for FL<5 with a fit:

CHL = 0.93 * FL + 0.40

For 2003-12 the result was CHL = 0.9 * FL - 0.03.

For FL>5 the fluorescence is << than the chlorophyll.

The outliers for FL<6 were associated with high values of CHL so were not flagged as it is likely the problem is with the CTD values. (See 2003-31_chl_comp.xls for details.)

- The chl values for cast #66 described in the section above were flagged “c” because of the general concern about identifying the depths from which the samples came. The results do not look bad for CHL, but this may be just a reflection that the gradient was low at the depths sampled.

##### SHIFT

Conductivity

Based on the results of 2003-26 (-0.2,-2 for primary/secondary) and 2003-27 (-0.2,0 records), tests were run using advancements of -0.2, -0.5, -0.8, -1 and -2 records. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the primary conductivity of -0.2 records for a net advancement of about 0.065s (since the deck unit had advanced it by +0.073s). For the secondary conductivity the results are variable with values from 0 to -1 looking best for different features. A choice of -0.5 records seems best overall for a net advancement of 0.052s.

All data were put through SHIFT using -0.2 and -0.5 records for the primary and secondary conductivity.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for 2 casts were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The range of values was from 0.8s to 1.6s. Since there are likely to be more flow-rate problems in the upcast than the downcast the conservative choice of 1s was made; a shift of +24 records was applied. This is the same figure that has been applied to all recent data sets.

Oxygen

- The dissolved oxygen sensor continues to show problems that lead to overly smoothed data and a lag. Both flow-rate to the sensor and time-response of the sensor may be sources of error. In January 2004 SeaBird found the sensor to be faulty and replaced the membrane.

- It was noted during 2003-41 that the problem was less severe than in the past, and this data set also seems a little better than we are used to seeing.

- To determine the best shift value to apply a study was made of the offset between downcast and upcast DOX; comparing it to the offset between downcast and upcast temperature leads to estimates of a delay of from +160 to +240 records. Tests using settings of +180, +200 and +220 records suggest that +180 is most suitable. While +220 has been the choice for most cruises in 2003, for 2003-41 a choice of +200 was made. All casts were shifted by +180 records.

##### REVERSE

Casts #21 and 66 were put through REVERSE. The output files were renamed 2003-31-9021.shfox and 2003-31-9066shfox to indicate that they are upcasts. After DELETE the results will be compared with the downcast data to see which are better choices for archiving.

##### DELETE

CLEAN was run to replace pad values in pressure with interpolated values and to fix the event numbers.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Salt: 5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: There were warnings for casts 29, 50, 52, 60 and 66. Some of the warnings are due to pressure jumps of from 2 to 5db between 2 records; there is also a step in temperature and salinity so this seems to be missing data, not bad data. For a few of the warnings there is no jump, just a single bad record with spikes in one or more channels

Similar observations were made during 2003-27 when the same CTD and deck unit were used. During 2003-31 the problems apparently cleared up after the deck unit was changed.

##### DETAILED EDITING

There were spiking problems in both pairs of sensors. The primary sensors are slightly less spiky than the secondary, so they will be selected for most casts. But there are casts during which the primary salinity responded to a spike by going to very high values and slowly returning to correct values leaving about 50 to 100db worth of bad salinity data. For individual casts if there are serious problems in the primary, the secondary sensors will be examined to see if the data are better. Upcast data will also be considered, but the problem casts were generally rosette casts so the upcast data are not expected to be good.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case. In a few cases the descent rate and pump status were examined to determine if unstable features should be removed or not.

Casts # 21, 53, 56, 66 were corrupted probably due to power surges. These were marked by spikes in temperature and salinity. These occurred for many casts, but for casts #21, 53 and 56 the primary salinity went off-scale when the spikes occurred and only slowly returned to normal values so that up to 75db of data were bad. The secondary salinity is not as good in general, but it was not as badly affected by the spiking, so it was selected for these casts. For cast #66 the primary salinity was bad from about 610 to 670db and 790 to 840db but the secondary was just as bad with corruption of the salinity from 560 to 660db and from 1150db to 1175db. The upcast data are poor, with the usual upcast problems plus frequent stops for bottles. Two files will be produced, the first with a patch of primary and secondary downcast salinity to produce as much data as possible and the second with upcast data. A number of steps are required to produce the downcast file. The primary salinity will be used from the surface to the spike around 790db, the secondary from 790 to 900db and the primary from 900db to the bottom. The steps taken were:

- Fracture the file into segments FR1, FR2, FR3 based on scan number

- Copy FR1 and FR3 to FCAL1, FCAL3

- Recalibrate the secondary salinity in FR2 by adding 0.0005 units to match the primary (calibration control file 2003-31-recal-sp.ccf). This salinity will be recalibrated later with all the other files that used primary salinity. Output: FCAL2.

- Remove the primary T and S from FCAL2 and the secondary from FCAL1 and FCAL3.

Note: After the editing was finished it was discovered that in fact only 0.0003 units should have been added to the salinity, but since the error is exceedingly small and affects only a small part of the profile, the data were not corrected.

- Rename the temperature and salinity in the fractured segments

- Join the segments, name output file NEW

- Edit the new file. This is the most difficult step because care is needed in the 650-700db region to decide when the salinity has settled to an acceptable value. Differences between the primary and secondary salinity settle down to expected values at about 680db. It may be that one or the other is acceptable above that, but there is no way of determining which is correct.

- Add a note to the headers to explain how the file was created and why

All casts required editing. The following casts were edited extensively: 21, 53, 57-66, 82, 86.

The descent rate was kept high for the most part, but it was extremely noisy for some casts.

Note that the secondary channels were edited for cast #21, 53 and 56; #66 contains a combination of primary and salinity.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exists with either edited data or data that does not require editing.

##### RECALIBRATION

The EDT, MRG and SAM files were recalibrated using file 2003-31-recal1.ccf to recalibrate salinity and dissolved oxygen using the results of COMPARE.

COMPARE was rerun after this recalibration (2003-31-sal-comp2.xls and 2003-31-dox-comp2.xls) and the results were satisfactory. (Output: COR1, MRGCOR1 and SAMCOR1)

##### FINAL OXYGEN COMPARISON

A set of downcast files were created (from COR1 files after metre-averaging) thinned to pressures close to those used for the bottles during this cruise. COMPARE was run comparing those files to the bottle channel in the MRG files. (See 2003-31-dox-comp3.xls.) A second recalibration was applied to the SBE dissolved oxygen channel in the COR1 files using 2003-31-recal2.ccf based on the results of the fit of differences versus pressure. COMPARE was run again and the results were reasonable. (See 2003-31-dox-comp4.xls.) This final recalibration was applied to the CTD files only, not the bottle files since the lag is not an issue in the bottle files. (Output: COR2)

##### Special Fluorometer Processing

The COR2 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

The full bottle files (SAMCOR1) were also saved after being put through REMOVE and HEADEDIT to produce files named BOF which were also saved to a CD-ROM. Copies of the processing report and file document were put on the CD-ROM as well.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – This equipment has been used 5 times since the last recalibration, but there was generally little or no salinity sampling and a lot of scatter. During 2003-27 the primary sensors were found to be high by 0.006 and the secondary by 0.004 units of salinity. No post-cruise calibration is available.

Local Climatology – The only excursions from the historic ranges were:

- A cast at station ODAS for which the temperature was slightly high from 700 to 800db. The salinity was within but on the low end of the range for the same depths.

- A few shallow casts in Queen Charlotte Sound area where local variations are expected to be very large and such excursions are not unexpected.

The casts with severe power-short problems were examined carefully and all data fell within the historic ranges. There is no evidence of problems with the calibration of the CTD.

Repeat Casts – Cast #55 was a repeat of cast 53 with about 2 hours between the two. The latter cast was only to 500db. A comparison of the two casts shows good correspondence along lines of constant density, with the largest differences < 0.02Cº and < 0.003units of salinity from 400-500db. Most of the differences are much less than that.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles.

Profile plots were produced with Transmissivity, Dissolved Oxygen and Fluorescence versus Pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The following channels were removed from all casts except #21, 53, 56 and 66:

Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag

The following channels were removed from cast #21, 53, 56:

Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag

The following channels were removed from cast #66:

Scan_Number, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag

HEADER EDIT was used to add a comment, fix formats and channel names and correct an error in the headers using file 2003-31-header.hdr. The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Final Bottle Files

The MRGCOR2 files were put through CLEAN to remove the SeaBird headers and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Flag, and Sample_Number.

HEADER EDIT was run to add a header describing analysis methods and quality flags and to fix channel names and formats. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. Cast #60 contains information from many bottle firings, but sampling was done for only one bottle; the others were done to test the rosette. So data from unsampled bottles were removed from the file. It was then run through CLEAN and renamed CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

19. Error light on several times, trouble tripping bottles, error spikes due to power shorts

21. Data spikes due to power shorts

29. Error light, computer crashed at bottom bottle trip. Cast restarted as #30 – upcast only.

29/30. Bottle cast logged as 2003-31-0030.BOT was renamed as 2003-31-0029.BOT to match rosette log

30. Upcast only. Downcast is Event #29.

32. Jellyfish strands noticed on CTD.

43. Winch re-terminated

53. Water sampler failure; only tripped 4 bottles.

55. Repeat of cast 53; full bottle sampling

56. Problems with bottle trips – 2 bottles tripped at surface at first trip

57. Problems with bottle firing – no confirm on fire. Seasave not transmitting to underwater unit. Acquisition computer changed.

60. Many bottles fired while CTD moving either as test or by accident; only one bottle sampled.

64. Data error light @450m; hit bottom @895m.

66. MEAN time local on screen display; error light on during upcast @430m. Deck unit changed from #0619 to  #0425 and converter changed from 4 to 2 pin after this cast. Some bottles did not fire.

67. Bottle fired while CTD moving – these were either tests or by accident, only one bottle was sampled.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-31

Dates: Start: September 22, 2003 End: September 30, 2003

Location: West Coast Vancouver Island / Queen Charlotte Sound

Vessel: Tully

Party Chief: Mackas D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550 	Cruise ID#:   2003-31

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 22/04/03 | Factory |  |  |
| Conductivity | 2173 | 24/04/03 | “ |  |  |
| Secondary Temp. | 2968 | 22/04/03 | “ |  |  |
| Secondary Cond. | 1729 | 24/04/03 | “ |  |  |
| Transmissometer | 498DR | 18/08/03 | IOS |  |  |
| SBE Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 75636 | 06/04/99 | Factory |  |  |