## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 2 January 2004 | 2nd oxygen recalibration applied to CTD files (see end of report) |

## PROCESSING NOTES

Cruise: 2003-32

Agency: OSAP

Location: Gulf of Alaska

Project: Argo

Party Chief: Robert M.

Platform: John P. Tully

Date: August 18, 2003 – August 26, 2003

Processed by: Germaine Gatien

Date of Processing: 11 December 2003 – 12 December 2003

Number of original CTD casts: 3

Number of casts processed: 3

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD ( #0550) was mounted with a SeaTech transmissometer (#498DR), a Seabird Dissolved Oxygen sensor, model 43 (#47) and a Seapoint Fluorometer (#2229) with a 10X cable. The pressure sensor was #75636.

The deck unit was a SeaBird Model 11+ (#0619).

# SUMMARY OF QUALITY AND CONCERNS

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained, read and entries about problems noted.

The cruise summary sheet was completed.

The sensor history was found.

##### Conversion of Raw Data

This was a very short cruise without rosette sampling. The same equipment was used for 2003-27 which immediately followed 2003-32. The same configuration file will be used as was used for many casts of the later cruise, namely, 2003-27-CTDa.con. There was a -1db offset in the pressure calibration for the original configuration files; this was kept in the file above, but it was found during 2003-27 that a +2db offset was needed. This will be applied later.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### CELLTM

Using the results of tests on data from 2003-26 and 2003-27, CELLTM was run on all casts using (0.03,9) for the primary and (0.02,9) for the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A few casts were plotted to check for agreement between the pairs of T and C sensors.

The differences between sensors were generally noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 2000 | +0.001 | +0.00017 | +0.0012 | Very Noisy, high |
| 2 | 2000 | +0.001 | +0.00020 | +0.0015 | Very Noisy, high |
| 3 | 2000 | +0.001 | +0.00019 | +0.0015 | Very Noisy, high |

The differences are similar to those of 2003-27.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

All IOS files were put through CLEAN to replace pad values in the Pressure channel.

##### Checking Headers

A header summary, cruise track and header check were produced and no errors were found.

The average surface pressure is 1.1db which is lower than usual as noted during 2003-27. Looking at the upcasts there is data that has in-water values but pressures of -2 to -2.5db. For cast #1 the surface looks like it is close to P = -2db. An offset of -1.0db was included in the original configuration files and those used for this conversion. Recalibration will be done later to add 2db to these values as was done for 2003-27.

The fluorometer data looks bad in the upper 60db in both downcasts and upcasts, so it will be removed later. Similar problems were noted during 2003-27 until a cable was changed. The oxygen and transmissivity data look ok.

##### 10. SHIFT

Conductivity

Using the results of 2003-27 all files were put through SHIFT using -0.2 records for the primary sensor only. The secondary sensor did not require a SHIFT

Dissolved Oxygen

Based on the results of tests for 2003-27 a shift of +220 records was applied to this data.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 20.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies to the whole range

Sample interval = 0.042 seconds. (taken from header)

There were no warnings.

##### DETAILED EDITING

Page plots were produced using the primary sensors since these were used for most of 2003-27. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. The descent rate was extremely noisy and heavy editing was needed for all casts to remove records corrupted by shed wakes.  Note was made of the editing details in the relevant files.

##### Other Comparisons

Previous experience with these sensors – Both conductivity sensors were recalibrated in April 2003. They were used during 2003-24 and 2003-26 but there was a lot of scatter in the bottle comparisons for both. During 2003-27 the primary salinity was found to be high by about 0.004 and the secondary by about 0.006.

Argo Float profiles – These 3 casts were each associated with the launch of an ARGO float so there is are CTD profiles available for comparison. However, those profiles come 10 days later and after some drift so we cannot expect extremely close correspondence. The profiles were checked and the differences in salinity along lines of constant density were within ±0.002psu and ±0.001around 2000db. This was before the CTD data was recalibrated. If the results of 2003-27 are applied to this data then the ARGO data will be within ±0.006psu of the CTD.

Historic ranges – No historic range data was available for this region.

##### RECALIBRATION

The data was recalibrated using file 2003-27-rcal1.ccf which added 2db to the pressure, subtracted 0.004psu from the primary salinity and applied the following correction to the dissolved oxygen data:

DOX = 1.1362 *DOX measured – 0.0143

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data, and profile plots were prepared of Temperature, Salinity, DOX and Transmissivity.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

For all casts the following channels were removed: Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Pump Status, Oxygen:Voltage, Fluorescence, Descent_Rate and Flag channels.

HEADER EDIT was used to fix formats and channel names.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

REVISION (2 January 2004) A second recalibration should have been applied to the dissolved oxygen channel, based on the analysis for 2003-27. File 2003-27-recal2.ccf was used to apply the following correction to the three CTD files:

DOX (corrected) = DOX (after 1st recal) -0.0988 + 9 E-5 * Pressure

(See 2003-27_proc.doc for details.)

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-32

Dates: Start: August 18, 2003 End: August 26, 2003

Location: Gulf of Alaska

Vessel: John P. Tully

Party Chief: Robert, M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | No | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550

Cruise ID#:		2003-32

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date(dd/mm/yy) | Location | Date | Location |
| Temperature | 2038 | 22/04/03 | Factory |  |  |
| Conductivity | 2173 | 24/04/03 | “ |  |  |
| Secondary Temp. | 2968 | 22/04/03 | “ |  |  |
| Secondary Cond. | 1729 | 24/04/03 | “ |  |  |
| Transmissometer | 498DR | 18/Aug/03 | IOS |  |  |
| Dissolved Oxygen | 47 | 21/Aug/02 |  |  |  |
| Fluorometer | 2229 |  | IOS |  |  |
| Pressure Sensor | 75636 | 11/01/96 | Factory |  |  |