## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2003-33

Agency: OSAP

Location:

Project:

Party Chief: Juhász T.

Platform: John P. Tully

Date: September 30, 2003 – October 12, 2003

Processed by: Germaine Gatien

Date of Processing: January 17, 2003 – February 10, 2003

Number of original CTD casts: 	80

Number of casts processed: 80

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlabs transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter (#1024), a Biospherical QSP200L PAR sensor (#4565) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0508) and the logging computer was #FS02. The salinometer was a Portasal model 8410 (#58879).

The temperature and conductivity sensors were changed after cast #6. There is mention of a reference PAR in the log, but details were not provided.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was generally in good order, but some equipment changes were not noted.

The rosette log sheets had some errors in cast numbers and some sample numbers were used twice.

The first 6 casts were done with different temperature and conductivity sensors from those used for the rest of the cruise. There was no salinity sampling for those 6 casts, but post-cruise calibrations were available.

There appears to have been a significant temporal drift in the secondary conductivity sensor used for casts #7 to 80. It also displayed pressure-dependence.

The dissolved oxygen analysis was done partly at sea (casts #1,3,6,9,12,16,17,19,22,25,28,45 and 48) and partly at IOS (casts #33-44 and 66.) There was an error in the concentration of reagent in the analysis at sea; this was corrected later. The dissolved oxygen sensor was not mounted until cast #29 so the only check on the corrections is from casts #45 and 48 which look fine in COMPARE.

The dissolved oxygen data in the CTD files should be considered:

- ±0.6ml/l from 0db–50db (except for the Effingham Inlet casts for which the error is ±5ml/l in the top 3db)

- ±0.2ml/l from 50 to 100db

- ±0.1ml/l below 100db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Problems with the T and C sensors were reported and those sensors were replaced after cast #6.

Salinity, chlorophyll and dissolved oxygen bottle calibration data were obtained. Nutrient data were not yet available. Flags and comments had been added as needed.

The cruise summary sheet was completed.

There were changes of equipment during the cruise with PAR, Surface PAR, DOX and altimeter not always being mounted. Most of the changes were recorded in the log, but not all, so it was necessary to go through all 80 configuration files to determine what was done.

The history of the conductivity sensors was obtained.

Six configuration files were prepared, all with the pressure offset = 0. The offset was checked mid-cruise so should be right for cast #7 to 80, but the earlier casts will have to be checked later and adjusted if necessary.

2003-33-A.con – For casts 1-6 using the pre-cruise calibrations

2003-33-B.con – For casts 1-6 using the post-cruise calibrations

2003-33-C.con – For casts 7-28 with Fluor, Trans and Alt

2003-33-D.con – For casts 29-65 and 69-80 with Fluor, Trans, Alt and DOX

2003-33-E.con – For casts 66-68 Fluor, Trans, PAR, Reference PAR and DOX

2003-33-F.con – For casts 17 Fluor, Trans, PAR, Reference PAR and Alt

The calibration constants were checked for all instruments except that reference PAR.

##### Conversion of Raw Data

The raw data were converted using the configuration files listed above.

There are no salinity calibration samples available for the sensors used for the first 6 casts. But there are post-cruise calibrations. The data were converted using both 2003-33-A.con and 2003-33-B.con to see if the differences are smaller for the post-cruise calibrations. Using the pre-cruise calibrations the differences between the two conductivity channels is on the order of 0.003 S/m below 200db. This is an order of magnitude higher than usual. The primary conductivity changes a little between the two calibrations, but the secondary changes dramatically. Clearly the secondary conductivity pre-cruise calibration should not be used. Since the sensors were not used between this cruise and their return to the factory for recalibration, it seems likely that the post-cruise calibration is the better choice for both sensors. So 2003-33-B.con will be used for casts #1-6.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. Cast #9 stops part way through the upcast.

Rosette files were converted using a start time of -2s and duration of 5s. There are inconsistencies between the casts noted as ROS in the log book and the files created. Some were created where there is no indication that bottles were fired and some were not created where sample numbers are recorded in the log.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

This step was skipped since the deck unit model is one of the newer ones that advances both conductivity channels by +0.073s. Tests will be done later to fine-tune the alignment.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.02,9), (0.03,9), (0.02,7) and (0.03,7); many casts were checked since there were many changes of equipment. The setting of (0.02,7.0) is clearly best for the primary for all casts except for one that looked about the same for all choices. For the secondary it is not quite so obvious with (0.02,7) looking best for many. One cast looked best with (0.03,9) and another looked the same with all choices; both of these casts were ones with very small range of T and S, so they are probably not good ones to judge by. It is unlikely that the secondary sensors will be selected, so this issue was not looked into any further. The secondary sensors have a lot of fine-scale noise in their signals.

All casts were put through CELLTM using (0.02,9) for both pairs of sensors.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

The differences between sensors were generally very noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 250 | -0.0013 | -0.0005 | -0.0035 | Ok |
| 2 | 290 | -0.0012 | -0.0003 | -0.0025 | Noisy,mostly high |
| 9 | 340 | -0.0008 | -0.0002 | -0.0017 | OK |
| 11 | 1150 | +0.0002 | -0.00015 | -0.002 | Good |
| 18 | 1600 | ~0 | -0.00012 | -0.0018 | Noisy, high |
| 29 | 1200 | -0.0004 | -0.00005 | -0.0005 | Mostly very high |
| 57 | 350 | ~0 | +0.0005 | +0.006 | Ok |
| 58 | 300 | -0.0002 | +0.00055 | +0.006 | Ok |

Cast #1 and 2 are had different sensors than the other casts.

Cast #18 was examined for pressure dependence of differences and there is clearly such dependence in the conductivity and salinity with the differences larger at depth.

Casts #57 and 58 seem notably different from the others. There were few deep casts late in the cruise so it is difficult to judge what is going on here. Looking at cast #30 is interesting. That cast was to about 440db and the descent rate was mostly quite steady and high. Yet the differences are very noisy. It looks like something was causing a great deal of variability. Perhaps the flow rate to one pump was not steady.

The secondary salinity is full of fine-scale noise with excursions of ±0.002psu.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

The rosette files were converted to IOS files and renamed BOT.

One cast was checked from each of the CTD and BOT casts to ensure the altimeter routine worked correctly; it did.

All BOT files were plotted and no significant outliers were found.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The cruise track was plotted and no problems noted.

A header summary and a header check were produced and no errors found.

The average surface pressure is 1.9db with a range of -0.1db to +5.5db which seems reasonable. The mixed layer depth calculation shows no casts having surface water sufficiently well-mixed to be useful for salinity calibration purposes.

All CLN files were examined on-screen and no evidence was found of serious problems in T and S. There was some spiking, mostly in salinity.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet was edited to remove the loop samples and was named 2003-33-nonloop.csv. That file was converted to individual SAL files. There were no flagged values so no comments needed in the headers. A number of errors were found in trying to merge these files with the BOT files. For two casts the data was not lined up with pressures and/or bottle numbers; that was fixed in the spreadsheet and in the SAL files.

The dissolved oxygen files (*.add) have comments in the headers to explain flagged values. The ADD files in the folder OXYGEN are incorrect. There was a mix-up over the concentration of a chemical used in the analysis. The files were corrected by Sheila Toews and those in the HYDRO folder are correct. There were a few problems with the data files:

- Cast #17 had samples out of order and had to be edited to enable MERGE to work

- CLEAN was used to change the pad values to -99. (Output ADD1)

- There are two channels with dissolved oxygen values; the first is incorrect but has the standard channel name. The second has the right values and the wrong names. So REMOVE was used to get rid of the wrong data and HEADEDIT was used to rename the good data channel. Some of the data had wrong channel names so that was corrected in a second run of HEADEDIT. (Output: DOX1)

- Many of the dissolved oxygen samples were brought back to IOS where they were analyzed. Those did not require the correction, but an error in how the comment was written had to be corrected in the data from cast #66. Unfortunately, some of the files created at sea got sent with the ones done at IOS, and those had not been corrected. So there was great confusion. Fortunately, these problems were made obvious by COMPARE.

- Another problem that COMPARE turned up is that the dissolved oxygen file from cast #47 is really from cast #48. This was no doubt caused by the mistake on the rosette sheet.

The chlorophyll analysis spreadsheet 2003-33arcchl.xls was edited to add cast numbers and named 2003-33-chl.xls. At the same time two typos in station names were fixed. Sample #291 was used twice during the cruise, so the second one was renamed #9291. The file was then converted to individual CHL files. The data do have flags for some points – a single comment covers all such flags so EDIT HEADER was used to add it to all the CHL files, which were then named CHL1.

The nutrient data in file 2003-33arcnut.xls was edited to add cast numbers and to change sample #271 of cast #47 to sample #9271 as was done for CHL and saved as 2003-33-nut.csv. There were flags and comments in the file. This file was converted to NUT files. Comments were transferred to the headers of casts #6, 12, 45, 67.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. A warning occurred for casts #9 and 18. A check of the files shows that the pump was off for the last few data points (at the surface), and these points were not in the output SAM file. This is probably something to do with an overflow in the descent rate column, but whatever the cause it is a good thing since we don’t want salinity data collected while the pumps were off. Records were rearranged in the SAMAVG, CHL1 and NUT files for cast #47 so that sample numbers are increasing.

The SAL, CHL1, DOX1 and NUT files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3, MRG)

##### 11. COMPARE

Salinity

There was no salinity sampling for the first 6 casts, so no comparison is possible for conductivity sensors #1763 and #2278.

For casts #7 through 80 COMPARE was run and differences above 400db plus a few outliers were excluded. The primary salinity appears high by 0.003psu with no obvious dependence on pressure or temperature. The secondary salinity is high by about 0.0015psu but there is both pressure-dependence and time-dependence in the differences. There is data from only 8 casts and only 7 with data below 400db. (See 2003-33-sal-comp.xls)

There were many bottles fired at one depth during cast #29. When COMPARE is run with pressure as the reference channel all the CTD data for those bottles are combined, so if changes occurred over the time of the firing we might think that some of the bottles were providing poor data. To investigate this, the MRG file for cast #29 was examined in EXCEL. This showed that the salinity drifted upwards by about 0.0008psu and the differences between the two pairs of sensors was quite steady at about 0.0011psu ±0.0002. Pressure changes account for much of the variation in salinity. There is no evidence of trouble with any of the bottles. (See cast29_salinity_study.xls.)

The initial study (section 8) of differences between the two CTD salinity channels suggests that it changed during the cruise. It appears to have been a steady drift, but this conclusion was based on a cursory study of the data. There are no salinity bottles after cast #29. To confirm whether there really was a drift, a study was made of the differences in the CTD files over the whole cruise. The files were thinned and a spreadsheet was created with cast number, temperature, salinity, salinity differences and pressure and this data was examined for any hints on what was happening. Data was limited to 150 to 400db since there is little deep data available for late in the cruise. When the differences were plotted against cast number there was a clear trend as S1-S0 moved from about -0.003 to +0.008psu. To study whether this is not time-dependence but a measure of how different the later casts were from the early ones the differences were plotted against primary salinity. On this plot most of the late data does indeed group around 31psu. However, the only points with salinity >33 from after cast #30 all fall well above other data with similar salinity values from earlier in the cruise. So this is probably a measure of time-dependence. Given what was learned from COMPARE the problem is likely to be with the secondary sensors, so the primary will be selected if possible. (See 2003-33-salinity_study.xls)

Fluorescence vs Chlorophyll-a

Plots of differences between the titrated bottle samples and the CTD fluorescence show that for FL<6 the fluorescence is approximately 3 times the chlorophyll. This was true both close to shore and offshore. (See 2003-33-chl-comp1.xls.)

Dissolved Oxygen

The dissolved oxygen sensor was mounted for casts 29 through 80.

There were many iterations of running COMPARE because of the problems outlined in the section above. Every outlier was traced to ensure that the data had been handled properly.

The differences between the titrated values and CTD values were displayed versus pressure, cast # and titrated values. The best fit was found versus the titrated values. When 5 severe outliers and 10 moderate outliers were excluded, the following fit was found:

Titrated DOX =1.1895 * DOX-CTD – 0.0896

The extreme outliers all come from near-surface samples in Effingham Inlet where the temperature and salinity gradients were high at the surface. This particular SBE oxygen sensor was not capable of resolving sharp gradients, so these results do not suggest a fault in the titrations. Most of the other outliers were also from the near-surface with large values of DOX. Five outliers were from pressures greater than 3db and those were only slight outliers.

Comparing this with the results of other cruises in the fall of 2003,

Titrated DOX = 1.1362 * DOX-CTD – 0.0143 (2003-27 – Sept. 2003)

Titrated DOX = 1.1553 * DOX-CTD – 0.0269 (2003-31 - Sept. 2003)

Titrated DOX = 1.1731 * DOX-CTD – 0.0289 (2003-41 – December 2003),

suggests that it is slightly out of line. There was no sampling below 250db. When only cast #66 was examined and two outliers excluded the result was

Titrated DOX = 1.1763 * DOX-CTD – 0.0513

which is quite close to the results of 2003-41. There is some slight hint of a drift with time, but given that most of the SBE DOX data was taken as the ship travelled along Effingham Inlet, this is probably more a measure of a drift in gradient along the inlet than a drift in the instrument.

The limited sampling and the large surface gradients lead to the conclusion that the 2003-33 results should not be used for the initial recalibration. The results of cruise 2003-41 will be used instead. (See 2003-33-dox-comp1.xls for details.)

##### 12. SHIFT

Conductivity

Tests were run on the primary conductivity only since it is unlikely that the secondary will be used for the archive. There were many changes of equipment which might entail some changes to the flow rates and hence the alignment, so tests were run on casts with different equipment arrangements.

The best results came from advancements of from -0.4 to -0.8 records with variations among casts with the same configuration and different configurations looking similar. Overall -0.6 looked best, on average. That would mean a net advancement of about 0.048s, since the channel was already advanced by +0.073s by the deck unit.

All casts were put through SHIFT using -0.6 records for the primary conductivity, a net advancement of +0.048s. (Output *.SHFC)

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. ALIGNCTD was used to advance the oxygen by from +180 to +220 records and the upcast and downcast traces compared with temperature traces. The alignment is judged by picking out how features line up and by looking at what happens during a bottle stop. Theoretically if the errors affect up and downcast data in the same way, then during a stop for bottles the dissolved oxygen values should drift to a spot about half-way between the two traces. A choice of +180s looks best overall using both approaches, although a larger shift might suit the deepest data. A choice that would make the upcast overlie the downcast better cannot be achieved because of the response time problems. Alignment will be done using SHIFT in IOS SHELL to advance the conductivity by +180 records.

(Because the oxygen sensor was not always mounted, the SHFC files from the previous step were copied to SHFO files. Then new SHFO files were created for those files with the oxygen sensor data.)

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was from 1 to 1.6s. A shift of 1 s (+24 records) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing since the secondary conductivity appears to have drifted significantly.

Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts except # 71 required at least some light editing. Only casts #13 and #48 were edited heavily.

Some of the casts from Effingham Inlet had unstable features in the top 5m. This may be real or might be a sign of slight misalignment of temperature and conductivity. Such features were left unedited.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –

Casts 1-6: 	Primary – used for 2003-29 (+0.002) and 2003-34 (-0.0004); based on assumption of steady drift using post-cruise calibration expect about -0.001psu.

Secondary – used for many cruises but large change noted between May 03 and Sept. 03.

In two previous Sept. cruises the salinity was low by 0.024psu. Based on assumption of steady drift using post-cruise calibration in Dec. 03 expect about -0.022psu.

Casts 7-80:	Primary – used during 2003-41 in Dec; high by 0.0006psu

Secondary – not used since previous recalibration

Historic ranges – All data fell within the historic ranges, however, there was no useful climatology available for the Effingham Inlet casts.

##### Initial Recalibration

File 2003-33-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of cruise 2003-41 and to correct the primary salinity by subtracting 0.003psu; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1) This step was not applied to casts #1 – 6 since there were different conductivity sensors for which there are no bottle comparisons available and there was no dissolved oxygen sensor mounted.

COMPARE was rerun and the corrections were found to be made correctly. (See 2003-33-dox-comp2.xls.)

### Final Dissolved Oxygen Recalibration

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but there remain significant errors due to the response-time problem. The best we can do to correct for this is try to make the downcast CTD data match the bottles from the upcast. A set of downcast files were prepared by metre-averaging the COR1 files, thinning the data to the depths generally used for bottles and reversing them. COMPARE was run comparing those files with the bottles in the MRG files. A fit of differences versus pressure was found and file 2003-33-recal2.ccf was prepared applying the following correction:

DOX(corrected) = DOX (after 1st recal) – 0.2184 +8*E-4 * Pressure

Comparing this with the results for 2003-41:

DOX(corrected) = DOX (after 1st recal) – 0.0503 +1*E-4 * Pressure

indicates that the correction is bigger for this cruise. However, this is likely just because there were no bottles below 250db so the large near-surface errors carry more weight in the fit. Since most casts were shallow this will give a good correction for most casts. However, there were some deep casts and this fit would cause corrections at depth that are inappropriate. To address this problem Formula 12 will be used in the recalibration program so that no correction will be made for pressures = or > 250db. We would normally expect corrections at 1000db to be on the order of 0.05 ml/l so this is not a serious problem.

File 2003-33-recal2.ccf was used to apply a second recalibration to casts #7-80 to the dissolved oxygen channel in the COR1 files. (Output: COR2) This was not applied to the bottle files since they are not subject to the lag.

It was also applied to the reversed thinned downcast data and COMPARE was rerun to check that the second recalibration was done properly. The results show that the DOX errors mostly fall within ±0.6ml/l in the top 50m and within ±0.2ml/l between 50and 100db and ±0.1ml/l below that. (See 2003-33-dox-comp4.xls.)

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: BOX)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, and Transmissivity profiles. Separate profile plots were prepared with DOX, FL, PAR and Surface PAR versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment, fix formats and channel names. For casts 1-6 the comment includes a warning that there were no salinity calibration bottles for the sensors used for those casts.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Bottle Files

The MRG files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1. The conductivity difference was large. A decision was made to change sensors at earliest convenience.

7. T and C sensors changed. Secondary salinity looks unresponsive. Pumping problem suspected.

22. Pressure check done. Pressure offset in configuration file reset from -1db to 0db.

29. DOX sensor added; configuration file changed.

37. DOX very low at bottom

64. Secondary salinity noisy.

65. Secondary conductivity cable changed. PAR on. Altimeter off. Reference PAR on.

67. Conductivity differences increasing – probably problem with secondary conductivity

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-33

Dates: Start: September 30, 2003 End: October 12, 2003

Location: LaPerouse, ECOHAB, Covariability, BIO-Moorings, GVRD, PATC-SARA, Effingham

Vessel: John P. Tully

Party Chief: Juhász T.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2003-33

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 2095 | 16/07/03 05/09/03 | Factory “ | 20/12/03 | Factory |
| Conductivity | 1763 1764 | 24/07/03 05/09/03 | “ “ | 19/12/03 | Factory |
| Secondary Temp. | 2710 2106 | 23/01/03 05/09/03 | “ “ | 20/12/03 | Factory |
| Secondary Cond. | 2278 2128 | 23/01/03 15/05/03 | “ “ | 19/12/03 | Factory |
| Transmissometer | 333DR | 18/08/03 | IOS |  |  |
| Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/99 | Factory |  |  |