## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 28-Sep-2017 | Corrected MISSION metadata field in header. R.H. |
| 22 Dec. 2004 | Recalibrated salinity based on post-cruise calibration; see note below. |
| 16-Dec-04 | Added loop data to the archive. J.L. |

## PROCESSING NOTES

Cruise: 2003-36

Agency: OSAP

Location: N. W. Pacific

Project: High Seas Salmon

Party Chief: Morris J.

Platform: W. E. RICKER

Date: October 6, 2003 – October 29, 2003

Processed by: Germaine Gatien

Date of Processing: 13 February 2004 – 25 February 2004

Number of original CTD casts: 98

Number of casts processed: 98

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N 0471). The salinity was analyzed using Portasal Model #8410 (S/N 58879).

# SUMMARY OF QUALITY AND CONCERNS

The files had non-standard names.

Due to the limited and scattered calibration information, the salinity should be considered ±0.005psu.

The descent rate of the CTD was extremely noisy for many of the offshore casts, but because it was kept high, on average, the loss of data due to shed wakes was minimized. For some shallow, inshore casts the descent rate was quite low so the quality of the salinity data is a little lower.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Bottle salinity and titrated chlorophyll data were obtained.

The cruise summary sheet was completed.

The CTD # is given as 0550 in the Daily Log book, but the pressure sensor is the one from CTD #0506; the sensors are also the ones normally on CTD #0506 and that is the CTD normally used on the Ricker. So it is presumed that the log entry is wrong and this CTD is #0506.

The configuration files were obtained and the calibration constants were checked. A small error in the calibration constants for the pressure sensor was corrected and the resulting file named 2003-36CTD.con.

The sensor history was found. All temperature and conductivity sensors were last used for cruise 2003-13.

##### Conversion of Raw Data

The raw data were converted using configuration file 2003-36-ctd.con.

A few casts were checked and all channels contained reasonable data.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2	Pass 2    Std Dev = 5	Points per block = 50

##### ALIGNCTD

Since the deck unit was one of the older versions that does not advance the secondary conductivity, all casts were put through ALIGNCTD to advance the secondary conductivity by +0.073 so that it matches the primary. Fine-tuning of the alignment will be done later using SHIFT.

##### CELLTM

Tests were run on two casts running CELLTM with choices of (0.02,9), (0.03,9), (0.02,7) and (0.03,7) and for (alpha, 1/beta). The best choice overall was found to be (0.02,7) for the primary and (0.03,9) for the secondary channel which were also the choices made when these sensors were last used.

CELLTM was run on all casts using those settings.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. The primary sensors are a little noisier than the secondary. The differences between conductivity sensors are reasonably steady, but the temperature and salinity differences are highly variable, both within casts and between casts.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 50 | 215 | -0.001 | -0.00025 | -0.0024 | Steady |
| 161 | 215 | -0.0013 | -0.00025 | -0.0015 | Steady |
| 197 | 215 | -0.0005 | -0.00033 | -0.0034 | Steady |
| 236 | 215 | -0.001 | -0.0003 | -0.0019 | Steady |

The transmissivity and fluorescence look reasonable and there is no serious spiking in any data.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers. Two casts had errors in the headers that had to be corrected before they could be converted.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header summary and a header check were produced. Errors in some cast numbers were fixed.

The cruise track was plotted and the positions look reasonable.

The average surface pressure is 5.5db, with a minimum of 3.9db; this is reasonable for the Ricker.

The mixed-layer depth calculation shows that the 10db salinity may be useful for calibration purposes for casts #142, 215, 278 and possibly #263. (However, there was no salinity sampling for #215.)

T0, T1, S0 and S1 were plotted for all casts. The primary and secondary salinity are reasonably close though the differences between sensors are notable during the upcasts. The secondary upcasts look like the downcasts, but there are significant differences in the primary, especially in the high gradient regions.

##### SHIFT

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles for 2 casts were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. As has been found in the past a choice of 1s looks reasonable although the descent rate was very noisy and the offsets highly variable; a shift of +24 records was applied.

Conductivity

Initial tests were run using advancements of +0.5, 0 and -0.5 records on cast #197. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were with a shift to the secondary conductivity of -0.5 records. Further tests were then tried for settings between -1.5 and -0.5 records, narrowing down the best choice to about -1.1 records, or -0.026s. This implies a net advancement of about +0.027s which is unusually low. SHIFT was run on several other casts using -1.1 records and that setting proved satisfactory. This odd alignment together with the noisy temperature suggests some problem in the plumbing or pump.

Similar tests were run for the secondary conductivity and the best results were with a setting of -0.5 records or -0.021s for a net advancement of +0.052. This value is in the usual range.

All data were put through SHIFT twice using -1.1 for the primary and -0.5 records for the secondary conductivity.

##### BOTTLE FILE PREPARATION

Bottle data were received from analysts in spreadsheet format:  2003-36chlarc.xls, 2003-36sal.csv and Nutrient summary 2003-36.xls.

The information from those spreadsheets was merged into 2003-36-Bottles.xls.

Subsets of this file were used to study comparisons of fluorescence with CHL and bottle salinity with CTD salinity.

Salinity comparison

There were only 4 deep salinity samples. The bottle was about 5db above the CTD. The casts were examined to see the depth at which the CTD stopped and 5 db were then subtracted from that to find the depth of the Niskin bottle. Plots were then prepared of salinity at that depth. The CTD moved up and down in the water during the stop so the estimated depth of the bottle is rough.

| Cast | CTD stop | Niskin level |
| --- | --- | --- |
| 83 | 252 | 247 |
| 194 | 301 | 296 |
| 230 | 401.5 | 396.5 |
| 289 | 489** | 484** |

During cast #289 there were stops at the bottom, 501db, and around 496db and 489db during the upcast. This cast is marked by a drop in salinity in the bottom 10 to 15db, and only the stop at 489db is in water with salinity remotely close to that of the bottle. Choosing the Niskin depth of 484db leads to differences close to that found for a number of other bottles, so will be assumed to be the correct value.

The differences between the deep bottles and the downcast CTD salinity shows a lot of variability, but the two deeper ones suggest that the primary salinity is low by about 0.007 and the primary by 0.009psu. This fits with the conclusion of section 8 that the primary is higher than the secondary by about 0.002psu.

| Deep Bottles |  |  |  |  |
| --- | --- | --- | --- | --- |
| Cast | pressure | S0-Sbot | S1-Sbot | S0-S1 |
| 83 | 250 | -0.0003 | -0.0027 | 0.0024 |
| 194 | 300 | -0.0039 | -0.0053 | 0.0014 |
| 230 | 400 | -0.0070 | -0.0088 | 0.0018 |
| 289 | 484 | -0.0078 | -0.0082 | 0.0004 |

Of the 10db bottles there were 3 in well-mixed surface waters. From those we would think that the primary is low by about 0.01psu and the secondary low by about 0.011psu. There are some questions about surface salinity samples having bubbles affect their values, so we should not put too much weight on this, but it is comforting that the results are reasonably close to those of the deep bottles.

| Shallow Bottles |  |  |  |  |
| --- | --- | --- | --- | --- |
| Cast | pressure | S0-Sbot | S1-Sbot | S0-S1 |
| 142 | ~10 | -0.0090 | -0.0104 | 0.0014 |
| 263 | ~10 | -0.0126 | -0.0145 | 0.0019 |
| 278 | ~10 | -0.0073 | -0.0094 | 0.0021 |
| Average |  | -0.0096 | -0.0114 |  |

When all 10db bottles were included in the comparison using CTD data from 11db there is a lot of variability. When points with differences >0.02 are excluded the median value gives the primary salinity low by 0.009psu and the secondary low by 0.011psu. There is a lot of scatter but no evidence of drift with time.  (See 2003-36sal-comp.xls and 2003-36sal-10db.xls.)

Fluorescence-chlorophyll comparison

For comparison with the extracted chlorophyll the shifted files were put through REVERSE and DELETE. These upcast files were then averaged in 0.5db-bins and thinned to just one point at 11db. This is 5db above the average stopping depth; there is some error associated with this assumption, but the only way to do better is to examine every cast in detail which was not done. The 11db values were exported in spreadsheet form and pasted into the table with CHL. One point with very low CHL was excluded from the comparison. There was a lot of scatter in the results, which is likely to be due to the uncertainty in depth of sampling, especially in deep casts where the ship motion was significant. When compared with the extracted chlorophyll and the results overall show a relationship of:

FL = 0.357 * CHL +0.345 (or FL = 0.7 * CHL when forced to have zero offset)

or CHL = 2.8 * FL - 1

For 2003-12 the result was CHL = 0.9 * FL - 0.03.

For 2003-13 the result was CHL = 0.925 * FL +1.19.

So these data look quite different.

To see if the differences are due to misjudging the depth of the bottle, the process was repeated using 9.5db, as has been used in the past. This led to a result of CHL = 1.76 * FL – 0.3, which is still very different from previous results.

All sampling was between 7am and 7pm. (See 2003-36chl_comp.xls.)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Surface Salinity:  5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maxiumum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: The only warnings referred to bottom or upcast data.

##### DETAILED EDITING

While the primary sensors appear to be slightly closer to the bottles they are noisier and the upcast primary data are notably different from the downcast. So secondary sensors will be selected, as was the case the last time these sensors were used.

Page plots were produced using T1,S1. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was extremely noisy for most of the offshore casts, but it was kept high on average, so the loss of data due to shed wakes was minimized. For shallow casts with low descent rate, there is a lot of noise that is indistinguishable from local mixing features, so the ability to remove bad data by editing is limited.

The following casts required no editing: 16, 59, 65, 95, 98, 107, 164, 170, 218, 221, 224, 245, 248.

The following casts required only light editing: 1-13, 19-28, 34-46, 53, 56, 62, 71, 74, 80-92, 101, 104, 140, 142, 158, 161, 167, 173-215, 227-242, 254, 257, 269, 272.

All other casts required heavier editing.

Note was made of the editing details in the relevant files. The edited files were copied to EDT files so that a complete set of files exist with either edited data or data that do not require editing.

##### Recalibration

The secondary salinity is believed to low by from 0.003 to 0.011psu. It was recalibrated by adding 0.008psu and should be considered ±0.005psu using file 2003-36-recal.ccf.

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. In all cases both sensors gave salinity that was too high. The best comparison was in October 2002 when there were surface bottles from well-mixed water. The primary salinity was found to be high by 0.004psu and the secondary salinity high by 0.002psu; however, both temperature sensors were different for those cruises so this information is of limited use. During 2003-13, when the temperature sensors were the same as used for this cruise, the primary was low by 0.006psu and the secondary high by 0.002psu, but there was only 1 deep bottle and the surface water was not well-mixed, so the results are suspect.

Historic ranges – There were some excursions from the historic ranges. Most were for temperature between 40 and 110db for casts that were very close to shore. There were 2 casts (#263 and 266) further offshore which also had temperature higher than expected between 40 and 110db, so this may imply a real change. The salinity data were all within the historic range with the exception of one near-shore cast for which the salinity was a little low between 70 and 100db. There is no indication of instrumental error.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and two header entries and to add the following comment:

There was a lot of scatter in the salinity calibration which was based on only

4 deep bottles and some shallow bottles.

Salinity should be considered +/- 0.005psu.

Transmissivity - The data are unedited.

Fluorescence - The data are unedited.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

49. Headers scrambled.

206. Log notes problems with NMEA date. Headers scrambled.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-36

Dates: Start: October 6, 2003 End: October 29, 2003

Location: North-West Pacific

Vessel: W.E. Ricker

Party Chief: Morris J.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2003-36

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory |  |  |
| Conductivity | 2424 | 16/04/02 | “ |  |  |
| Secondary Temp. | 2374 | 20/06/02 | “ |  |  |
| Secondary Cond. | 2399 | 16/04/02 | “ |  |  |
| Transmissometer | 197 | 22/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |

Dec. 22, 2004: All CTD files were recalibrated using file 2003-36-recal2.ccf to add 0.0027psu to the secondary salinity based on post-cruise calibration of the conductivity sensor. It was assumed that the drift was linear with time. G. Gatien