## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 22 April 2005 | Cast #42 reprocessed – see end of report for details. |

## PROCESSING NOTES

Cruise: 2003-39

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait

Project: Paleo-oceanography

Party Chief: Spear D.

Platform: Vector

Date: October 23, 2003 – November 1, 2003

Processed by: Germaine Gatien

Date of Processing: 10 February 2004 – 18 February 2004

Number of original CTD casts: 	91

Number of casts processed: 91

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Chelsea transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter (#1024) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0424). The salinometer was a Portasal model 8410 (#58879).

# SUMMARY OF QUALITY AND CONCERNS

There were severe noise problems in the primary sensors during cast #26.

The DOX and altimeter channels were not logged for cast #42 due to a configuration file error.

There were problems in the logging of data during casts #70 and 85 with sudden skips in pressure.

For many casts the rosette sheets were only partly filled in; no indication is given of what sampling was done. This can be a problem when we try to track down errors and there is no way to determine if any bottle data were missing. There were a few errors in sample numbers.

The CTD hit the bottom hard several times; there may have been damage to the secondary conductivity sensor and/or dirt got into the pumps. This was not noted in the log and considerable effort was put into understanding some odd things in the data. It would have saved time had the incidents been noted in the log book.

The transmissivity, dissolved oxygen and fluorescence are unedited, except where records were removed in the editing of temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

- ±0.6ml/l in the top 50m

- ±0.4ml/l from 50m-100m

- ±0.2ml/l below 100m

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity, dissolved oxygen and nutrient bottle data were obtained. Flags and comments had been added to the oxygen files as needed. The spreadsheet for nutrients contains a comment that should go in all headers and there are specific flag comments that will have to be transferred to some cast headers. There were no flagged values of salinity, so no need to transfer comments.

There was no chlorophyll sampling.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. There is a -1db pressure offset. The pressure offset will be checked after conversion to IOS SHELL.

The history of the conductivity sensors and dissolved oxygen sensor were found.

##### Conversion of Raw Data

The raw data were converted using the configuration file 2003-39-0001.con except for cast #42. For that cast an error was made in the acquisition and DOX and altimeter were not recorded. So 2003-30-0042.con was used for the conversion of that cast only.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close.

Rosette files were converted using a start time of -4s and duration of 5s. This should capture the records from 4s before firing to 1s after. In the past -2/5 was used but on a recent VECTOR cruise there was a problem with the CTD starting up so quickly that upcast data were captured in the rosette file.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit model is one of the older ones that does not advance the secondary conductivity channel. So alignctd was used to advance it by +0.073s to match the primary conductivity. Tests will be done later to fine-tune the alignments.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.1,7), (0.02,9), (0.03,9), (0.02,7), (0.03,7) and (0.0245, 9.5). Leaving the data alone seems to provide the best results for both sensors.

This step will be skipped.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. Because there was a hint of drift the number of casts examined was increased.

Most of the casts were relatively shallow.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 4 | 440 | -0.0004 | -0.0005 | -0.005 | Steady |
| 6 | 610 | -0.0004 | -0.00043 | -0.0042 | Steady |
| 8 | 272 | -0.0004 | -0.0004 | -0.004 | Steady |
| 13 | 235 | -0.0004 | -0.00017 | -0.0015 | Steady |
| 15 | 290 | -0.0004 | ~0 | ~0 | Steady |
| 25 | 275 | -0.0004 | +0.0006 | +0.0068 | Steady |
| 27 | 260 | -0.0004 | +0.00045 | +0.0052 | Steady |
| 28 | 254 | -0.0005 | +0.0005 | +0.0055 | Steady |
| 43 | 221 | -0.0005 | +0.00045 | +0.005 | Steady |
| 45 | 225 | -0.0005 | +0.0005 | +0.0055 | Steady |
| 55 | 218 | -0.001 | +0.00047 | +0.0055 | Steady |

Casts #4 and #6 were examined for depth dependence and none is obvious.

The temperature looks quite steady; the increase at the end is in very shallow water, so not significant.

The conductivity and salinity differences seem reasonably steady from cast #27 through #55, but between cast #8 and #27 there is a lot of change. It is hoped that the bottle comparisons will explain what happened but there were not many salinity bottles taken.

##### Conversion to IOS Headers

The ROS and CNV files for cast #34 had errors in the SeaBird headers that needed to be corrected before they could be converted.

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files.

CLEAN was run to add event numbers and the output file from this was renamed BOT.

All BOT files were plotted and no significant outliers were found.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary and a header check were produced and one error in station name was found and corrected in the IOS and CLN files.

The average surface pressure is 0.3db with many values less than zero. Checks were made within the files to see if the CTD really recorded out-of-water. It appears that the offset applied in the con file was inappropriate, and that even without that setting the pressures look a little low. File 2003-33-recal-press.ccf will be applied to add 1.2db to the pressure in all files later. This step will be applied to the full CTD files before running DELETE to ensure that DELETE does not remove good data.

All CLN files were examined on-screen for problems in T and S. There are some differences that may indicate problems. These will be checked in the editing process. Casts 28, 43, 70 and 85 need examining.

However, the region is one of active mixing so we do not expect close agreement between up and downcasts.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheets were converted to individual SAL files. There were no flagged values. A few problems arose in merging files and a first run of COMPARE turned up some severe outliers. Investigation of these problems led to the following discoveries and actions:

- The oxygen file named 2003-39-0001.ADD should be 2003-39-0002.ADD.

- Bottle 12, cast #12 had been fired but not sampled. The sample number 999 was assigned to this bottle since putting a pad value there caused trouble in MERGE.

- For cast #16 there is no record of sampling on the rosette sheet. One sample is labelled #118, but the salinity looks like it is from a bottle much deeper. The bottom Niskin seems the most likely one to have been sampled, with sample #111, although #112 is also possible. I changed the sample number in the SAL file to 111, flagged it “”d” and put a comment in the header.

- For cast #18 there is no record of what sampling was done. The analyst notes that there was no sample number on one bottle; she decided that it should be #128, and that looks reasonable. The sample was flagged “d” and an explanation entered into the header as a comment.

The dissolved oxygen files (*.add) have comments in the headers to explain flagged values. One cast needed editing because samples were out of order.

There was no chlorophyll sampling.

The spreadsheet QF2003-39-nut.xls was edited to produce QF2003-33-nut.csv; channel names were changed to standard format. This file was then converted to NUT files. Those files were put through HEADEDIT to add a general comment on the analysis (2003-39-nut-header.txt) (Output NUT1) and specific comments explaining flags were copied from the spreadsheet and written into the individual files. SORT was used to arrange the data with increasing sample numbers (Output NUT2.)

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files.

The SAL, ADD and NUT2 files were merged with SAMAVG in three steps. (Output: MRG1, MRG2, MRG)

##### 11. COMPARE

Salinity

There were only 12 samples, of which 4 were from above 100db and the deepest was only 440db.

After a few initial errors were fixed as described above, COMPARE had only minor outliers in shallow water. The primary salinity showed no significant pressure-dependence but does show some time-dependence. The primary salinity appears to be high by ~0.0008psu. The secondary salinity shows significant pressure-dependence and time-dependence (-0.003psu to +0.003psu) and is, on average, high by ~0.001psu. Examining the plot of differences versus pressure and writing the cast numbers beside the points makes it quite clear that the significant drift is with time, not pressure. (See 2003-39-sal-comp.xls)

No salinity recalibration will be necessary for the primary sensors and while there was drift in the secondary sensors, it is not so large that the data cannot be used.

Dissolved Oxygen

The differences between the titrated values and CTD values were displayed versus pressure, cast # and titrated values. The best fit was found versus titrated values, but there is a lot of scatter. The differences versus file pair number show that most of the outliers are from casts #8 through #14, especially from #10, 11, 12 and 13. When casts #10 – 13 were excluded plus a few outliers the following fit was found:

Titrated DOX = 1.1756 * DOX-CTD – 0.0162 (2003-39)

This is in reasonable agreement with the results of 2003-41. It is a little out of line with cruise 2003-33 but that was considered an unreliable result. It is also quite different from 2003-27 and combined with the 2003-41 results there is evidence of greater change in calibration through the autumn of 2003 than has been noted previously:

Titrated DOX = 1.1362 * DOX-CTD – 0.0143 (2003-27)

Titrated DOX = 1.1895 * DOX-CTD – 0.0896 (2003-33)

Titrated DOX = 1.1731 * DOX-CTD – 0.0289 (2003-41)

(Late in the processing errors were found in the oxygen data in the merged files. After fixing these errors COMPARE was rerun and produced much better results for casts #10 to 13. However, the fit was still very close with

Titrated DOX = 1.178 * DOX-CTD – 0.0221

so the recalibration was not redone. The errors did not affect the salinity comparison)

In the course of investigating the problems in the oxygen data, it was discovered that the CTD hit the bottom hard during cast #9. The altimetry was very noisy near the bottom of the cast, but read about 1.2m a short time before stopping and the descent speed was about 0.6m/s at that point. There is no note of this in the log, but the transmissivity went to zero and the salinity is bad at the bottom. The timing of the first hit fits the time when the salinity differences started to change. Whether the secondary sensor was damaged or the pump had dirt in it is unknown at this point. This event was later confirmed by Steve Mihaly who says that the CTD hit bottom on several occasions. The CTD record suggests hard hits for casts #9, 50 and 52.

##### 12. SHIFT

Conductivity

During 2003-33 the primary conductivity was shifted by +0.8 records and during 2003-41 it was shifted by -0.4 records. Tests were run on several casts, and -0.2 records looked best.

All casts were put through SHIFT using -0.2 records for the primary sensor.

Tests were done for the secondary channel only for cast #26 which had poor primary data. For that cast the secondary conductivity was shifted by -1.2 records.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the test was hard to apply because there was so little variation for many casts. So a shift of 1 s (+24 records) was applied since this is the shift that has been used in most other cruises.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. Using SHIFT the dissolved oxygen channel was advanced by +160, +180 and +200 records and the results examined. For 2003-33 an advancement by +180 records produced the best results for shallow data and it looks like a good choice for this data too. SHIFT was applied to advance the dissolved oxygen channel +180 records.

##### DELETE

File 2003-33-recal-press.ccf was applied to add 1.2db to the pressure in all SHFFL files. (Output: COR1)

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were warnings for casts #70 and #85. Investigation shows that there were equipment problems that affected all channels, probably just not logging anything for a few seconds. For cast #85 the problem occurs only in the top 12db of the downcast and extensively during the upcast, so the downcast data will probably be used. However, for cast #70 the problems are all in the downcast and the temperature and salinity data look very odd during the downcast. The upcast data look better, so will be used.

File 2003-39-0070.shffl was put through reverse, recalibrated and rerun through DELETE. A quick check  was made of the alignment of the conductivity to ensure that the setting chosen works well on upcast data and it does.

##### DETAILED EDITING

The primary sensors were chosen for further processing. They are closer to the bottles and show little drift with time or pressure.

Page plots were produced using (T0,S0). These were used to guide the editing. With a few exceptions the descent rate was generally very steady and high enough.

The following casts required no editing: #17, 38, 66, 81, 84, 89.

The following cast required extensive editing: #2.

All other casts required only light editing. Note was made of the editing details in the headers of the relevant files. And for casts #70 and note was made of the fact that the data came from the upcast.

Because this is an area of active mixing no attempt was made to remove small instabilities that were not clearly due to instrument problems.

##### Other Comparisons

Previous experience with these sensors – The primary conductivity sensor was used during 2003-33 in early October 2003 when the salinity was found to be high by about 0.003psu and during 2003-41 in December when it was high by 0.0006psu. The secondary conductivity was used during 2003-41 in December 2003 when it was found to be low by 0.003 with slight pressure dependence. There was no deep sampling during the December cruise.

Historic ranges – There was little local climatology available for the region; all data fell within or very close to the historic ranges.

Repeat Casts – There were many repeat casts at stations BE1, BE2, SE1 and NS0 though they were in areas of great variability. Plots were prepared with such casts overlaid in the following groups:

BE1:34, 36, 38, 40, 43, 45, 47, 49, 51, 53, 55.

BE2: 44, 46, 48, 50, 52, 54, 56.

SE1: 58, 60, 62, 64, 66, 69, 71, 73, 75, 77, 79, 81, 83, 85, 87.

NS0: 59, 61, 63, 65, 67, 68, 70, 72, 74, 76, 78, 80, 82, 84, 86.

These sites are in areas of expected tidal variation and show considerable variability, as expected, although the total range of T and S is quite small. There is nothing to suggest problems with the data.

##### Initial Recalibration

File 2003-39-press.ccf had already been applied to the downcast files, but was now applied to the MRG and SAM files (Output: MRGCOR1, SAMCOR1).

No recalibration is needed for the primary salinity. The secondary salinity for cast #26 needs to be reduced by 0.005psu to match the primary, based on the differences noted in section 8.

File 2003-39-recal2.ccf was prepared to recalibrate the secondary salinity for cast #26 and to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRGCOR1 and SAMCOR1 files. (Output: COR2, MRGCOR2, SAMCOR2)

COMPARE was rerun for DOX and the corrections were found to be made correctly. (See 2003-39-dox-comp2.xls.) Cast #26 was checked to ensure that the secondary salinity was recalibrated properly; there was no bottle file for that cast, and the CTD file looks right.

### Final Dissolved Oxygen Recalibration

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but there remain significant errors due to the response-time problem. The best we can do to correct for this is try to make the downcast CTD data match the bottles from the upcast. A set of downcast files were prepared by metre-averaging the COR2 files, thinning the data to the depths generally used for bottles and reversing them. COMPARE was run comparing those files with the bottles in the MRGCOR1 files. The fit of differences versus pressure was found using all the data, but there are few bottles below 200db and the fit does not look very good there. Leaving the deep bottles out of the fit produced poor results around 200 to 250db, so they were left in. The fit to be applied from 0 to 210db is:

DOX(corrected) = DOX (after 1st recal) – 0.148 +7*E-4 * Pressure

There will be no corrections applied below 210db. File 2003-39-recal3.ccf was used to apply this second recalibration to the dissolved oxygen channel in the COR2 files. (Output: COR3) The bottle files do not need this recalibration since they are not subject to the lag.

This 2nd recalibration of DOX was also applied to the reversed thinned downcast data and COMPARE was rerun to check that the second recalibration was done properly. The results show that the DOX errors mostly fall within ±0.6ml/l in the top 50m and within ±0.4ml/l from 50 to 100 and ±0.2ml/l below that. (See 2003-39-dox-comp4.xls.)

##### Special Fluorometer Processing

The COR3 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T & S profiles. Separate profile plots were prepared with Temperature, Transmissivity, DOX and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts except #26.

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from l cast #26 only.

HEADER EDIT was used to add a comment, fix formats and channel names using file 2003-39-header.hdr. A specific comment was added to cast #26 to warn that the quality of data below 200m is in doubt due to equipment problems.

The Standards Check routine was run and no problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRGCOR2 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats and to add a standard comment about quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

REVISION: April 22, 2005. Cast #42 was found to contain bad data. When the data was acquired the wrong configuration file was used. The same file was used to convert the data, but in fact the sensors were the wrong ones. The data could not be converted using the proper configuration file because the headers contain the wrong sensor numbers. A new configuration file was created 2003-39-test2.con which contains the proper calibration co-efficients, but the wrong sensor numbers, to match those used in acquisition. It was also necessary to suppress two voltage channels, so DO and Altimetry are not available.

The steps as described in the report were repeated for this cast. There was no rosette for this cast.

Particulars

9. CTD hit bottom hard

26. Primary temperature and salinity bad in places – also DOX. Presume something interfered with pump. The next cast was ok.

33. Beginning of time series

40-42. There was significant prop wash over the CTD at the surface as the ship tried to keep station in windy conditions.

41. According to log, data lost in crash of SeaSave but file looks ok

42. According to log, rerun of cast 42, but aborted due to con file error, but file is complete, though lacking two channels.

57. Beginning of time series

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-39

Dates: Start: October 23, 2003 End: November 1, 2003

Location: B.C.Inlets

Vessel: Vector

Party Chief: Johannessen S.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2003-39

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 05/09/03 | Factory |  |  |
| Conductivity | 1764 | 05/09/03 | “ |  |  |
| Secondary Temp. | 2106 | 05/09/03 | “ |  |  |
| Secondary Cond. | 2128 | 15/05/03 | “ |  |  |
| Transmissometer | 333DR | 15/08/03 | IOS |  |  |
| Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/99 | Factory |  |  |