## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 31-May-2018 | Added DIC and Alkalinity data to 9 CHE casts. For details see document Carbon_Data_Addition.docx. |
| 08-Nov-2004 | The silicate and phosphate files were transposed in the rosette files so they were reprocessed from the original spreadsheet. J. L. |
| 30 Jan. 2004 | Error in transmissivity offset corrected using 2003-41-recal3.ccf |
| 17 Feb. 2004 | Error in calibration of bottle data corrected using 2003-41-recal4.ccf |

## PROCESSING NOTES

Cruise: 2003-41

Agency: OSAP

Location: Strait of Georgia / Juan de Fuca Strait

Project: SoG

Party Chief: Johannessen S.

Platform: Vector

Date: December 1, 2003 – December 6, 2003

Processed by: Germaine Gatien

Date of Processing: 2 January 2004 – 12 January 2004

Number of original CTD casts: 	68

Number of casts processed: 68

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Chelsea transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter (#1024) and a Seapoint Fluorometer (#2356) with a 10X cable. The deck unit was a model 911 (#0508). The salinometer was a Portasal model 8410 (#59724).

# SUMMARY OF QUALITY AND CONCERNS

There were problems in the numbering of samples. In a few cases bottles had no sample numbers on them. In one case no number is listed in the Daily log or rosette log, in other cases it is listed in the logs, but apparently was not written on the sample bottles. So the analyst assigned numbers. Unfortunately, the numbers chosen had already been used. In such cases either the rosette sheets should be checked before assigning numbers or numbers should be chosen that are highly unlikely to be duplicates, for example, 9991.

There was an altimeter mounted, but this is not listed on the first page of the Daily Log. It was mentioned in the comments.

There was an offset of -2db entered in the pressure calibration. This was inappropriate and led to negative pressures for values that were clearly in the water. This is easy to fix in processing, but may be confusing at sea. This could have been seen in the first cast in Saanich Inlet and it is recommended that operators get in the habit of checking this on the first cast.

The dissolved oxygen data in the CTD files should be considered ±0.6ml/l in the top 50m and ±0.2ml/l below that. The rosette data are mostly within ±0.2ml/l.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity and oxygen bottle calibration data was obtained. Flags and comments had been added as needed.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. Errors were corrected in the dates of the pressure, secondary conductivity and transmissometer calibrations. There was an offset of -2db entered in the con files used at sea; initial tests suggest that the pressures are too low by at least 1.5db, so the pressure offset was removed from the con file. The corrected file was named 2003-41-CTD.con.

There was an altimeter listed in the con file, but this was not listed in the appropriate area of the log book. It was mentioned in comments in the log and the channel does contain useful data.

There is no history available for the two conductivity sensors since they have not been used since the latest factory recalibration.

##### Conversion of Raw Data

The raw data was converted using the configuration file 2003-41-CTD.con.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close.

Rosette files were converted using a start time of -4s and duration of 5s. This should capture the records from 4s before firing to 1s after. In the past -2/5 was used but on a recent VECTOR cruise there was a problem with the CTD starting up so quickly that upcast data was captured in the rosette file.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

This step was skipped since the deck unit model is one of the newer ones that advances both conductivity channels by +0.073s. Tests will be done later to fine-tune the alignment.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.02,9), (0.03,9), (0.02,7) and (0.03,7). Leaving the data alone seems to provide the best results. The secondary salinity has a lot of fine-scale noise so it is hard to detect any change. This step will be skipped.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

The differences between sensors were generally very noisy so the figures that follow are rough averages:

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 1 | 290 | -0.0003 | -0.00075 | -0.0075 | ~.7, steady |
| 14 | 250 | -0.0006 | -0.0075 | -0.007 | Extremely noisy |
| 39 | 310 | -0.0003 | -0.0005 | -0.005 | Extremely noisy |
| 43 | 290 | -0.0004 | -0.0004 | -0.004 | Low, fairly noisy |
| 59 | 320 | -0.0002 | -0.00034 | -0.003 | Ok |
| 63 | 340 | -0.0002 | -0.00035 | -0.003 | Ok |

The differences in conductivity and salinity are larger than we usually see, especially for the early casts. But these casts are not very deep and are in areas of large variability, so this may not be particularly significant.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. The NMEA time in the SeaBird headers lacks a date, so the Convert routine selected the Start Time from the computer. It was noted in the log that casts #1-10 have the start time in Pacific Standard Time, so the first 10 casts were renamed IOS1, then the ADD TIME CHANNEL routine was used to change the start times to UTC. The output was named IOS, so all IOS files have UTC time.

The rosette files were converted to IOS files. The ADD TIME routine was used to change the time zone to UTC by adding 8 hours for cast #3 only. The output file from this was renamed BOT so that all BOT files have UTC times.

All BOT files were plotted and significant outliers were found for casts 17, 42, 53 and 64. CTDEDIT was used to clean the primary salinity for cast #53 and the secondary salinity for casts #17 and #42. Cast #64 proved to be very noisy for both salinity channels for bottle #1. This bottle should probably not be used for calibration purposes. No attempt was made to edit that data. The edited files were renamed *.BOT.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The cruise track was plotted and no problems noted.

A header summary and a header check were produced and one error in station name was found and corrected in the IOS and CLN files.

The average surface pressure is 2.4db with a range of 0.5db to 3.6db which seems reasonable. The mixed layer depth calculation shows only cast #54 having surface water sufficiently well-mixed to be useful for salinity calibration purposes.

All CLN files were examined on-screen for problems in T and S. The only problems noted (other than some spiking) were in casts #60 and 64 during which the secondary salinity and fluorescence looked bad during the upcasts, as though the pump was not working properly. There is no mention in the Daily Log of the CTD hitting bottom. The following cast looked ok in each case.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheets were converted to individual SAL files. There were no flagged values so no comments needed in the headers. The dissolved oxygen files (*.add) have comments in the headers to explain flagged values. There was no chlorophyll sampling.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. A few problems were found in the salinity data:

- One deep bottle had been taken at the bottom of cast #50, but no sample number was assigned. The analyst assigned the number #126, but that one had been used for cast #46 and in the spreadsheet the bottle salinity value from cast #50 was assigned to cast #46. After checking with the analyst, I changed the sample number to 9126 and changed the cast number and station name in the salinity spreadsheet.

- Two bottles fired during cast #66 had been given sample #210 and 211 on the rosette sheet, but the analyst assigned #198 and 199 because there were no #s on the sample bottles. Unfortunately those sample numbers had been used on the previous cast. The sample numbers were changed in the salinity spreadsheet file to match the rosette log records.

- Two bottles fired during cast #34 had been given sample #73 and 74 on the rosette sheet, but the analyst assigned #067 and 068 because there were no #s on the sample bottles. Unfortunately those sample numbers had been used on the previous cast. The sample numbers were changed in the salinity spreadsheet file to match the rosette log records.

- There was a salinity bottle shown as being from cast #24 (station 69), but there is no bottle file to go with it. In fact while the rosette log shows cast #24, the Daily log shows the bottle cast was #23 and the station name for cast #23 is station 69, so that makes sense. The salinity spreadsheet was changed to agree with the Daily log. The name of the dissolved oxygen file for that cast was renamed 2003-41-0023.ADD and the cast number changed within that file.

Problems were found in the dissolved oxygen data for casts # 29 and 42:

- For cast #29 – a value of -99.000 was substituted for 0.00 for a sample flagged “e”

- For cast #42 two sample numbers were wrong in the ADD file (the DOX values for samples #122 and #123 were said to be from #121 and #122. There was no sampling from #121.) This was corrected in the ADD file.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 11. COMPARE

Salinity

When a few outliers are excluded, the primary sensors shows no significant pressure or time-dependence and the average salinity difference is about 0.0001psu. The secondary salinity shows a little pressure and time-dependence and the salinity is low by about 0.003psu. Sample #196 from cast #64 looks like a severe outlier, but it appears to be the CTD data that is in doubt since the deepest bottle has salinity lower than the one above. The transmissivity is ~0. It is likely that the CTD hit bottom.

(See 2003-41-sal-comp.xls)

No salinity recalibration will be necessary for the primary sensors.

Dissolved Oxygen

The differences between the titrated values and CTD values were displayed versus pressure, cast # and titrated values. The best fit was found versus the titrated values. When points with residuals <-0.6 and >0.6 were excluded as well as a few other outliers, the following fit was found:

Titrated DOX = 1.1731 * DOX-CTD – 0.0289

This is a little out of line with other recent cruises such as 2003-27 (Line P – September 2003) for which it was found that

Titrated DOX = 1.1362 * DOX-CTD – 0.0143

However, this cruise sampled nothing deeper than 400db and no DOX lower than 2ml/l and the area is one of active mixing, so the differences may not be due to drift in the sensor.

The only severe outlier was already flagged. (See 2003-41-dox-comp1.xls.)

##### 12. SHIFT

Conductivity

Tests were run on the deep waters of casts #40 and 46 since they display the sort of variations needed to analyze the alignment. The results were examined in T-S space with the best results those that minimize unstable spiking without oversmoothing. The best results were found using an advancement of -0.4 records for the primary conductivity for a net advancements of about 0.056s, since the channel was already advanced by +0.073s by the deck unit. The secondary was not checked since it is unlikely to be archived. All casts were put through SHIFT using -0.4 records for the primary sensor.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was from 0.7 to 2.5s. A shift of 1 s (+24 records) was applied. This is the shift that has been used in most other cruises.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. ALIGNCTD was used to advance the oxygen by from 8 to 12s and the upcast and downcast traces examined. A choice of 8 or 9s makes the difference between upcast and downcast dissolved oxygen traces look most like those of the temperature traces in the top 200m and 10s looks better at depth. The downcast trace looks most like the upcast at the end of a bottle stop with a choice of about 8s. Aligning downcast features in T and O2 suggests a lower setting, but since the dissolved oxygen probably never has a chance to reach maximum and minimum values in distinctive features, this may be misleading. A choice of about +8s seems best overall. Alignment will be done using SHIFT in IOS SHELL to advance the conductivity by +200 records.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing. They are closer to the bottles and show little drift with time or pressure.

Page plots were produced using (T0,S0). These were used to guide the editing. The descent rate was quiet and often very low in the early part of the cruise. For casts #13-20 and 35-54 it was very noisy and the average descent rate was often much lower than 1m/s. DELETE plus CTDEDIT removed much of the data corrupted by slowdowns or reversals of the CTD, but in this region editing is limited by the difficulty in telling good data from bad. Only if the data is clearly bad was it removed.

The following casts required light editing: #1-5, 12,21,22,24,25,27-29,31-34,37,44,48,55-61,63,64,66-68.

The following casts were edited more extensively: # 13-20, 35,36,38-43,45,46,49-54.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – Both conductivity sensors have been recalibrated since their last use so there is no history available.

Historic ranges – All data fell within the historic ranges except for 4 casts in Juan de Fuca Strait for which the salinity was low; this is probably due to real conditions rather than a calibration problem.

##### Initial Recalibration

No recalibration was needed for the primary salinity.

File 2003-41-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun and the corrections were found to be made correctly. (See 2003-41-dox-comp2.xls.)

### Final Dissolved Oxygen Recalibration

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but there remain significant errors due to the response-time problem. The best we can do to correct for this is try to make the downcast CTD data match the bottles from the upcast. A set of downcast files were prepared by metre-averaging the COR1 files, thinning the data to the depths generally used for bottles and reversing them. COMPARE was run comparing those files with the bottles in the MRG files. A fit of differences versus pressure was found and file 2003-41-recal2.ccf was prepared applying the following correction:

DOX(corrected) = DOX (after 1st recal) – 0.0503 +1*E-4 * Pressure

File 2003-41-recal2.ccf was used to apply a second recalibration to the dissolved oxygen channel in the COR1, MRGCOR1 and SAMCOR1 files. (Output: COR2, MRGCOR2 and SAMCOR2)***

It was also applied to the reversed thinned downcast data and COMPARE was rerun to check that the second recalibration was done properly. The results show that the DOX errors mostly fall within ±0.6ml/l in the top 50m and within ±0.2ml/l below that. (See 2003-41-dox-comp4.xls.)

February 17, 2004: The 2nd DOX calibration should not have been applied to the bottle files since the wait before firing means there is no time-response problem. To “undo” the recalibration file 2003-41-recal4.ccf was applied to the CHE files. [DOX = DOX(with error) +0.0503 -1E-4*Pressure]

##### Special Fluorometer Processing

The EDT files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: BOX)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, and Transmissivity profiles. Separate profile plots were prepared with DOX and FL versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment, fix formats and channel names using file 2003-41-header.hdr. The Standards Check routine was run and no problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRG files were put through CLEAN to remove the SeaBird headers, REVERSE to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1-10. Local time entered in START time in files; correct time in NMEA time in files and in Daily Log

5. CTD hit bottom

12. Rosette hit stern of ship (Bottle #3)

15. Cast aborted part-way through due to ship drift.

# 64. CTD may have hit bottom.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2003-41

Dates: Start: December 1, 2003 End: December 6, 2003

Location: Strait of Georgia / Juan de Fuca Strait

Vessel: Vector

Party Chief: Johannessen S.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2003-41

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2095 | 16/07/03 | Factory |  |  |
| Conductivity | 1764 | 24/07/03 | “ |  |  |
| Secondary Temp. | 2106 | 23/01/03 | “ |  |  |
| Secondary Cond. | 2128 | 23/01/03 | “ |  |  |
| Transmissometer | 333DR | 18/08/03 | IOS |  |  |
| Dissolved Oxygen | 0047 | 21/08/02 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 77511 | 30/09/99 | Factory |  |  |