## File List
Number of Files= |  | 8
Pair # | Primary File | Secondary File | First Row in Main | Last Row in Main | Comparison
1 | 2004-05-0001 | 2004-05-0001 | 11 | 22 | 2004-05-0001 & 2004-05-0001
2 | 2004-05-0003 | 2004-05-0003 | 24 | 31 | 2004-05-0003 & 2004-05-0003
3 | 2004-05-0007 | 2004-05-0007 | 33 | 51 | 2004-05-0007 & 2004-05-0007
4 | 2004-05-0013 | 2004-05-0013 | 53 | 72 | 2004-05-0013 & 2004-05-0013
5 | 2004-05-0018 | 2004-05-0018 | 74 | 96 | 2004-05-0018 & 2004-05-0018
6 | 2004-05-0025 | 2004-05-0025 | 98 | 121 | 2004-05-0025 & 2004-05-0025
7 | 2004-05-0032 | 2004-05-0032 | 123 | 146 | 2004-05-0032 & 2004-05-0032
8 | 2004-05-0047 | 2004-05-0047 | 148 | 171 | 2004-05-0047 & 2004-05-0047

## (DOX-CTD - DOX-BOT)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
1
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=Pressure
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00060363 |  | X | CALC Y
m= | -9.4109e-06 | 3.4527e-05 |  | Std Dev of Resid |  | 0.15213 |  | 5.088 | 0.0069981
b= | 0.007046 | 0.01777 |  | Resid Sum of Squares |  | 2.8468 |  | 1499 | -0.0070609
# of Values Extracted = |  |  | 154
# of Values Used in Fit = |  |  | 125
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | Pressure
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 222.372 | 0.043 | 0 | 0.043 | 0.0049533 | -0.038047 |  | 222.372 | 0.043
1 | YES | 200.997 | 0.046 | 0 | 0.046 | 0.0051545 | -0.040846 |  | 200.997 | 0.046
1 | YES | 175.992 | 0.049 | 0 | 0.049 | 0.0053898 | -0.04361 |  | 175.992 | 0.049
1 | YES | 175.992 | 0.049 | 0 | 0.049 | 0.0053898 | -0.04361 |  | 175.992 | 0.049
1 | YES | 150.999 | 0.051 | 0 | -0.001 | 0.005625 | 0.006625 |  | 150.999 | -0.001
1 | YES | 126.004 | 0.0745 | 0.006364 | 0.0225 | 0.0058602 | -0.01664 |  | 126.004 | 0.0225
1 | YES | 101.007 | 1.1185 | 0.074246 | 0.0975 | 0.0060954 | -0.091405 |  | 101.007 | 0.0975
1 | YES | 75.9925 | 4.438 | 0.11597 | 0.162 | 0.0063309 | -0.15567 |  | 75.9925 | 0.162
1 | YES | 51.0615 | 4.389 | 0.08061 | 0.126 | 0.0065655 | -0.11943 |  | 51.0615 | 0.126
1 | YES | 25.9955 | 5.2975 | 0.065761 | 0.1165 | 0.0068014 | -0.1097 |  | 25.9955 | 0.1165
1 | YES | 10.4907 | 5.423 | 0.14636 | 0.018 | 0.0069473 | -0.011053 |  | 10.4907 | 0.018
1 | YES | 7.262 | 5.733 | 0 | 0.054 | 0.0069777 | -0.047022 |  | 7.262 | 0.054
2 | NO | 0 | 0 | 0 | -4.676 | 0.007046 | 4.683 |  | 101.03 | -0.085
2 | YES | 101.03 | 4.145 | 0.0056571 | -0.085 | 0.0060952 | 0.091095 |  | 75.9635 | -0.1155
2 | YES | 75.9635 | 4.5955 | 0.041719 | -0.1155 | 0.0063311 | 0.12183 |  | 75.9635 | -0.1115
2 | YES | 75.9635 | 4.5955 | 0.041719 | -0.1115 | 0.0063311 | 0.11783 |  | 51.007 | -0.339
2 | YES | 51.007 | 5.568 | 0.074953 | -0.339 | 0.006566 | 0.34557 |  | 25.9815 | 0.046
2 | YES | 25.9815 | 6.64 | 0.0028286 | 0.046 | 0.0068015 | -0.039199 |  | 10.3832 | 0.0162
2 | YES | 10.3832 | 6.7162 | 0.0054037 | 0.0162 | 0.0069483 | -0.0092518 |  | 5.088 | -0.013
2 | YES | 5.088 | 6.757 | 0.011533 | -0.013 | 0.0069981 | 0.019998 |  | 1324.86 | -0.0075
3 | YES | 1324.86 | 0.4605 | 0.0021213 | -0.0075 | -0.0054221 | 0.0020779 |  | 1250.98 | -0.016
3 | YES | 1250.98 | 0.364 | 0 | -0.016 | -0.0047268 | 0.011273 |  | 1000.99 | 0.005
3 | YES | 1000.99 | 0.232 | 0 | 0.005 | -0.0023742 | -0.0073742 |  | 1000.99 | -0.01
3 | YES | 1000.99 | 0.232 | 0 | -0.01 | -0.0023742 | 0.0076258 |  | 802.038 | -0.005
3 | YES | 802.038 | 0.236 | 0 | -0.005 | -0.00050186 | 0.0044981 |  | 600.988 | -0.002
3 | YES | 600.988 | 0.445 | 0.0014142 | -0.002 | 0.0013902 | 0.0033902 |  | 402.992 | -0.002
3 | YES | 402.992 | 1.271 | 0 | -0.002 | 0.0032535 | 0.0052535 |  | 300.996 | 0.039
3 | YES | 300.996 | 1.732 | 0 | 0.039 | 0.0042134 | -0.034787 |  | 250.996 | 0.061
3 | YES | 250.996 | 2.09 | 0.0056569 | 0.061 | 0.0046839 | -0.056316 |  | 200.962 | -0.0335
3 | YES | 200.962 | 2.5065 | 0.15486 | -0.0335 | 0.0051548 | 0.038655 |  | 176.021 | 0.0065002
3 | YES | 176.021 | 3.0135 | 0.020506 | 0.0065002 | 0.0053895 | -0.0011107 |  | 148.993 | 0.169
3 | YES | 148.993 | 3.716 | 0.041012 | 0.169 | 0.0056439 | -0.16336 |  | 126.137 | 0.04
3 | YES | 126.137 | 4.407 | 0.057983 | 0.04 | 0.005859 | -0.034141 |  | 102.018 | 0.136
3 | YES | 102.018 | 4.891 | 0 | 0.136 | 0.0060859 | -0.12991 |  | 76.0105 | 0.0175
3 | YES | 76.0105 | 6.6175 | 0.041719 | 0.0175 | 0.0063307 | -0.011169 |  | 51.0015 | 0.0515
3 | YES | 51.0015 | 6.6645 | 0.0021212 | 0.0515 | 0.006566 | -0.044934 |  | 25.9945 | 0.00099993
3 | YES | 25.9945 | 6.677 | 0.0014141 | 0.00099993 | 0.0068014 | 0.0058015 |  | 10.5122 | 0.04475
3 | YES | 10.5122 | 6.6878 | 0.0092151 | 0.04475 | 0.0069471 | -0.037803 |  | 5.9755 | -0.016
3 | YES | 5.9755 | 6.642 | 0.070711 | -0.016 | 0.0069898 | 0.02299 |  | 1248.49 | -0.008
4 | NO | 0 | 0 | 0 | -1.677 | 0.007046 | 1.684 |  | 1251 | -0.017
4 | NO | 1500.99 | 0.695 | 0 | -0.06 | -0.0070796 | 0.05292 |  | 1000.98 | -0.0105
4 | YES | 1248.49 | 0.408 | 0.0028284 | -0.008 | -0.0047034 | 0.0032967 |  | 800.987 | -0.0185
4 | YES | 1251 | 0.41 | 0 | -0.017 | -0.004727 | 0.012273 |  | 601 | 0.0055
4 | YES | 1000.98 | 0.2465 | 0.0007071 | -0.0105 | -0.0023741 | 0.0081259 |  | 401.013 | 0.114
4 | YES | 800.987 | 0.2265 | 0.00070711 | -0.0185 | -0.00049197 | 0.018008 |  | 300.982 | 0.245
4 | YES | 601 | 0.3995 | 0.0035355 | 0.0055 | 0.0013901 | -0.0041099 |  | 249.005 | 0.2325
4 | YES | 401.013 | 1.394 | 0.0028285 | 0.114 | 0.0032721 | -0.11073 |  | 201.003 | -0.468
4 | YES | 300.982 | 2.479 | 0.018385 | 0.245 | 0.0042135 | -0.24079 |  | 175.991 | -0.3835
4 | YES | 249.005 | 3.0075 | 0.016264 | 0.2325 | 0.0047027 | -0.2278 |  | 150.991 | -0.3155
4 | YES | 201.003 | 2.892 | 0.059397 | -0.468 | 0.0051544 | 0.47315 |  | 125.986 | -0.1645
4 | YES | 175.991 | 3.4985 | 0.16334 | -0.3835 | 0.0053898 | 0.38889 |  | 100.993 | 0.0385
4 | YES | 150.991 | 3.3255 | 0.017678 | -0.3155 | 0.0056251 | 0.32113 |  | 76.034 | 0.184
4 | YES | 125.986 | 3.6725 | 0.054447 | -0.1645 | 0.0058604 | 0.17036 |  | 51.019 | 0.013
4 | YES | 100.993 | 4.3345 | 0.033234 | 0.0385 | 0.0060956 | -0.032404 |  | 24.017 | 0.017
4 | YES | 76.034 | 6.729 | 0.0098995 | 0.184 | 0.0063305 | -0.17767 |  | 10.9907 | 0.017
4 | YES | 51.019 | 6.762 | 0.0042427 | 0.013 | 0.0065659 | -0.0064341 |  | 6.0625 | 0.015
4 | YES | 24.017 | 6.732 | 0.0042427 | 0.017 | 0.00682 | -0.01018 |  | 1498.99 | -0.0435
4 | YES | 10.9907 | 6.736 | 0.0052914 | 0.017 | 0.0069426 | -0.010058 |  | 1251 | -0.00099999
4 | YES | 6.0625 | 6.74 | 0.0028286 | 0.015 | 0.006989 | -0.0080109 |  | 1000.99 | 0.0015
5 | NO | 0 | 0 | 0 | -2.484 | 0.007046 | 2.491 |  | 800.986 | 0.0155
5 | NO | 0 | 0 | 0 | -2.244 | 0.007046 | 2.251 |  | 598.997 | -0.041
5 | NO | 0 | 0 | 0 | -1.931 | 0.007046 | 1.938 |  | 398.978 | -0.0089999
5 | NO | 0 | 0 | 0 | -1.405 | 0.007046 | 1.412 |  | 299.003 | 0.0565
5 | YES | 1498.99 | 0.6005 | 0.0021213 | -0.0435 | -0.0070608 | 0.036439 |  | 249.992 | 0.068
5 | NO | 1501 | 0.6035 | 0.0021213 | -0.0395 | -0.0070797 | 0.03242 |  | 201.005 | 0.227
5 | YES | 1251 | 0.337 | 0.0014142 | -0.00099999 | -0.004727 | -0.003727 |  | 173.996 | 0.384
5 | YES | 1000.99 | 0.2555 | 0.00070712 | 0.0015 | -0.0023742 | -0.0038742 |  | 150.974 | 0.178
5 | YES | 800.986 | 0.3465 | 0.0021213 | 0.0155 | -0.00049196 | -0.015992 |  | 126.002 | 0.137
5 | YES | 598.997 | 0.613 | 0.0014142 | -0.041 | 0.0014089 | 0.042409 |  | 98.98 | 0.0215
5 | YES | 398.978 | 1.502 | 0.012728 | -0.0089999 | 0.0032913 | 0.012291 |  | 75.9945 | 0.02
5 | YES | 299.003 | 2.4445 | 0.0035356 | 0.0565 | 0.0042321 | -0.052268 |  | 48.9815 | 0.0089998
5 | YES | 249.992 | 3.21 | 0 | 0.068 | 0.0046934 | -0.063307 |  | 24.0075 | -0.012
5 | YES | 201.005 | 4.405 | 0.02687 | 0.227 | 0.0051544 | -0.22185 |  | 10.483 | 0.01325
5 | YES | 173.996 | 4.916 | 0.014142 | 0.384 | 0.0054086 | -0.37859 |  | 6.2655 | 0.0019999
5 | YES | 150.974 | 5.061 | 0.018385 | 0.178 | 0.0056252 | -0.17237 |  | 1252.01 | -0.008
5 | YES | 126.002 | 5.309 | 0.024042 | 0.137 | 0.0058602 | -0.13114 |  | 1001 | 0.0145
5 | YES | 98.98 | 6.7495 | 0.0035353 | 0.0215 | 0.0061145 | -0.015386 |  | 800.999 | 0.0515
5 | YES | 75.9945 | 6.756 | 0.0014141 | 0.02 | 0.0063308 | -0.013669 |  | 601.003 | 0.0105
5 | YES | 48.9815 | 6.759 | 0.0070709 | 0.0089998 | 0.0065851 | -0.0024148 |  | 401.01 | 0.03
5 | YES | 24.0075 | 6.765 | 0.0028286 | -0.012 | 0.0068201 | 0.01882 |  | 298.998 | -0.0255
5 | YES | 10.483 | 6.7893 | 0.0012582 | 0.01325 | 0.0069474 | -0.0063025 |  | 252.008 | 0.019
5 | YES | 6.2655 | 6.791 | 0.0014141 | 0.0019999 | 0.006987 | 0.0049872 |  | 199 | 0.5805
6 | NO | 0 | 0 | 0 | -3.102 | 0.007046 | 3.109 |  | 176 | 0.2605
6 | NO | 0 | 0 | 0 | -2.758 | 0.007046 | 2.765 |  | 150.959 | -0.087
6 | NO | 0 | 0 | 0 | -2.419 | 0.007046 | 2.426 |  | 125.974 | -0.0345
6 | NO | 0 | 0 | 0 | -1.964 | 0.007046 | 1.971 |  | 100.98 | -0.436
6 | NO | 0 | 0 | 0 | -1.347 | 0.007046 | 1.354 |  | 75.992 | -0.0505
6 | NO | 0 | 0 | 0 | -1.548 | 0.007046 | 1.555 |  | 50.95 | -0.0075002
6 | NO | 1500.98 | 0.5325 | 0.0035356 | -0.0095 | -0.0070795 | 0.0024205 |  | 25.9585 | 0.0090003
6 | YES | 1252.01 | 0.294 | 0 | -0.008 | -0.0047365 | 0.0032635 |  | 10.4867 | -0.0325
6 | YES | 1001 | 0.2565 | 0.0007071 | 0.0145 | -0.0023743 | -0.016874 |  | 5.464 | 0.063667
6 | YES | 800.999 | 0.3775 | 0.0021213 | 0.0515 | -0.00049208 | -0.051992 |  | 1499 | -0.0415
6 | YES | 601.003 | 0.6815 | 0.0007071 | 0.0105 | 0.0013901 | -0.0091099 |  | 1251 | -0.017
6 | YES | 401.01 | 1.496 | 0.0070711 | 0.03 | 0.0032722 | -0.026728 |  | 1000.86 | -0.0095
6 | YES | 298.998 | 1.9325 | 0.012021 | -0.0255 | 0.0042322 | 0.029732 |  | 798.99 | 0.012
6 | YES | 252.008 | 2.615 | 0 | 0.019 | 0.0046744 | -0.014326 |  | 600.904 | -0.015
6 | YES | 199 | 3.6975 | 0.024749 | 0.5805 | 0.0051732 | -0.57533 |  | 399.035 | -0.0325
6 | YES | 176 | 4.2125 | 0.053033 | 0.2605 | 0.0053897 | -0.25511 |  | 298.982 | -0.181
6 | YES | 150.959 | 5.205 | 0.043841 | -0.087 | 0.0056254 | 0.092625 |  | 248.987 | -0.0585
6 | YES | 125.974 | 5.6915 | 0.0091924 | -0.0345 | 0.0058605 | 0.040361 |  | 198.99 | -0.1335
6 | YES | 100.98 | 6.067 | 0.093338 | -0.436 | 0.0060957 | 0.4421 |  | 175.96 | -0.0365
6 | YES | 75.992 | 6.7835 | 0.0021212 | -0.0505 | 0.0063309 | 0.056831 |  | 149.007 | -0.2965
6 | YES | 50.95 | 6.8395 | 0.0021215 | -0.0075002 | 0.0065665 | 0.014067 |  | 123.989 | -0.3125
6 | YES | 25.9585 | 6.852 | 0.0014145 | 0.0090003 | 0.0068017 | -0.0021986 |  | 101 | -0.1705
6 | YES | 10.4867 | 6.8185 | 0.013128 | -0.0325 | 0.0069473 | 0.039447 |  | 75.967 | -0.2585
6 | YES | 5.464 | 6.8407 | 0.0050333 | 0.063667 | 0.0069946 | -0.056672 |  | 51.019 | -0.088
7 | NO | 0 | 0 | 0 | -3.247 | 0.007046 | 3.254 |  | 24.002 | -0.0955
7 | NO | 0 | 0 | 0 | -2.898 | 0.007046 | 2.905 |  | 9.475 | -0.07925
7 | NO | 0 | 0 | 0 | -2.888 | 0.007046 | 2.895 |  | 1251 | 0.0075
7 | NO | 0 | 0 | 0 | -2.534 | 0.007046 | 2.541 |  | 1002 | -0.032
7 | NO | 0 | 0 | 0 | -1.945 | 0.007046 | 1.952 |  | 798.983 | -0.0305
7 | NO | 0 | 0 | 0 | -1.266 | 0.007046 | 1.273 |  | 602.018 | -0.035
7 | YES | 1499 | 0.5115 | 0.0021213 | -0.0415 | -0.0070609 | 0.034439 |  | 401.005 | -0.045
7 | YES | 1251 | 0.335 | 0 | -0.017 | -0.004727 | 0.012273 |  | 299.009 | 0.018
7 | YES | 1000.86 | 0.2945 | 0.0007071 | -0.0095 | -0.0023729 | 0.0071271 |  | 251.027 | 0.0165
7 | YES | 798.99 | 0.376 | 0.0014142 | 0.012 | -0.00047317 | -0.012473 |  | 200.984 | -0.036
7 | YES | 600.904 | 0.591 | 0.0014142 | -0.015 | 0.001391 | 0.016391 |  | 174.033 | 0.115
7 | YES | 399.035 | 1.2315 | 0.0035355 | -0.0325 | 0.0032907 | 0.035791 |  | 151.99 | 0.235
7 | YES | 298.982 | 1.678 | 0.0098995 | -0.181 | 0.0042323 | 0.18523 |  | 123.996 | 0.5395
7 | YES | 248.987 | 2.1985 | 0.019092 | -0.0585 | 0.0047028 | 0.063203 |  | 100.994 | 0.495
7 | YES | 198.99 | 2.8865 | 0.037477 | -0.1335 | 0.0051733 | 0.13867 |  | 73.9945 | -0.0825
7 | YES | 175.96 | 3.4045 | 0.0021212 | -0.0365 | 0.0053901 | 0.04189 |  | 48.9945 | -0.1345
7 | YES | 149.007 | 4.0155 | 0.0021215 | -0.2965 | 0.0056437 | 0.30214 |  | 23.9955 | -0.1215
7 | YES | 123.989 | 4.9175 | 0.072832 | -0.3125 | 0.0058792 | 0.31838 |  | 9.47 | -0.08675
7 | YES | 101 | 5.5875 | 0.044547 | -0.1705 | 0.0060955 | 0.1766 |  | 5.3487 | -0.058
7 | YES | 75.967 | 6.6935 | 0.017677 | -0.2585 | 0.0063311 | 0.26483
7 | YES | 51.019 | 6.878 | 0.0014145 | -0.088 | 0.0065659 | 0.094566
7 | YES | 24.002 | 6.8615 | 0.019092 | -0.0955 | 0.0068201 | 0.10232
7 | YES | 9.475 | 6.8717 | 0.0044253 | -0.07925 | 0.0069568 | 0.086207
7 | NO | 4.081 | 6.8715 | 0.0035356 | -0.0645 | 0.0070076 | 0.071508
8 | NO | 0 | 0 | 0 | -3.644 | 0.007046 | 3.651
8 | NO | 0 | 0 | 0 | -2.99 | 0.007046 | 2.997
8 | NO | 0 | 0 | 0 | -2.578 | 0.007046 | 2.585
8 | NO | 0 | 0 | 0 | -2.025 | 0.007046 | 2.032
8 | NO | 0 | 0 | 0 | -2.012 | 0.007046 | 2.019
8 | NO | 0 | 0 | 0 | -1.381 | 0.007046 | 1.388
8 | NO | 1501.02 | 0.631 | 0.0014142 | 0.001 | -0.0070799 | -0.0080799
8 | YES | 1251 | 0.4365 | 0.00070712 | 0.0075 | -0.004727 | -0.012227
8 | YES | 1002 | 0.397 | 0 | -0.032 | -0.0023837 | 0.029616
8 | YES | 798.983 | 0.4485 | 0.0007071 | -0.0305 | -0.00047311 | 0.030027
8 | YES | 602.018 | 0.555 | 0 | -0.035 | 0.0013805 | 0.036381
8 | YES | 401.005 | 0.795 | 0.0014142 | -0.045 | 0.0032722 | 0.048272
8 | YES | 299.009 | 1.139 | 0.0056569 | 0.018 | 0.0042321 | -0.013768
8 | YES | 251.027 | 1.5185 | 0.0035355 | 0.0165 | 0.0046836 | -0.011816
8 | YES | 200.984 | 2.085 | 0.028284 | -0.036 | 0.0051546 | 0.041155
8 | YES | 174.033 | 2.61 | 0.02687 | 0.115 | 0.0054082 | -0.10959
8 | YES | 151.99 | 3.18 | 0 | 0.235 | 0.0056157 | -0.22938
8 | YES | 123.996 | 4.9005 | 0.1096 | 0.5395 | 0.0058791 | -0.53362
8 | YES | 100.994 | 6.456 | 0.14425 | 0.495 | 0.0060956 | -0.4889
8 | YES | 73.9945 | 6.9515 | 0.0035356 | -0.0825 | 0.0063497 | 0.08885
8 | YES | 48.9945 | 6.9275 | 0.0063642 | -0.1345 | 0.0065849 | 0.14108
8 | YES | 23.9955 | 6.9435 | 0.0049497 | -0.1215 | 0.0068202 | 0.12832
8 | YES | 9.47 | 6.9822 | 0.0041932 | -0.08675 | 0.0069569 | 0.093706
8 | YES | 5.3487 | 6.998 | 0.00099993 | -0.058 | 0.0069957 | 0.064996
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## (DOX-CTD - DOX-BOT) (2)
Calculation of Linear Least Squares Fit For
(DOX-CTD - DOX-BOT)
2
Fit Equation: Diff = m * Prop + b
where: Diff =Diff (DOX-CTD - DOX-BOT), Prop=DOX-CTD
Coefficients |  |  |  | Error Estimates |  |  |  | Regression Line Limits
| Value | Std Err. |  | Coeff of Determination |  | 0.00027848 |  | X | CALC Y
m= | -0.00097215 | 0.0052519 |  | Std Dev of Resid |  | 0.15216 |  | 0.043 | 0.0072764
b= | 0.0073182 | 0.022805 |  | Resid Sum of Squares |  | 2.8477 |  | 6.998 | 0.00051507
# of Values Extracted = |  |  | 154
# of Values Used in Fit = |  |  | 125
File | Use | Pressure | DOX-CTD | DOX-CTD | Diff (DOX-CTD - DOX-BOT) |  |  |  | Fit vs. | DOX-CTD
Pair # | in Fit? | Average | Average | Std Dev | Measured | Fit | Resid |  | X | Y
1 | YES | 222.372 | 0.043 | 0 | 0.043 | 0.0072764 | -0.035724 |  | 0.043 | 0.043
1 | YES | 200.997 | 0.046 | 0 | 0.046 | 0.0072734 | -0.038727 |  | 0.046 | 0.046
1 | YES | 175.992 | 0.049 | 0 | 0.049 | 0.0072705 | -0.041729 |  | 0.049 | 0.049
1 | YES | 175.992 | 0.049 | 0 | 0.049 | 0.0072705 | -0.041729 |  | 0.049 | 0.049
1 | YES | 150.999 | 0.051 | 0 | -0.001 | 0.0072686 | 0.0082686 |  | 0.051 | -0.001
1 | YES | 126.004 | 0.0745 | 0.006364 | 0.0225 | 0.0072457 | -0.015254 |  | 0.0745 | 0.0225
1 | YES | 101.007 | 1.1185 | 0.074246 | 0.0975 | 0.0062308 | -0.091269 |  | 1.1185 | 0.0975
1 | YES | 75.9925 | 4.438 | 0.11597 | 0.162 | 0.0030038 | -0.159 |  | 4.438 | 0.162
1 | YES | 51.0615 | 4.389 | 0.08061 | 0.126 | 0.0030514 | -0.12295 |  | 4.389 | 0.126
1 | YES | 25.9955 | 5.2975 | 0.065761 | 0.1165 | 0.0021682 | -0.11433 |  | 5.2975 | 0.1165
1 | YES | 10.4907 | 5.423 | 0.14636 | 0.018 | 0.0020462 | -0.015954 |  | 5.423 | 0.018
1 | YES | 7.262 | 5.733 | 0 | 0.054 | 0.0017448 | -0.052255 |  | 5.733 | 0.054
2 | NO | 0 | 0 | 0 | -4.676 | 0.0073182 | 4.6833 |  | 4.145 | -0.085
2 | YES | 101.03 | 4.145 | 0.0056571 | -0.085 | 0.0032886 | 0.088289 |  | 4.5955 | -0.1155
2 | YES | 75.9635 | 4.5955 | 0.041719 | -0.1155 | 0.0028507 | 0.11835 |  | 4.5955 | -0.1115
2 | YES | 75.9635 | 4.5955 | 0.041719 | -0.1115 | 0.0028507 | 0.11435 |  | 5.568 | -0.339
2 | YES | 51.007 | 5.568 | 0.074953 | -0.339 | 0.0019052 | 0.34091 |  | 6.64 | 0.046
2 | YES | 25.9815 | 6.64 | 0.0028286 | 0.046 | 0.0008631 | -0.045137 |  | 6.7162 | 0.0162
2 | YES | 10.3832 | 6.7162 | 0.0054037 | 0.0162 | 0.00078902 | -0.015411 |  | 6.757 | -0.013
2 | YES | 5.088 | 6.757 | 0.011533 | -0.013 | 0.00074936 | 0.013749 |  | 0.4605 | -0.0075
3 | YES | 1324.86 | 0.4605 | 0.0021213 | -0.0075 | 0.0068705 | 0.01437 |  | 0.364 | -0.016
3 | YES | 1250.98 | 0.364 | 0 | -0.016 | 0.0069643 | 0.022964 |  | 0.232 | 0.005
3 | YES | 1000.99 | 0.232 | 0 | 0.005 | 0.0070926 | 0.0020926 |  | 0.232 | -0.01
3 | YES | 1000.99 | 0.232 | 0 | -0.01 | 0.0070926 | 0.017093 |  | 0.236 | -0.005
3 | YES | 802.038 | 0.236 | 0 | -0.005 | 0.0070887 | 0.012089 |  | 0.445 | -0.002
3 | YES | 600.988 | 0.445 | 0.0014142 | -0.002 | 0.0068856 | 0.0088856 |  | 1.271 | -0.002
3 | YES | 402.992 | 1.271 | 0 | -0.002 | 0.0060826 | 0.0080825 |  | 1.732 | 0.039
3 | YES | 300.996 | 1.732 | 0 | 0.039 | 0.0056344 | -0.033366 |  | 2.09 | 0.061
3 | YES | 250.996 | 2.09 | 0.0056569 | 0.061 | 0.0052864 | -0.055714 |  | 2.5065 | -0.0335
3 | YES | 200.962 | 2.5065 | 0.15486 | -0.0335 | 0.0048815 | 0.038381 |  | 3.0135 | 0.0065002
3 | YES | 176.021 | 3.0135 | 0.020506 | 0.0065002 | 0.0043886 | -0.0021116 |  | 3.716 | 0.169
3 | YES | 148.993 | 3.716 | 0.041012 | 0.169 | 0.0037057 | -0.16529 |  | 4.407 | 0.04
3 | YES | 126.137 | 4.407 | 0.057983 | 0.04 | 0.0030339 | -0.036966 |  | 4.891 | 0.136
3 | YES | 102.018 | 4.891 | 0 | 0.136 | 0.0025634 | -0.13344 |  | 6.6175 | 0.0175
3 | YES | 76.0105 | 6.6175 | 0.041719 | 0.0175 | 0.00088497 | -0.016615 |  | 6.6645 | 0.0515
3 | YES | 51.0015 | 6.6645 | 0.0021212 | 0.0515 | 0.00083928 | -0.050661 |  | 6.677 | 0.00099993
3 | YES | 25.9945 | 6.677 | 0.0014141 | 0.00099993 | 0.00082713 | -0.0001728 |  | 6.6878 | 0.04475
3 | YES | 10.5122 | 6.6878 | 0.0092151 | 0.04475 | 0.00081668 | -0.043933 |  | 6.642 | -0.016
3 | YES | 5.9755 | 6.642 | 0.070711 | -0.016 | 0.00086116 | 0.016861 |  | 0.408 | -0.008
4 | NO | 0 | 0 | 0 | -1.677 | 0.0073182 | 1.6843 |  | 0.41 | -0.017
4 | NO | 1500.99 | 0.695 | 0 | -0.06 | 0.0066425 | 0.066643 |  | 0.2465 | -0.0105
4 | YES | 1248.49 | 0.408 | 0.0028284 | -0.008 | 0.0069215 | 0.014922 |  | 0.2265 | -0.0185
4 | YES | 1251 | 0.41 | 0 | -0.017 | 0.0069196 | 0.02392 |  | 0.3995 | 0.0055
4 | YES | 1000.98 | 0.2465 | 0.0007071 | -0.0105 | 0.0070785 | 0.017579 |  | 1.394 | 0.114
4 | YES | 800.987 | 0.2265 | 0.00070711 | -0.0185 | 0.007098 | 0.025598 |  | 2.479 | 0.245
4 | YES | 601 | 0.3995 | 0.0035355 | 0.0055 | 0.0069298 | 0.0014298 |  | 3.0075 | 0.2325
4 | YES | 401.013 | 1.394 | 0.0028285 | 0.114 | 0.005963 | -0.10804 |  | 2.892 | -0.468
4 | YES | 300.982 | 2.479 | 0.018385 | 0.245 | 0.0049082 | -0.24009 |  | 3.4985 | -0.3835
4 | YES | 249.005 | 3.0075 | 0.016264 | 0.2325 | 0.0043944 | -0.22811 |  | 3.3255 | -0.3155
4 | YES | 201.003 | 2.892 | 0.059397 | -0.468 | 0.0045067 | 0.47251 |  | 3.6725 | -0.1645
4 | YES | 175.991 | 3.4985 | 0.16334 | -0.3835 | 0.0039171 | 0.38742 |  | 4.3345 | 0.0385
4 | YES | 150.991 | 3.3255 | 0.017678 | -0.3155 | 0.0040853 | 0.31959 |  | 6.729 | 0.184
4 | YES | 125.986 | 3.6725 | 0.054447 | -0.1645 | 0.0037479 | 0.16825 |  | 6.762 | 0.013
4 | YES | 100.993 | 4.3345 | 0.033234 | 0.0385 | 0.0031044 | -0.035395 |  | 6.732 | 0.017
4 | YES | 76.034 | 6.729 | 0.0098995 | 0.184 | 0.00077658 | -0.18322 |  | 6.736 | 0.017
4 | YES | 51.019 | 6.762 | 0.0042427 | 0.013 | 0.0007445 | -0.012256 |  | 6.74 | 0.015
4 | YES | 24.017 | 6.732 | 0.0042427 | 0.017 | 0.00077366 | -0.016227 |  | 0.6005 | -0.0435
4 | YES | 10.9907 | 6.736 | 0.0052914 | 0.017 | 0.00076977 | -0.01623 |  | 0.337 | -0.00099999
4 | YES | 6.0625 | 6.74 | 0.0028286 | 0.015 | 0.00076589 | -0.014234 |  | 0.2555 | 0.0015
5 | NO | 0 | 0 | 0 | -2.484 | 0.0073182 | 2.4913 |  | 0.3465 | 0.0155
5 | NO | 0 | 0 | 0 | -2.244 | 0.0073182 | 2.2513 |  | 0.613 | -0.041
5 | NO | 0 | 0 | 0 | -1.931 | 0.0073182 | 1.9383 |  | 1.502 | -0.0089999
5 | NO | 0 | 0 | 0 | -1.405 | 0.0073182 | 1.4123 |  | 2.4445 | 0.0565
5 | YES | 1498.99 | 0.6005 | 0.0021213 | -0.0435 | 0.0067344 | 0.050234 |  | 3.21 | 0.068
5 | NO | 1501 | 0.6035 | 0.0021213 | -0.0395 | 0.0067315 | 0.046231 |  | 4.405 | 0.227
5 | YES | 1251 | 0.337 | 0.0014142 | -0.00099999 | 0.0069906 | 0.0079905 |  | 4.916 | 0.384
5 | YES | 1000.99 | 0.2555 | 0.00070712 | 0.0015 | 0.0070698 | 0.0055698 |  | 5.061 | 0.178
5 | YES | 800.986 | 0.3465 | 0.0021213 | 0.0155 | 0.0069813 | -0.0085187 |  | 5.309 | 0.137
5 | YES | 598.997 | 0.613 | 0.0014142 | -0.041 | 0.0067222 | 0.047722 |  | 6.7495 | 0.0215
5 | YES | 398.978 | 1.502 | 0.012728 | -0.0089999 | 0.005858 | 0.014858 |  | 6.756 | 0.02
5 | YES | 299.003 | 2.4445 | 0.0035356 | 0.0565 | 0.0049417 | -0.051558 |  | 6.759 | 0.0089998
5 | YES | 249.992 | 3.21 | 0 | 0.068 | 0.0041976 | -0.063803 |  | 6.765 | -0.012
5 | YES | 201.005 | 4.405 | 0.02687 | 0.227 | 0.0030359 | -0.22396 |  | 6.7893 | 0.01325
5 | YES | 173.996 | 4.916 | 0.014142 | 0.384 | 0.0025391 | -0.38146 |  | 6.791 | 0.0019999
5 | YES | 150.974 | 5.061 | 0.018385 | 0.178 | 0.0023981 | -0.1756 |  | 0.294 | -0.008
5 | YES | 126.002 | 5.309 | 0.024042 | 0.137 | 0.002157 | -0.13484 |  | 0.2565 | 0.0145
5 | YES | 98.98 | 6.7495 | 0.0035353 | 0.0215 | 0.00075665 | -0.020743 |  | 0.3775 | 0.0515
5 | YES | 75.9945 | 6.756 | 0.0014141 | 0.02 | 0.00075033 | -0.01925 |  | 0.6815 | 0.0105
5 | YES | 48.9815 | 6.759 | 0.0070709 | 0.0089998 | 0.00074741 | -0.0082524 |  | 1.496 | 0.03
5 | YES | 24.0075 | 6.765 | 0.0028286 | -0.012 | 0.00074158 | 0.012741 |  | 1.9325 | -0.0255
5 | YES | 10.483 | 6.7893 | 0.0012582 | 0.01325 | 0.00071801 | -0.012532 |  | 2.615 | 0.019
5 | YES | 6.2655 | 6.791 | 0.0014141 | 0.0019999 | 0.00071631 | -0.0012836 |  | 3.6975 | 0.5805
6 | NO | 0 | 0 | 0 | -3.102 | 0.0073182 | 3.1093 |  | 4.2125 | 0.2605
6 | NO | 0 | 0 | 0 | -2.758 | 0.0073182 | 2.7653 |  | 5.205 | -0.087
6 | NO | 0 | 0 | 0 | -2.419 | 0.0073182 | 2.4263 |  | 5.6915 | -0.0345
6 | NO | 0 | 0 | 0 | -1.964 | 0.0073182 | 1.9713 |  | 6.067 | -0.436
6 | NO | 0 | 0 | 0 | -1.347 | 0.0073182 | 1.3543 |  | 6.7835 | -0.0505
6 | NO | 0 | 0 | 0 | -1.548 | 0.0073182 | 1.5553 |  | 6.8395 | -0.0075002
6 | NO | 1500.98 | 0.5325 | 0.0035356 | -0.0095 | 0.0068005 | 0.0163 |  | 6.852 | 0.0090003
6 | YES | 1252.01 | 0.294 | 0 | -0.008 | 0.0070324 | 0.015032 |  | 6.8185 | -0.0325
6 | YES | 1001 | 0.2565 | 0.0007071 | 0.0145 | 0.0070688 | -0.0074312 |  | 6.8407 | 0.063667
6 | YES | 800.999 | 0.3775 | 0.0021213 | 0.0515 | 0.0069512 | -0.044549 |  | 0.5115 | -0.0415
6 | YES | 601.003 | 0.6815 | 0.0007071 | 0.0105 | 0.0066556 | -0.0038444 |  | 0.335 | -0.017
6 | YES | 401.01 | 1.496 | 0.0070711 | 0.03 | 0.0058638 | -0.024136 |  | 0.2945 | -0.0095
6 | YES | 298.998 | 1.9325 | 0.012021 | -0.0255 | 0.0054395 | 0.030939 |  | 0.376 | 0.012
6 | YES | 252.008 | 2.615 | 0 | 0.019 | 0.004776 | -0.014224 |  | 0.591 | -0.015
6 | YES | 199 | 3.6975 | 0.024749 | 0.5805 | 0.0037236 | -0.57678 |  | 1.2315 | -0.0325
6 | YES | 176 | 4.2125 | 0.053033 | 0.2605 | 0.003223 | -0.25728 |  | 1.678 | -0.181
6 | YES | 150.959 | 5.205 | 0.043841 | -0.087 | 0.0022581 | 0.089258 |  | 2.1985 | -0.0585
6 | YES | 125.974 | 5.6915 | 0.0091924 | -0.0345 | 0.0017852 | 0.036285 |  | 2.8865 | -0.1335
6 | YES | 100.98 | 6.067 | 0.093338 | -0.436 | 0.0014201 | 0.43742 |  | 3.4045 | -0.0365
6 | YES | 75.992 | 6.7835 | 0.0021212 | -0.0505 | 0.0007236 | 0.051224 |  | 4.0155 | -0.2965
6 | YES | 50.95 | 6.8395 | 0.0021215 | -0.0075002 | 0.00066916 | 0.0081693 |  | 4.9175 | -0.3125
6 | YES | 25.9585 | 6.852 | 0.0014145 | 0.0090003 | 0.000657 | -0.0083433 |  | 5.5875 | -0.1705
6 | YES | 10.4867 | 6.8185 | 0.013128 | -0.0325 | 0.00068957 | 0.033189 |  | 6.6935 | -0.2585
6 | YES | 5.464 | 6.8407 | 0.0050333 | 0.063667 | 0.00066802 | -0.062999 |  | 6.878 | -0.088
7 | NO | 0 | 0 | 0 | -3.247 | 0.0073182 | 3.2543 |  | 6.8615 | -0.0955
7 | NO | 0 | 0 | 0 | -2.898 | 0.0073182 | 2.9053 |  | 6.8717 | -0.07925
7 | NO | 0 | 0 | 0 | -2.888 | 0.0073182 | 2.8953 |  | 0.4365 | 0.0075
7 | NO | 0 | 0 | 0 | -2.534 | 0.0073182 | 2.5413 |  | 0.397 | -0.032
7 | NO | 0 | 0 | 0 | -1.945 | 0.0073182 | 1.9523 |  | 0.4485 | -0.0305
7 | NO | 0 | 0 | 0 | -1.266 | 0.0073182 | 1.2733 |  | 0.555 | -0.035
7 | YES | 1499 | 0.5115 | 0.0021213 | -0.0415 | 0.0068209 | 0.048321 |  | 0.795 | -0.045
7 | YES | 1251 | 0.335 | 0 | -0.017 | 0.0069925 | 0.023992 |  | 1.139 | 0.018
7 | YES | 1000.86 | 0.2945 | 0.0007071 | -0.0095 | 0.0070319 | 0.016532 |  | 1.5185 | 0.0165
7 | YES | 798.99 | 0.376 | 0.0014142 | 0.012 | 0.0069526 | -0.0050474 |  | 2.085 | -0.036
7 | YES | 600.904 | 0.591 | 0.0014142 | -0.015 | 0.0067436 | 0.021744 |  | 2.61 | 0.115
7 | YES | 399.035 | 1.2315 | 0.0035355 | -0.0325 | 0.006121 | 0.038621 |  | 3.18 | 0.235
7 | YES | 298.982 | 1.678 | 0.0098995 | -0.181 | 0.0056869 | 0.18669 |  | 4.9005 | 0.5395
7 | YES | 248.987 | 2.1985 | 0.019092 | -0.0585 | 0.0051809 | 0.063681 |  | 6.456 | 0.495
7 | YES | 198.99 | 2.8865 | 0.037477 | -0.1335 | 0.0045121 | 0.13801 |  | 6.9515 | -0.0825
7 | YES | 175.96 | 3.4045 | 0.0021212 | -0.0365 | 0.0040085 | 0.040508 |  | 6.9275 | -0.1345
7 | YES | 149.007 | 4.0155 | 0.0021215 | -0.2965 | 0.0034145 | 0.29991 |  | 6.9435 | -0.1215
7 | YES | 123.989 | 4.9175 | 0.072832 | -0.3125 | 0.0025376 | 0.31504 |  | 6.9822 | -0.08675
7 | YES | 101 | 5.5875 | 0.044547 | -0.1705 | 0.0018863 | 0.17239 |  | 6.998 | -0.058
7 | YES | 75.967 | 6.6935 | 0.017677 | -0.2585 | 0.00081109 | 0.25931
7 | YES | 51.019 | 6.878 | 0.0014145 | -0.088 | 0.00063173 | 0.088632
7 | YES | 24.002 | 6.8615 | 0.019092 | -0.0955 | 0.00064777 | 0.096148
7 | YES | 9.475 | 6.8717 | 0.0044253 | -0.07925 | 0.0006378 | 0.079888
7 | NO | 4.081 | 6.8715 | 0.0035356 | -0.0645 | 0.00063805 | 0.065138
8 | NO | 0 | 0 | 0 | -3.644 | 0.0073182 | 3.6513
8 | NO | 0 | 0 | 0 | -2.99 | 0.0073182 | 2.9973
8 | NO | 0 | 0 | 0 | -2.578 | 0.0073182 | 2.5853
8 | NO | 0 | 0 | 0 | -2.025 | 0.0073182 | 2.0323
8 | NO | 0 | 0 | 0 | -2.012 | 0.0073182 | 2.0193
8 | NO | 0 | 0 | 0 | -1.381 | 0.0073182 | 1.3883
8 | NO | 1501.02 | 0.631 | 0.0014142 | 0.001 | 0.0067047 | 0.0057047
8 | YES | 1251 | 0.4365 | 0.00070712 | 0.0075 | 0.0068938 | -0.0006062
8 | YES | 1002 | 0.397 | 0 | -0.032 | 0.0069322 | 0.038932
8 | YES | 798.983 | 0.4485 | 0.0007071 | -0.0305 | 0.0068822 | 0.037382
8 | YES | 602.018 | 0.555 | 0 | -0.035 | 0.0067786 | 0.041779
8 | YES | 401.005 | 0.795 | 0.0014142 | -0.045 | 0.0065453 | 0.051545
8 | YES | 299.009 | 1.139 | 0.0056569 | 0.018 | 0.0062109 | -0.011789
8 | YES | 251.027 | 1.5185 | 0.0035355 | 0.0165 | 0.005842 | -0.010658
8 | YES | 200.984 | 2.085 | 0.028284 | -0.036 | 0.0052912 | 0.041291
8 | YES | 174.033 | 2.61 | 0.02687 | 0.115 | 0.0047809 | -0.11022
8 | YES | 151.99 | 3.18 | 0 | 0.235 | 0.0042267 | -0.23077
8 | YES | 123.996 | 4.9005 | 0.1096 | 0.5395 | 0.0025542 | -0.53695
8 | YES | 100.994 | 6.456 | 0.14425 | 0.495 | 0.001042 | -0.49396
8 | YES | 73.9945 | 6.9515 | 0.0035356 | -0.0825 | 0.00056028 | 0.08306
8 | YES | 48.9945 | 6.9275 | 0.0063642 | -0.1345 | 0.00058361 | 0.13508
8 | YES | 23.9955 | 6.9435 | 0.0049497 | -0.1215 | 0.00056805 | 0.12207
8 | YES | 9.47 | 6.9822 | 0.0041932 | -0.08675 | 0.00053038 | 0.08728
8 | YES | 5.3487 | 6.998 | 0.00099993 | -0.058 | 0.00051507 | 0.058515
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
YES
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
NO
NO
NO
NO
NO
NO
NO
NO
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
YES
FIT vs. Pressure
FIT vs. DOX-CTD
FIT vs. File Pair #
MAKE/UPDATE CHART
CREATE CAL CTRL
INCLUDE/EXCLUDE

## Main
# Channels= | 2
| Primary | Secondary |  | Comparison Name
Ref | Pressure | Pressure |  | Pressure - Pressure
Comp 1 | DOX-BOT | DOX-CTD |  | DOX-CTD - DOX-BOT
Average Size: | 2
|  |  | Primary | Secondary | Secondary |  | Primary | Secondary | Secondary
Primary | Secondary | Secondary | Pressure | Pressure | Pressure | (S-P) | DOX-BOT | DOX-CTD | DOX-CTD | (S-P)
File Name | File Name | Count | Value | Avg | Std Dev | Diff | Value | Avg | Std Dev | Diff
2004-05-0001 | 2004-05-0001 | 2 | 223.078 | 222.372 | 0.55296 | -0.70599 | 0 | 0.043 | 0 | 0.043
2004-05-0001 | 2004-05-0001 | 2 | 200.755 | 200.997 | 1.4093 | 0.24249 | 0 | 0.046 | 0 | 0.046
2004-05-0001 | 2004-05-0001 | 2 | 175.734 | 175.992 | 1.4057 | 0.25801 | 0 | 0.049 | 0 | 0.049
2004-05-0001 | 2004-05-0001 | 2 | 175.736 | 175.992 | 1.4057 | 0.25601 | 0 | 0.049 | 0 | 0.049
2004-05-0001 | 2004-05-0001 | 2 | 150.58 | 150.999 | 1.41 | 0.41899 | 0.052 | 0.051 | 0 | -0.001
2004-05-0001 | 2004-05-0001 | 2 | 125.507 | 126.004 | 1.4036 | 0.4965 | 0.052 | 0.0745 | 0.006364 | 0.0225
2004-05-0001 | 2004-05-0001 | 2 | 100.294 | 101.007 | 1.41 | 0.713 | 1.021 | 1.1185 | 0.074246 | 0.0975
2004-05-0001 | 2004-05-0001 | 2 | 75.447 | 75.9925 | 1.4064 | 0.5455 | 4.276 | 4.438 | 0.11597 | 0.162
2004-05-0001 | 2004-05-0001 | 2 | 50.245 | 51.0615 | 1.5068 | 0.8165 | 4.263 | 4.389 | 0.08061 | 0.126
2004-05-0001 | 2004-05-0001 | 2 | 25.169 | 25.9955 | 1.4121 | 0.8265 | 5.181 | 5.2975 | 0.065761 | 0.1165
2004-05-0001 | 2004-05-0001 | 4 | 10.579 | 10.4907 | 1.2958 | -0.088251 | 5.405 | 5.423 | 0.14636 | 0.018
2004-05-0001 | 2004-05-0001 | 1 | 5.593 | 7.262 | 0 | 1.669 | 5.679 | 5.733 | 0 | 0.054
2004-05-0003 | 2004-05-0003 | 0 | 113.796 | 0 | 0 | -113.796 | 4.676 | 0 | 0 | -4.676
2004-05-0003 | 2004-05-0003 | 2 | 101.238 | 101.03 | 1.4411 | -0.208 | 4.23 | 4.145 | 0.0056571 | -0.085
2004-05-0003 | 2004-05-0003 | 2 | 76.233 | 75.9635 | 1.4531 | -0.2695 | 4.711 | 4.5955 | 0.041719 | -0.1155
2004-05-0003 | 2004-05-0003 | 2 | 75.077 | 75.9635 | 1.4531 | 0.8865 | 4.707 | 4.5955 | 0.041719 | -0.1115
2004-05-0003 | 2004-05-0003 | 2 | 51.408 | 51.007 | 1.4255 | -0.401 | 5.907 | 5.568 | 0.074953 | -0.339
2004-05-0003 | 2004-05-0003 | 2 | 26.255 | 25.9815 | 1.4291 | -0.2735 | 6.594 | 6.64 | 0.0028286 | 0.046
2004-05-0003 | 2004-05-0003 | 5 | 10.928 | 10.3832 | 1.107 | -0.5448 | 6.7 | 6.7162 | 0.0054037 | 0.0162
2004-05-0003 | 2004-05-0003 | 3 | 6.162 | 5.088 | 0.85259 | -1.074 | 6.77 | 6.757 | 0.011533 | -0.013
2004-05-0007 | 2004-05-0007 | 2 | 1324.18 | 1324.86 | 1.287 | 0.68494 | 0.468 | 0.4605 | 0.0021213 | -0.0075
2004-05-0007 | 2004-05-0007 | 2 | 1250.16 | 1250.98 | 1.4135 | 0.82849 | 0.38 | 0.364 | 0 | -0.016
2004-05-0007 | 2004-05-0007 | 2 | 1000.92 | 1000.99 | 1.3583 | 0.073486 | 0.227 | 0.232 | 0 | 0.005
2004-05-0007 | 2004-05-0007 | 2 | 1001.49 | 1000.99 | 1.3583 | -0.49951 | 0.242 | 0.232 | 0 | -0.01
2004-05-0007 | 2004-05-0007 | 1 | 802.207 | 802.038 | 0 | -0.16895 | 0.241 | 0.236 | 0 | -0.005
2004-05-0007 | 2004-05-0007 | 2 | 600.939 | 600.988 | 1.4283 | 0.049011 | 0.447 | 0.445 | 0.0014142 | -0.002
2004-05-0007 | 2004-05-0007 | 2 | 402.651 | 402.992 | 1.4262 | 0.34052 | 1.273 | 1.271 | 0 | -0.002
2004-05-0007 | 2004-05-0007 | 2 | 300.968 | 300.996 | 1.4163 | 0.028503 | 1.693 | 1.732 | 0 | 0.039
2004-05-0007 | 2004-05-0007 | 2 | 250.478 | 250.996 | 1.4376 | 0.51849 | 2.029 | 2.09 | 0.0056569 | 0.061
2004-05-0007 | 2004-05-0007 | 2 | 200.576 | 200.962 | 1.3866 | 0.38649 | 2.54 | 2.5065 | 0.15486 | -0.0335
2004-05-0007 | 2004-05-0007 | 2 | 176.465 | 176.021 | 1.4121 | -0.4435 | 3.007 | 3.0135 | 0.020506 | 0.0065002
2004-05-0007 | 2004-05-0007 | 2 | 149.429 | 148.993 | 1.4453 | -0.436 | 3.547 | 3.716 | 0.041012 | 0.169
2004-05-0007 | 2004-05-0007 | 2 | 126.741 | 126.137 | 1.5507 | -0.6035 | 4.367 | 4.407 | 0.057983 | 0.04
2004-05-0007 | 2004-05-0007 | 1 | 102.079 | 102.018 | 0 | -0.061005 | 4.755 | 4.891 | 0 | 0.136
2004-05-0007 | 2004-05-0007 | 2 | 75.349 | 76.0105 | 1.3881 | 0.6615 | 6.6 | 6.6175 | 0.041719 | 0.0175
2004-05-0007 | 2004-05-0007 | 2 | 51.234 | 51.0015 | 1.422 | -0.2325 | 6.613 | 6.6645 | 0.0021212 | 0.0515
2004-05-0007 | 2004-05-0007 | 2 | 26.363 | 25.9945 | 1.3909 | -0.3685 | 6.676 | 6.677 | 0.0014141 | 0.00099993
2004-05-0007 | 2004-05-0007 | 4 | 10.524 | 10.5122 | 1.3321 | -0.01175 | 6.643 | 6.6878 | 0.0092151 | 0.04475
2004-05-0007 | 2004-05-0007 | 2 | 5.755 | 5.9755 | 1.364 | 0.2205 | 6.658 | 6.642 | 0.070711 | -0.016
2004-05-0013 | 2004-05-0013 | 0 | 2001.07 | 0 | 0 | -2001.07 | 1.677 | 0 | 0 | -1.677
2004-05-0013 | 2004-05-0013 | 2 | 1500.85 | 1500.99 | 1.4354 | 0.13501 | 0.755 | 0.695 | 0 | -0.06
2004-05-0013 | 2004-05-0013 | 2 | 1248.64 | 1248.49 | 2.1256 | -0.14893 | 0.416 | 0.408 | 0.0028284 | -0.008
2004-05-0013 | 2004-05-0013 | 2 | 1250.28 | 1251 | 1.4305 | 0.72339 | 0.427 | 0.41 | 0 | -0.017
2004-05-0013 | 2004-05-0013 | 2 | 1001.88 | 1000.98 | 1.4319 | -0.89648 | 0.257 | 0.2465 | 0.0007071 | -0.0105
2004-05-0013 | 2004-05-0013 | 2 | 801.144 | 800.987 | 1.4107 | -0.15747 | 0.245 | 0.2265 | 0.00070711 | -0.0185
2004-05-0013 | 2004-05-0013 | 2 | 601.429 | 601 | 1.4312 | -0.42902 | 0.394 | 0.3995 | 0.0035355 | 0.0055
2004-05-0013 | 2004-05-0013 | 2 | 400.882 | 401.013 | 1.463 | 0.1315 | 1.28 | 1.394 | 0.0028285 | 0.114
2004-05-0013 | 2004-05-0013 | 2 | 301.349 | 300.982 | 1.4262 | -0.36652 | 2.234 | 2.479 | 0.018385 | 0.245
2004-05-0013 | 2004-05-0013 | 2 | 249.591 | 249.005 | 1.4333 | -0.5865 | 2.775 | 3.0075 | 0.016264 | 0.2325
2004-05-0013 | 2004-05-0013 | 2 | 201.248 | 201.003 | 1.4269 | -0.24501 | 3.36 | 2.892 | 0.059397 | -0.468
2004-05-0013 | 2004-05-0013 | 2 | 175.993 | 175.991 | 1.4128 | -0.0019989 | 3.882 | 3.4985 | 0.16334 | -0.3835
2004-05-0013 | 2004-05-0013 | 2 | 150.926 | 150.991 | 1.4248 | 0.065491 | 3.641 | 3.3255 | 0.017678 | -0.3155
2004-05-0013 | 2004-05-0013 | 2 | 126.898 | 125.986 | 1.3435 | -0.912 | 3.837 | 3.6725 | 0.054447 | -0.1645
2004-05-0013 | 2004-05-0013 | 2 | 101.336 | 100.993 | 1.4354 | -0.343 | 4.296 | 4.3345 | 0.033234 | 0.0385
2004-05-0013 | 2004-05-0013 | 2 | 76.236 | 76.034 | 1.4849 | -0.202 | 6.545 | 6.729 | 0.0098995 | 0.184
2004-05-0013 | 2004-05-0013 | 2 | 50.683 | 51.019 | 1.4156 | 0.336 | 6.749 | 6.762 | 0.0042427 | 0.013
2004-05-0013 | 2004-05-0013 | 2 | 24.5 | 24.017 | 1.4199 | -0.483 | 6.715 | 6.732 | 0.0042427 | 0.017
2004-05-0013 | 2004-05-0013 | 3 | 11.763 | 10.9907 | 1.0095 | -0.77233 | 6.719 | 6.736 | 0.0052914 | 0.017
2004-05-0013 | 2004-05-0013 | 2 | 5.323 | 6.0625 | 1.3216 | 0.7395 | 6.725 | 6.74 | 0.0028286 | 0.015
2004-05-0018 | 2004-05-0018 | 0 | 3200.43 | 0 | 0 | -3200.43 | 2.484 | 0 | 0 | -2.484
2004-05-0018 | 2004-05-0018 | 0 | 3000.14 | 0 | 0 | -3000.14 | 2.244 | 0 | 0 | -2.244
2004-05-0018 | 2004-05-0018 | 0 | 2500.09 | 0 | 0 | -2500.09 | 1.931 | 0 | 0 | -1.931
2004-05-0018 | 2004-05-0018 | 0 | 1999.18 | 0 | 0 | -1999.18 | 1.405 | 0 | 0 | -1.405
2004-05-0018 | 2004-05-0018 | 2 | 1499.98 | 1498.99 | 1.4114 | -0.98511 | 0.644 | 0.6005 | 0.0021213 | -0.0435
2004-05-0018 | 2004-05-0018 | 2 | 1501.59 | 1501 | 1.4213 | -0.59204 | 0.643 | 0.6035 | 0.0021213 | -0.0395
2004-05-0018 | 2004-05-0018 | 2 | 1250.68 | 1251 | 1.4269 | 0.32105 | 0.338 | 0.337 | 0.0014142 | -0.00099999
2004-05-0018 | 2004-05-0018 | 2 | 1000.97 | 1000.99 | 1.4128 | 0.013062 | 0.254 | 0.2555 | 0.00070712 | 0.0015
2004-05-0018 | 2004-05-0018 | 2 | 800.816 | 800.986 | 1.4086 | 0.16998 | 0.331 | 0.3465 | 0.0021213 | 0.0155
2004-05-0018 | 2004-05-0018 | 2 | 599.794 | 598.997 | 1.4001 | -0.797 | 0.654 | 0.613 | 0.0014142 | -0.041
2004-05-0018 | 2004-05-0018 | 2 | 399.56 | 398.978 | 1.3881 | -0.58249 | 1.511 | 1.502 | 0.012728 | -0.0089999
2004-05-0018 | 2004-05-0018 | 2 | 299.57 | 299.003 | 1.4199 | -0.56702 | 2.388 | 2.4445 | 0.0035356 | 0.0565
2004-05-0018 | 2004-05-0018 | 1 | 249.991 | 249.992 | 0 | 0.0010071 | 3.142 | 3.21 | 0 | 0.068
2004-05-0018 | 2004-05-0018 | 2 | 200.278 | 201.005 | 1.4269 | 0.72701 | 4.178 | 4.405 | 0.02687 | 0.227
2004-05-0018 | 2004-05-0018 | 2 | 174.298 | 173.996 | 1.4277 | -0.30151 | 4.532 | 4.916 | 0.014142 | 0.384
2004-05-0018 | 2004-05-0018 | 2 | 151.632 | 150.974 | 1.4694 | -0.658 | 4.883 | 5.061 | 0.018385 | 0.178
2004-05-0018 | 2004-05-0018 | 2 | 125.804 | 126.002 | 1.4383 | 0.198 | 5.172 | 5.309 | 0.024042 | 0.137
2004-05-0018 | 2004-05-0018 | 2 | 99.136 | 98.98 | 1.4199 | -0.15601 | 6.728 | 6.7495 | 0.0035353 | 0.0215
2004-05-0018 | 2004-05-0018 | 2 | 76.393 | 75.9945 | 1.4149 | -0.39849 | 6.736 | 6.756 | 0.0014141 | 0.02
2004-05-0018 | 2004-05-0018 | 2 | 49.932 | 48.9815 | 1.4192 | -0.9505 | 6.75 | 6.759 | 0.0070709 | 0.0089998
2004-05-0018 | 2004-05-0018 | 2 | 24.829 | 24.0075 | 1.4432 | -0.8215 | 6.777 | 6.765 | 0.0028286 | -0.012
2004-05-0018 | 2004-05-0018 | 4 | 10.145 | 10.483 | 1.2957 | 0.338 | 6.776 | 6.7893 | 0.0012582 | 0.01325
2004-05-0018 | 2004-05-0018 | 2 | 5.585 | 6.2655 | 1.1519 | 0.6805 | 6.789 | 6.791 | 0.0014141 | 0.0019999
2004-05-0025 | 2004-05-0025 | 0 | 3674.55 | 0 | 0 | -3674.55 | 3.102 | 0 | 0 | -3.102
2004-05-0025 | 2004-05-0025 | 0 | 3501.17 | 0 | 0 | -3501.17 | 2.758 | 0 | 0 | -2.758
2004-05-0025 | 2004-05-0025 | 0 | 2999.45 | 0 | 0 | -2999.45 | 2.419 | 0 | 0 | -2.419
2004-05-0025 | 2004-05-0025 | 0 | 2498.94 | 0 | 0 | -2498.94 | 1.964 | 0 | 0 | -1.964
2004-05-0025 | 2004-05-0025 | 0 | 2000.87 | 0 | 0 | -2000.87 | 1.347 | 0 | 0 | -1.347
2004-05-0025 | 2004-05-0025 | 0 | 2001.03 | 0 | 0 | -2001.03 | 1.548 | 0 | 0 | -1.548
2004-05-0025 | 2004-05-0025 | 2 | 1501.6 | 1500.98 | 1.417 | -0.61194 | 0.542 | 0.5325 | 0.0035356 | -0.0095
2004-05-0025 | 2004-05-0025 | 1 | 1253.11 | 1252.01 | 0 | -1.094 | 0.302 | 0.294 | 0 | -0.008
2004-05-0025 | 2004-05-0025 | 2 | 1001.28 | 1001 | 1.4389 | -0.2785 | 0.242 | 0.2565 | 0.0007071 | 0.0145
2004-05-0025 | 2004-05-0025 | 2 | 801.384 | 800.999 | 1.4389 | -0.38544 | 0.326 | 0.3775 | 0.0021213 | 0.0515
2004-05-0025 | 2004-05-0025 | 2 | 600.599 | 601.003 | 1.434 | 0.40405 | 0.671 | 0.6815 | 0.0007071 | 0.0105
2004-05-0025 | 2004-05-0025 | 2 | 400.794 | 401.01 | 1.4213 | 0.216 | 1.466 | 1.496 | 0.0070711 | 0.03
2004-05-0025 | 2004-05-0025 | 2 | 299.167 | 298.998 | 1.4135 | -0.16852 | 1.958 | 1.9325 | 0.012021 | -0.0255
2004-05-0025 | 2004-05-0025 | 1 | 252.285 | 252.008 | 0 | -0.27701 | 2.596 | 2.615 | 0 | 0.019
2004-05-0025 | 2004-05-0025 | 2 | 199.301 | 199 | 1.4347 | -0.30148 | 3.117 | 3.6975 | 0.024749 | 0.5805
2004-05-0025 | 2004-05-0025 | 2 | 176.2 | 176 | 1.4192 | -0.19951 | 3.952 | 4.2125 | 0.053033 | 0.2605
2004-05-0025 | 2004-05-0025 | 2 | 150.82 | 150.959 | 1.4821 | 0.13899 | 5.292 | 5.205 | 0.043841 | -0.087
2004-05-0025 | 2004-05-0025 | 2 | 126.37 | 125.974 | 1.3874 | -0.396 | 5.726 | 5.6915 | 0.0091924 | -0.0345
2004-05-0025 | 2004-05-0025 | 2 | 100.426 | 100.98 | 1.4453 | 0.55399 | 6.503 | 6.067 | 0.093338 | -0.436
2004-05-0025 | 2004-05-0025 | 2 | 75.972 | 75.992 | 1.3647 | 0.020004 | 6.834 | 6.7835 | 0.0021212 | -0.0505
2004-05-0025 | 2004-05-0025 | 2 | 51.668 | 50.95 | 1.451 | -0.718 | 6.847 | 6.8395 | 0.0021215 | -0.0075002
2004-05-0025 | 2004-05-0025 | 2 | 26.469 | 25.9585 | 1.4361 | -0.5105 | 6.843 | 6.852 | 0.0014145 | 0.0090003
2004-05-0025 | 2004-05-0025 | 4 | 10.665 | 10.4867 | 1.3224 | -0.17825 | 6.851 | 6.8185 | 0.013128 | -0.0325
2004-05-0025 | 2004-05-0025 | 3 | 5.912 | 5.464 | 1.3846 | -0.448 | 6.777 | 6.8407 | 0.0050333 | 0.063667
2004-05-0032 | 2004-05-0032 | 0 | 3999.48 | 0 | 0 | -3999.48 | 3.247 | 0 | 0 | -3.247
2004-05-0032 | 2004-05-0032 | 0 | 3502.3 | 0 | 0 | -3502.3 | 2.898 | 0 | 0 | -2.898
2004-05-0032 | 2004-05-0032 | 0 | 3501.01 | 0 | 0 | -3501.01 | 2.888 | 0 | 0 | -2.888
2004-05-0032 | 2004-05-0032 | 0 | 3000.97 | 0 | 0 | -3000.97 | 2.534 | 0 | 0 | -2.534
2004-05-0032 | 2004-05-0032 | 0 | 2500.1 | 0 | 0 | -2500.1 | 1.945 | 0 | 0 | -1.945
2004-05-0032 | 2004-05-0032 | 0 | 2000.3 | 0 | 0 | -2000.3 | 1.266 | 0 | 0 | -1.266
2004-05-0032 | 2004-05-0032 | 2 | 1498.56 | 1499 | 1.4022 | 0.4425 | 0.553 | 0.5115 | 0.0021213 | -0.0415
2004-05-0032 | 2004-05-0032 | 2 | 1251.76 | 1251 | 1.4071 | -0.76001 | 0.352 | 0.335 | 0 | -0.017
2004-05-0032 | 2004-05-0032 | 2 | 1001.33 | 1000.86 | 1.6193 | -0.47504 | 0.304 | 0.2945 | 0.0007071 | -0.0095
2004-05-0032 | 2004-05-0032 | 2 | 798.273 | 798.99 | 1.4093 | 0.71649 | 0.364 | 0.376 | 0.0014142 | 0.012
2004-05-0032 | 2004-05-0032 | 2 | 600.711 | 600.904 | 1.5097 | 0.19348 | 0.606 | 0.591 | 0.0014142 | -0.015
2004-05-0032 | 2004-05-0032 | 2 | 399.141 | 399.035 | 1.4397 | -0.10599 | 1.264 | 1.2315 | 0.0035355 | -0.0325
2004-05-0032 | 2004-05-0032 | 2 | 299.095 | 298.982 | 1.4432 | -0.11252 | 1.859 | 1.678 | 0.0098995 | -0.181
2004-05-0032 | 2004-05-0032 | 2 | 249.239 | 248.987 | 1.4368 | -0.252 | 2.257 | 2.1985 | 0.019092 | -0.0585
2004-05-0032 | 2004-05-0032 | 2 | 199.341 | 198.99 | 1.4093 | -0.3515 | 3.02 | 2.8865 | 0.037477 | -0.1335
2004-05-0032 | 2004-05-0032 | 2 | 175.572 | 175.96 | 1.3251 | 0.38799 | 3.441 | 3.4045 | 0.0021212 | -0.0365
2004-05-0032 | 2004-05-0032 | 2 | 149.177 | 149.007 | 1.4184 | -0.17001 | 4.312 | 4.0155 | 0.0021215 | -0.2965
2004-05-0032 | 2004-05-0032 | 2 | 124.453 | 123.989 | 1.3986 | -0.46401 | 5.23 | 4.9175 | 0.072832 | -0.3125
2004-05-0032 | 2004-05-0032 | 2 | 101.936 | 101 | 1.3704 | -0.936 | 5.758 | 5.5875 | 0.044547 | -0.1705
2004-05-0032 | 2004-05-0032 | 2 | 75.818 | 75.967 | 1.4015 | 0.14899 | 6.952 | 6.6935 | 0.017677 | -0.2585
2004-05-0032 | 2004-05-0032 | 2 | 50.066 | 51.019 | 1.4156 | 0.953 | 6.966 | 6.878 | 0.0014145 | -0.088
2004-05-0032 | 2004-05-0032 | 2 | 24.543 | 24.002 | 1.393 | -0.541 | 6.957 | 6.8615 | 0.019092 | -0.0955
2004-05-0032 | 2004-05-0032 | 4 | 9.897 | 9.475 | 1.3203 | -0.422 | 6.951 | 6.8717 | 0.0044253 | -0.07925
2004-05-0032 | 2004-05-0032 | 2 | 4.116 | 4.081 | 1.2728 | -0.035 | 6.936 | 6.8715 | 0.0035356 | -0.0645
2004-05-0047 | 2004-05-0047 | 0 | 4000.37 | 0 | 0 | -4000.37 | 3.644 | 0 | 0 | -3.644
2004-05-0047 | 2004-05-0047 | 0 | 3501.53 | 0 | 0 | -3501.53 | 2.99 | 0 | 0 | -2.99
2004-05-0047 | 2004-05-0047 | 0 | 3001.69 | 0 | 0 | -3001.69 | 2.578 | 0 | 0 | -2.578
2004-05-0047 | 2004-05-0047 | 0 | 2500.83 | 0 | 0 | -2500.83 | 2.025 | 0 | 0 | -2.025
2004-05-0047 | 2004-05-0047 | 0 | 2501.18 | 0 | 0 | -2501.18 | 2.012 | 0 | 0 | -2.012
2004-05-0047 | 2004-05-0047 | 0 | 2000.99 | 0 | 0 | -2000.99 | 1.381 | 0 | 0 | -1.381
2004-05-0047 | 2004-05-0047 | 2 | 1501.34 | 1501.02 | 1.4545 | -0.32044 | 0.63 | 0.631 | 0.0014142 | 0.001
2004-05-0047 | 2004-05-0047 | 2 | 1250.28 | 1251 | 1.4262 | 0.72742 | 0.429 | 0.4365 | 0.00070712 | 0.0075
2004-05-0047 | 2004-05-0047 | 1 | 1002.85 | 1002 | 0 | -0.84796 | 0.429 | 0.397 | 0 | -0.032
2004-05-0047 | 2004-05-0047 | 2 | 799.416 | 798.983 | 1.4029 | -0.43298 | 0.479 | 0.4485 | 0.0007071 | -0.0305
2004-05-0047 | 2004-05-0047 | 1 | 602.617 | 602.018 | 0 | -0.599 | 0.59 | 0.555 | 0 | -0.035
2004-05-0047 | 2004-05-0047 | 2 | 400.188 | 401.005 | 1.4163 | 0.81653 | 0.84 | 0.795 | 0.0014142 | -0.045
2004-05-0047 | 2004-05-0047 | 2 | 298.565 | 299.009 | 1.4029 | 0.444 | 1.121 | 1.139 | 0.0056569 | 0.018
2004-05-0047 | 2004-05-0047 | 2 | 250.406 | 251.027 | 1.4397 | 0.621 | 1.502 | 1.5185 | 0.0035355 | 0.0165
2004-05-0047 | 2004-05-0047 | 2 | 200.921 | 200.984 | 1.4206 | 0.063492 | 2.121 | 2.085 | 0.028284 | -0.036
2004-05-0047 | 2004-05-0047 | 2 | 174.735 | 174.033 | 1.4305 | -0.7025 | 2.495 | 2.61 | 0.02687 | 0.115
2004-05-0047 | 2004-05-0047 | 1 | 152.107 | 151.99 | 0 | -0.11699 | 2.945 | 3.18 | 0 | 0.235
2004-05-0047 | 2004-05-0047 | 2 | 123.544 | 123.996 | 1.4418 | 0.4525 | 4.361 | 4.9005 | 0.1096 | 0.5395
2004-05-0047 | 2004-05-0047 | 2 | 100.287 | 100.994 | 1.3895 | 0.7075 | 5.961 | 6.456 | 0.14425 | 0.495
2004-05-0047 | 2004-05-0047 | 2 | 74.409 | 73.9945 | 1.3697 | -0.4145 | 7.034 | 6.9515 | 0.0035356 | -0.0825
2004-05-0047 | 2004-05-0047 | 2 | 49.38 | 48.9945 | 1.405 | -0.3855 | 7.062 | 6.9275 | 0.0063642 | -0.1345
2004-05-0047 | 2004-05-0047 | 2 | 24.787 | 23.9955 | 1.4079 | -0.7915 | 7.065 | 6.9435 | 0.0049497 | -0.1215
2004-05-0047 | 2004-05-0047 | 4 | 9.144 | 9.47 | 1.3303 | 0.326 | 7.069 | 6.9822 | 0.0041932 | -0.08675
2004-05-0047 | 2004-05-0047 | 3 | 5.351 | 5.3487 | 1.5096 | -0.0023332 | 7.056 | 6.998 | 0.00099993 | -0.058
Create Fit Sheet