## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 6-Jun-2013 | Added Iron profile files with cast numbers 8xxx from Keith Johnson’s spreadsheet file which can be found in the cruise .DOC directory. |
| 20-Feb-2012 | Added one productivity cast, (9001), containing Primary Productivity, Chlorophyll, POC and PON, from Frank Whitney’s Productivity spreadsheet. |
| 27-Sep-2010 | Errors corrected in Loop File – Silicate and Phosphate were reversed and phosphate values edited to match original source file to get 2 decimal places. G. Gatien |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |
| 5-Apr-2010 | Added Lisa Miller’s Dissolved Inorganic Carbon and Alkalinity data to the rosette files. J.L. |
| 19-Jan-2006 | Added loop data to archive. |

## PROCESSING NOTES

Cruise: 2004-05

Agency: OSAP

Location: Gulf of Alaska

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: February 18, 2004 – February 28, 2004

Processed by: Germaine Gatien

Date of Processing: March 22, 2004 – April 7, 2004

Number of original CTD casts: 31

Number of casts processed: 31

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlab transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter OA-916D (#1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0424) and the logging computer was #FS02. The salinometers were Portasal models 8410 (#59724 and 58879). A mid-ship winch was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order.

Notes from the chief scientist on problems during the cruise were very helpful.

The primary temperature and conductivity sensors both appear to be working poorly.

The transmissivity and fluorometer worked badly below 150db for casts #1- #23.

This was the first cruise using this dissolved oxygen sensor since it was repaired. The quality of the data is much higher than for previous uses.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l from 0db–150db.

- ±0.1ml/l below 150db – 500db

- ±0.05ml/l from 500 – 1500db

- unreliable below 1500db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Problems were noted with the surface PAR.

The station name is wrong for cast #26.

Salinity, chlorophyll, dissolved oxygen and nutrient bottle calibration data were obtained. Flags and comments had been added as needed.

The cruise summary sheet was completed.

The history of the conductivity sensors was obtained.

The calibration constants were checked for all instruments. The dates of the pressure and transmissivity calibrations were wrong. There was an offset of -0.2db in the pressure configuration. During 2004-04, which immediately followed this cruise, a pressure test was done which showed that there should be an offset of +0.2db. It is presumed that the entry for 2004-05 had a sign error. To be sure, a check was done to ensure that +0.2db is a good choice for this cruise as well. For cast #47 using an offset of -0.2db leads to data points with pressures less than zero that have “in-water” values including points during the upcast with the pumps running and pressures less than -0.3db.

A new configuration file (2004-05-ctd.con) was prepared with the dates of calibrations corrected and the pressure offset changed to +0.2db.

There were problems with the transmissometer for casts #1-23 due to a missing piece in the mounting block.

The fluorometer data was poor (large spikes) for the first part of the cruise until the cable was changed.

It was noted at sea that the salinity differences were strongly pressure-dependent. The secondary sensors were found to be closer to the bottles.

##### Conversion of Raw Data

All data were converted using 2004-05-ctd.con. The Bottom Contact channel was not converted.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. The dissolved oxygen data is the first seen since repairs were done on this instrument and it looks notably better than has been seen before. The altimeter data is bad for many casts; it is noted in the log that it was not useful because it was “seeing the weight.” While it was clearly not useful for at-sea purposes there does seem to be useful data for a few casts to see how close to the bottom the first bottles were fired. Casts 1,2,3,7 and 50 have useful data. Since other casts may have false low values due to the weight, the channel will be stripped from the files before conversion to IOS HEADERS, to avoid putting bad information into the headers.

Rosette files were converted using a start time of -2s and duration of 5s. A few errors mentioned in the Daily Log book and/or notes from M. Robert, were fixed in both CNV and ROS files:

- The station name for cast #44 was changed to P26.

- Cast #48 was changed to #50.

Because of the problems with the altimeter and because STRIP will not work on ROS files, ROS files were recreated without the altimeter channel for casts #5 and #10-47. This is most easily done early so the altimeter value does not go into the IOS HEADERS.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit was one of the old ones that do not advance the secondary conductivity channels so ALIGNCTD was used to advance it by +0.073s. Tests will be done later to fine-tune the alignment.

The dissolved oxygen channel was aligned by +5s based on tests on a few casts. Again, this can be fine-tuned later if found necessary.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.01,7) , (0.01,9),  (0.02,9), (0.03,9), (0.02,7) and (0.03,7. The setting of (0.02,7) improved the primary data. The secondary sensors have a lot of fine-scale noise in their signals and it was hard to see much effect from running CELLTM, but no setting seemed to make the data better.

All casts were put through CELLTM using (0.02,7) for the primary conductivity only.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

STRIP was run to remove the Altimeter channel from the CNV files in the DERIVE folder for casts #5 and #10-#47.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

The differences between T and C sensors were generally very noisy. The descent rate was frequently very noisy, but was usually kept fairly high. The secondary salinity is full of fine-scale noise with average excursions of ±0.001psu at 2000db, while the noise level of the primary is about ±0.0002psu. There salinity differences show significant time and pressure dependence.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 17 | 1800 | -0.0008 | -0.00035 | -0.004 | High, very.noisy |
| 29 | 1800 | -0.0008 | -0.0004 | -0.0045 | High, x.noisy |
| 47 | 1800 | -0.0005 | -0.0006 | -0.0065 | High, noisy |
| 47 | 3700 | -0.0012 | -0.0006 | -0.005 | High, noisy |

There appear to be two problems – conductivity differences are increasing with time, and temperature differences are depth-dependent. This is likely to be due to faulty primary sensors since the initial observations at sea suggest that the secondary salinity is close to the bottle salinity.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. The 5 casts for which altimetry was converted were checked. The altimeter approached minimum values in a reasonable way, so the data looks useful.

The rosette files were converted to IOS files and renamed BOT.

All BOT files were plotted and significant outliers were found only in cast #50. CTDEDIT was used to clean the secondary temperature and salinity for bottle #19, cast #50. The output file was renamed BOT.

Cast #45 had very noisy data in bottle #1, but simple editing will not change this. This bottle is unlikely to be useful for calibration purposes.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The header check and header summary were run. The station name was missing from cast #50; this was added to the CLN file.

The cruise track was plotted and no problems noted.

The average surface pressure is 2.6db.

All CLN files were examined on-screen. There was no evidence of serious problems in T and S; both salinity channels have occasional spikes. The transmissivity and fluorescence have odd shifts in value during casts 5 to 23. These problems appear below 150db, so are of no significance for the fluorescence except in establishing the dark value. The transmissivity problems are worst in the upcast data. The top 150db of the upcasts looks like the downcasts, with no such shifts; however, there are some large spikes in both the FL and TR in the top 150db.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet was edited to remove the loop samples and was named 2004-05 salinity.csv. That file was converted to individual SAL files. Comments for 3 flagged values were added to the headers of the SAL files for casts #13 and 18.

The dissolved oxygen files (*.add) were created by the analyst and had a quality channel and comments for flagged values. There were errors in the sample numbers for cast #47. These were fixed in the ADD file.

The extracted chlorophyll and Phaeo-Pigment analysis spreadsheet, 2004-05arcchl.xls, was edited to change headers to standard names, to remove loop samples, to add cast numbers and to add flag channels for CHL and PHAEO and was saved as 2004-05-chl-nonloop.csv. The file was then converted to individual CHL files. There were no flagged values.

The nutrient data in file 2004-05nuts.xls was edited to fix channel names and units and saved as 2004-05-nuts.csv. There were flags and comments in the file. This file was converted to NUT files. It was then put through SORT to arrange the data by increasing sample number. (Output: NUT1)

Comments from the nutrient analysis spreadsheet were transferred to the headers of casts #1 and 13.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #50 had 24 bottles but there are no sample numbers in the log book and no data in the spreadsheets. These are presumably being done at UBC.

The SAL, CHL1, DOX and NUT1 files were merged with SAMAVG in four steps. (Output: MRG1, MRG2, MRG3, MRG)

##### 11. COMPARE

Salinity

COMPARE was run and differences above 500db plus a few outliers were excluded. The primary salinity is both pressure-dependent and time-dependent, and high by an average of 0.0052psu. The secondary is lower than the bottles by about 0.0008psu with no obvious dependence on pressure or temperature. No recalibration is necessary for the secondary salinity. (See 2004-05-sal-comp1.xls)

Fluorescence vs Chlorophyll-a

Most of the chlorophyll-a values were low, less than 0.5ug/l. There was a lot of variability in the fluorescence for those points, with FL/CHL ranging from 1 to 6. For the few chlorophyll values greater than 1 (maximum about 4) the ratio FL/CHL ranges from 0.6 to 1.2. Plots of differences between the titrated bottle samples and the CTD fluorescence show little pressure-dependence. (See 2004-05-chl-comp1.xls and chl-vs-fl.xls.)

Dissolved Oxygen

COMPARE was run using titrated dissolved oxygen data from 8 casts and plots were prepared of differences between CTD and titrated values versus pressure, DOX and cast number. The best fit was versus DOX value. Two outliers were identified from that plot and removed from the comparison, as were values from 1500db downwards since the sensor is not expected to give good values at those pressures. Using those values the fit is:

DOX (Titrated) = 1.0526 * DOX (CTD) + 0.0291

The fit against cast number suggests time dependence, but this is probably not real. The first cast had several anoxic samples and for the rest of the cruise the changes are more likely geographic in origin since the distance offshore steadily increased. (See 2004-05-dox-comp1.xls)

The above fit will later be used to recalibrate the CTD DOX channel.

Outliers from all these comparisons were examined; most were not extreme enough to warrant flagging or it appeared that the CTD value might be wrong rather than the bottle value.

##### 12. SHIFT

Conductivity

Tests were run on the secondary conductivity channel only (since the primary salinity is unlikely to be used) to determine what shift reduces instabilities best without oversmoothing; results were examined in T-S plots after running DELETE. The best results came from an advancement of -2.25. That would mean a net advancement of about -0.021s, since the channel was already advanced by +0.073s earlier. This is an unusually low value. This may imply that the secondary channel was advanced by the deck unit. In the past deck units of this vintage did not do this, but something may have changed. If the deck unit advanced the conductivity by 0.073 then the net alignment is +0.052 which seems more reasonable.

Another test of the shift was to look at the noise level in the resulting salinity and the same choice greatly improved the noise level, though the basic fine-scale noise continues to be much higher than in the primary salinity.

All casts were put through SHIFT using -2.25 records for the secondary conductivity. (Output *.SHFC2)

Dissolved Oxygen

Tests were run using advancements of +80 records to +160 records and the best choice varied from cast to cast and from one depth to another. Overall the choice of +120 records (~5s) was found best. That shift was applied to all casts.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was about 1s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

##### DELETE

During the downcast section of cast #1 a winch problem led to multiple up and down sections between 40 and 68db. DELETE will select the first section from the top to 68db and the final downcast section 68db to the bottom. The patch leads to a notable shift in values at 68db, A test was done to see if editing the file for cast #1 before DELETE could improve the match, but the results were no better.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: The only warning pertained to the upcast section of cast #30, so is of no significance to the downcast file.

##### DETAILED EDITING

The secondary sensors were chosen for further processing since the primary conductivity appears to have drifted significantly. The secondary salinity has fine-scale noise but the calibration is good.

Page plots were produced using (T1,S1). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing, and most required heavy editing. The average CTD descent rate was kept high most of the time which reduced the incidence of shed wakes, but conditions were very rough so there was considerable corruption of the data by such wakes.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –

The primary sensors have been recalibrated since their last use. The secondary sensors were last used for 2003-39 and 2003-41 (October and December 2003) when the salinity was within 0.001psu of bottles. On an earlier cruise the salinity was low by 0.003psu.

Historic ranges – All data fell within the historic ranges.

##### Initial Recalibration

File 2004-05-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun. The fit of differences was much improved versus pressure and DOX. The fit versus file pair number still shows time-dependence, though it is not large enough to justify a time-dependent correction and there could still be some geographic factor that complicates this issue. (See 2004-05-dox-comp2.xls.)

##### Final DOX comparison

Another run of COMPARE was made using metre-averaged, thinned downcast files. The differences versus pressure show an offset of about 0.03ml/l with the CTD reading a little high and no significant pressure-dependence. When the differences are plotted against DOX values the results are similar. When plotted against cast number, there appears to be some time dependence. But when the first cast is removed from the comparison it disappears. The first cast was shallow and had lower DOX values.

The offset does not seem likely to be due to an imperfect alignment since a small error would cause little change at depth. It is likely due to the response time being a little slow, so the sensor gives slightly higher values than expected on the downcast. Only the downcast files are affected by this, the bottle files do not require recalibration because the sensor had time to equilibrate.

The COR1 files and the thinned downcast files were recalibrated by subtracting 0.03ml/l using file 2004-05-recal2.ccf. (Output: COR2)

COMPARE was rerun using the recalibrated thinned files and the average difference was +0.004ml/l. (See 2004-05-DOX-comp4.xls)

##### Special Fluorometer Processing

The COR2 files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. Since there was chlorophyll sampling for this cruise, the full bottle files (SAMCOR1) were also put through REMOVE and HEADEDIT (output: BOF). The BOF and FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective, though the large spikes noted earlier are not entirely removed. (Output: BOX)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the BOX files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Replacement of bad data with pad values

Three runs of CLEAN were used to replace the dissolved oxygen data for pressures greater than 1500db, and the fluorescence and transmissivity data for pressures greater than 150db for casts #1-23 only. (Output: CLN4)

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles. Separate profile plots were prepared with DOX, FL and Transmissivity versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

For casts 5 through 23 the fluorescence and transmissivity data

were bad below 150db with sudden shifts in values. Those channels

have been blanked for pressures greater than 140db.

The dissolved oxygen data was removed for pressures greater than

1500db because the sensor is unreliable at depth. The DO data in

the CTD files should be considered:

•	±0.2ml/l from 0db–150db.

•	±0.1ml/l below 150db – 500db

•	±0.05ml/l from 500 – 1500db

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods (2004-05-sal_oxy_chl_fresh_nut_analysis.txt). Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1-23. Very noisy transmissivity signal

1. Winch problems during upcast.

7.   Altimeter problem – sensing weight attached to rosette

10. Transmissometer problem?

12. Reconnected transmissometer cable

22. Wrong event number in original file – fixed in converted files.

23. Transmissivity cable changed.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-05

Dates: Start: February 18, 2001 End: February 28, 2004

Location: Gulf of Alaska

Vessel: John P. Tully

Party Chief: Robert M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-05

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2710 | 20/12/03 | Factory “ |  |  |
| Conductivity | 2278 | 19/12/03 | “ |  |  |
| Secondary Temp. | 2095 | 05/09/03 | “ |  |  |
| Secondary Cond. | 1764 | 05/09/03 | “ |  |  |
| Transmissometer | 333DR | 06/01/04 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13/01/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |