## 2004-06-sal-comp
Salinity:Bottle | Sample_Number | Number of points | Flag:Salinity:Bottle | Bottle_Number | Cast | Pressure | Station | Comments | Pre-S0 | Pre-S1 | Post-S0 | Post-S1 | preS0-bot | preS1-bot | postS0-bot | postS1-bot | local grad
30.9103 | 1 | 3 |  |  | 1 | 336 |  |  | 30.9135 | 30.9122 | 30.9305 | 30.9335 | 0.0032 | 0.0019 | 0.0202 | 0.0232 | 0.005
31.1433 | ????? |  |  |  | 8 | 499 |  |  | 31.1421 | 31.1405 | 31.159 | 31.1626 | -0.0012 | -0.0028 | 0.0157 | 0.0193 | 0.0004
31.1152 | 14 | 2 |  |  | 14 | 374 |  | ctd 14 | 31.1232 | 31.1219 | 31.141 | 31.1435 | 0.008 | 0.0067 | 0.0258 | 0.0283 | 0.004
30.8848 | 15 | 3 |  |  | 15 | 327 |  |  | 30.889 | 30.8875 | 30.9058 | 30.9093 | 0.0042 | 0.0027 | 0.021 | 0.0245 | 0.005
31.156 | 22 | 2 |  |  | 22 | 400 |  | lot of motion | 31.1552 | 31.1549 | 31.1723 | 31.1765 | -0.0008 | -0.0011 | 0.0163 | 0.0205 | 0.003
[10 empty] | average all |  |  | 0.00268 | 0.00148 | 0.0198 | 0.02316
[10 empty] | low gradient |  |  | -0.0012 | -0.0028 | 0.0157 | 0.0193
[9 empty] | post-preS0 | post-pre S1
[9 empty] | 0.017 | 0.0213
[9 empty] | 0.0178 | 0.0216
[9 empty] | 0.0168 | 0.0218
[9 empty] | 0.0171 | 0.0216
[9 empty] | 0.017175 | 0.021575
average
Drift expected between the two calibrations based on conductivity cal change, not temp, at C=3.4
[6 empty] | S0 | drift per month |  | -0.0006 |  | -0.0186
[6 empty] | S1 | at Cond.~3.4 |  | -0.0007 |  | -0.0217
The temperature cal. change was larger for the primary sensors than for the secondary
[6 empty] | if we use the post-cruise calibration and there is linear drift expect salinity too high by | [7 empty] | 0.0051
0.00595
[6 empty] | if we use the pre-cruise calibration and there is linear drift expect sal low by | [7 empty] | 0.0135
0.01575
[9 empty] | 31.1362 | 31.1351
31.1433 | 9 | 3 |  |  | 9 | 499 |  |  |  |  | 31.1533 | 31.1568 | -0.0071 | -0.0082 | 0.01 | 0.0135 | 5e-05