## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-06

Agency: OSAP

Location: Strait of Georgia, Johnstone Strait

Project: SoG Hake - QCS Skate Survey

Party Chief: Cooke K.

Platform: W. E. RICKER

Date: March 2, 2004 – March 30, 2004

Processed by: Germaine Gatien

Date of Processing: 11 February 2005 – 17 February 2005

Number of original CTD casts: 25

Number of casts processed: 25

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0506) was mounted with Transmissometer #197 and Seapoint Fluorometer (#2228). The deck unit was a SeaBird model 11 (S/N presumed to be #0471.) The salinity was analyzed using Portasal Model #8410 (S/N 59724).

# SUMMARY OF QUALITY AND CONCERNS

The log book was missing at the time of processing, but photocopies of the log sheets were obtained. No list of equipment was available.

The descent rate was generally smooth and high. The overall quality of the data is high, but the salinity calibration is in some doubt so that absolute values should be considered ±0.005psu.

There were problems with the secondary conductivity and the fluorescence from the surface to about 45db of the downcast for 3 casts in the area around Sechelt and Lasqueti Island. The downcast transmissivity also looks odd near the surface for those casts. It seems likely that there was some clogging of the secondary pump system that cleared when the pressure was high enough. The transmissometer lens may have been covered by whatever clogged the pump and gradually cleared.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was unavailable, but copies of the sheets were obtained from Steve Romaine. The cruise plan was available as was a text file with times and positions.

Bottle salinity data were obtained.

The configuration files were obtained and the calibration constants were checked.

An error in the calibration constants and date for the pressure sensor were corrected. The resulting file was named 2004-06-CTD.con.

There were post-cruise calibrations from November 2004. These were entered into a file named 2004-06-CTD-post.con.

The sensor history was found. All temperature and conductivity sensors were last used for cruise 2004-03 and were used for 2004-17 in June 2004.

No rosette casts were expected, but a check was made by doing a conversion of rosette files. This produced one rosette file, but the data was very strange, with negative pressures and large differences between sensors. When the CTD cast file was examined similar data was found at the beginning of that file. Later in the file the data looks reasonable. There were also scrambled header entries which is often associated with pressure spikes, so this is assumed to be what happened. The upcast data looks ok, but this should be checked later to ensure the pressure is reliable. The rosette file was deleted.

##### Salinity Bottle Comparison

The pre-cruise calibrations are very old and it is known that there was a lot of drift over the years. However, the post-cruise calibrations are from 8 months later and the CTD was used on at least 4 cruises during that time. There is no assurance the drift was linear, so it was decided to do the bottle comparison early in order to pick the best calibrations to use.

Recently problems have been noted when samples are from the bottom. Cruises in the Gulf of Georgia with bottles from a variety of depths show the bottom samples out of line from others. It is not known if the problem is with the bottle samples, the CTD or both. It is known that the presence of mud will affect conductivity. The CTD tends to look high compared to the bottles in these circumstances. These samples mostly seem to be near the bottom, and the transmissivity suggests some are in a bottom boundary layer, but the values are not so low as to suggest significant problems from mud.

It has also been noted from other Ricker cruises that near-surface samples in well-mixed waters tend to give calibration results that are more reliable than those from the bottom. This could be a question of how well the bottles have flushed. For this cruise in particular, conditions were very tranquil. Without the bobbing up and down that usually occurs, the Niskin may not have flushed as well as we expect which would make the CTD look high relative to bottles.

Cast #9 is given as a calibration sample in the salinity file, but in the log the sample is given as from cast #8. The maximum depth sampled is about the same for the two casts. The time spent at the bottom would suggest that it was cast #8. Cast #8 is said to be in water much deeper than the sample depth but there is a question mark beside the bottom depth, so again, this is uncertain.

Casts for which there was salinity bottle sampling were converted using both sets of calibrations, namely casts 1, 8, 9, 14, 15 and 22. The resulting salinity was read off plots at the appropriate depths and displayed in a spreadsheet with the bottle results, 2004-06-sal-comp.xls. The results show so much scatter that it is hard to conclude anything. Because the depth of firing is only an estimate, it is important to look at data in a low salinity gradient. The local gradient was measured over the 4db block centered on a point 5db above bottom. There is only one casts with a very low gradient, 0.0005psu variation over 4db and we aren’t sure whether it is cast #8 or 9. The other 4 casts had variations of from 0.003 to 0.005psu over the same distance.

For the low-gradient sample, the pre-cruise primary and secondary salinity values were low by 0.007 and 0.008psu, respectively, if it is cast #9. However, if it is cast #8 then they are low by only 0.001 and 0.003psu. Using the post-cruise calibrations the salinity was high by 0.010 and 0.0135psu for cast #8, respectively or 0.016 and 0.019psu for cast #9. The casts with higher gradients all suggest larger differences for the post-cruise calibration. If we assume a linear drift with time we would expect that using the post-cruise calibrations would lead to primary salinity high by about 0.005psu and secondary salinity high by 0.006psu. It is possible that the drift was considerably higher over the last few months.

Looking at differences between the pairs of sensors the temperature channels are closer if the post-cruise calibration is used, but the conductivity channels are further apart. The secondary channel was lower than the primary if the pre-cruise calibrations were used, but higher if the post-cruise ones were applied. The secondary channel drifted more over the last 8 months than the primary channel did.

The primary salinity is less noisy than the secondary and is the one likely to be archived. The temperatures look better if the post-cruise calibration is used, so that will be done. If we trust the bottle results we would then subtract about 0.015psu. If we base recalibration on a presumption of linear drift of about -0.0006psu per month for 8.5 months, we would subtract 0.005psu. We have many reasons to doubt the bottle comparison, but are equally uncertain that linear drift occurred. So the wisest choice seems to be to split the difference. A choice of -0.01psu allows for the possibility of non-linear drift, gives some weight to the bottles, and leaves an uncertainty of ±0.005psu.

##### Conversion of Raw Data

The raw data were converted using configuration file 2004-06-CTD-post.con.

One cast had a non-standard name, CTD2004-06-00246.cnv; the time and position is right for cast #26 and there is no other file for that cast, so it was renamed 2004-06-0026.cnv.

All expected channels were present. The differences between the pairs of channels look fairly good for downcasts but during upcasts there is a misalignment of the two temperature and conductivity channels. The transmissivity and fluorescence look reasonable, except late in the cruise when the fluorescence looks very different between downcast and upcast.

##### WILDEDIT

Program WILDEDIT was used to remove spikes in the pressure and temperature channels only.  Parameters used were: 	Pass 1    Std Dev = 2

Pass 2    Std Dev = 5

Points per block = 50

##### ALIGNCTD

ALIGNCTD was not run. The deck unit serial number is unknown, but the results of recent cruises suggest that the secondary conductivity is being advanced even by older deck units. If this is not so, SHIFT can be used later to do this correction.

##### CELLTM

Tests were run on two casts running CELLTM with choices of (0.01,7), (0.01,9), (0.02,7), (0.02,9), (0.03,7),(0.03,7) and (0.0245,9.5) for (alpha, 1/beta). The primary sensors were not improved by any setting. Closer examination of the files suggests that the primary salinity was very noisy on the upcast. So no setting could be expected to make them look alike since they really are different. The settings used on other recent cruises using the same equipment were (0.01, 9) or (0.01, 7). The secondary data looks best with (0.03, 9). CELLTM was run using (0.01, 9) or (0.03, 9) for the primary and secondary, respectively, for all casts.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity and to calculate the descent rate. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

Four casts were plotted to check for agreement between the pairs of T and C sensors. The differences were about 0.0002Cº, 0.0004and 0.004psu for temperature, conductivity and salinity, respectively; the primary was higher than the secondary for temperature and lower for salinity and conductivity.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ data to IOS Headers.

CLEAN was used to add event numbers and to remove pad values in the pressure channel using linear interpolation based on record number.

##### Checking Headers

A header check was produced and showed there was an error in the headers for cast #14. The cause was scrambled headers in the CNV file. This was corrected and reconverted.

A header summary was produced and some inconsistencies in time were found between the headers and the log. The file 2004-06 log.txt confirms that the positions in the headers are consistent with the times in the headers.

The cruise track was plotted and the positions look reasonable.

There were station names in the headers for only a few casts, but the log has entries for a few more, though not for all casts. The entries from the log were added to the headers.

The average surface pressure is 1.8db. One cast has pressures around zero with conductivity very low. The pumps were off at that point, but the conductivity started to look like in-water values at about 0.2db. On the upcast the conductivity has in-water values very close to the surface. So the pressures appear to be reasonable.

As noted earlier there is a problem with pressure for cast #14. This appears to clear up while the CTD is soaking.

T0, T1, S0 and S1 were plotted for all casts. The pairs of sensors are reasonably close for most downcasts but differ more during upcasts due to poor alignment. This has been observed frequently in other data collected from this CTD on W.E. Ricker cruises.

Big problems were found in casts #28, 29 and 30. As noted in the log the fluorescence is odd, with very large differences between upcast and downcast. If it had occurred at only one site, it might possibly reflect real changes or drift during the cast, though the log positions don’t support the latter theory. But the same thing happens for 3 downcasts. Investigation showed that the secondary conductivity is also very odd from the surface to about 45db of the downcast, but the temperature looks ok. It is unknown on which pump the fluorometer was mounted, but it seems likely that it was on the secondary. Cast #34 also looks fairly odd in the top 10db but quickly settles down. The transmissivity is very low near the surface for cast 28. The upcast has higher values by about 15%/m. The downcast data is unusually noisy as well.

I have seen problems like this in Effingham Inlet and suspect that some sort of surface slick is clogging up the pump, but that by 50db there is sufficient pressure to clear the line. The low transmissivity may mean that some sort of film has also coated the lens. The fluorescence and transmissivity are not reliable for these casts. The primary temperature and salinity look fine.

##### SHIFT

Fluorescence

A shift of +24 records has been used for most cruises in the past, including the most recent uses of this particular equipment. A few casts were checked and that setting looks appropriate here as well, so it was applied to all casts.

Conductivity

For three recent cruises using this equipment a setting of -1.2 records was found effective in minimizing primary salinity spiking without oversmoothing. A test was run on cast #4 from this cruise and that setting was effective. The secondary salinity is bad for 3 casts and is unlikely to be archived so no alignment studies will be done.

All data were put through SHIFT using -1.2 for the primary channel only.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00	Minimum Salinity: 5

Pressure Tolerance: 1.0                                        Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over  11 points) will be deleted.

Drop rate applies in the range    10.00 dbars to 10 dbars less than the maximum pressure

Sample interval = 0.042 seconds. (from header)

COMMENTS ON WARNINGS: none

During cast #8 the CTD was lowered to 20db, then returned to the surface after which a full cast was run. DELETE selected the initial 20db lowering and patched it to the 20db-bottom section of the full cast. There is no note in the log to indicate why this was done, nor is any of the data obviously bad. To avoid patching a text editor was used to remove the initial downcast section. DELETE was then rerun on the edited data file.

##### DETAILED EDITING

The primary sensors were selected for editing.

Page plots were produced using T0,S0. These plots were examined for spikes and instabilities and used to guide the use of CTDEDIT. Where unstable features were clearly due to shed wakes the data were removed. Salinity was cleaned where large spikes occurred. Small spikes (mostly “overshoots” in large T gradient areas) were cleaned only if it was clear they were due to imperfect alignment of T and C. Small two-sided spikes in salinity will mostly be removed by metre-averaging.  Editing of salinity was done where it appeared that would not be the case.

The descent rate was kept high and steady. All casts required light editing only.

Note was made of the editing details in the relevant files.

##### Other comparisons

Previous experience with these sensors – This equipment has been used many times since the last recalibration, but there was generally little salinity sampling and a lot of scatter. There was a large drift to lower values and the post-cruise calibration showed a drift of about 0.0007psu per month for the primary salinity at conductivity values around 3.4

Historic ranges – All data fell comfortably within the historic ranges.

##### Recalibration

Based on the bottle data and the history of the instrument, the primary salinity is believed to be reading high. The primary salinity was recalibrated by subtracting 0.01psu using file 2004-06-recal.ccf. If the bottles are right this will be high by about 0.005psu and if the linear-drift estimate is right then it will be low by 0.005psu.

##### Special Fluorometer Processing

The COR files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. Casts #28, 29 and 30 were not included since the fluorescence data is bad. The files were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective.

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T, S, Transmissivity and Fluorescence profiles.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate and Flag channels were removed from all casts.

Transmissivity and Fluorescence were removed from casts 28, 29 and 30.

HEADER EDIT was used to fix formats and header entries and to add the following comments:

Transmissivity: The data are unedited except where records were removed in editing temperature and salinity.

Fluorescence:URU:Seapoint - The data are unedited except as above.

*************************************************************************

The salinity should be considered +/- 0.005psu due to doubts about the calibration.

For casts #28 to 30 a different comment was entered to explain why transmissivity and fluorescence channels were missing.

The final files were named CTD. The Standards Check routine was run and HEADEDIT adjusted until all format problems were removed.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

8. Salinity sample taken according to log, but bottle label indicates it is from cast #9.

29. Note in log about fluorometer drift.

28-30. Fluorometer and transmissivity look bad.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-06

Dates: Start: February 11, 2004 End: February 17, 2004

Location: Strait of Georgia / Queen Charlotte Strait / Hecate Strait

Vessel: W.E. Ricker

Party Chief: Cooke K.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0506 | No | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0506	Cruise ID#:		2004-06

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2668 | 20/06/02 | Factory | 18/11/04 |  |
| Conductivity | 2424 | 16/04/02 | “ | 19/11/04 |  |
| Secondary Temp. | 2374 | 20/06/02 | “ | 18/11/04 |  |
| Secondary Cond. | 2399 | 16/04/02 | “ | 19/11/04 |  |
| Transmissometer | 197 | 16/01/03 | IOS |  |  |
| Fluorometer | 2228 |  | IOS |  |  |