## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 27-May-2010 | An error was found in the calibration parameters used in processing this cruise. It is estimated that pressure is low by <0.5db, so no correction was applied. For details see file “Report on Calibration Errors for Pressure Sensor #77511, CTD 0585 “ in Osd_Date_Archive\Cruise_Data\DOCUMENTS |

## PROCESSING NOTES

Cruise: 2004-07

Agency: OSAP

Location: West Coast Vancouver Island/Strait of Georgia/Juan de Fuca Strait

Project: La Perouse, Covariability, Ecohab, GURD, Juan de Fuca, Georgia Strait

Party Chief: Tom Juhasz

Platform: John P. Tully

Date: April 26, 2004 – May 8, 2004

Processed by: Germaine Gatien

Date of Processing: June 8, 2004 – June 29, 2004

Number of original CTD casts: 83

Number of casts processed: 83

Number of rosette casts: 28

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with a Wetlab transmissometer (#333DR), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter OA-916D (#1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911+ (#0508). The salinometer was a Portasal models 8410 (#59724).

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was generally in good order with a good record of equipment used, but there were a few problems. There were changes of equipment noted in the log that did not correspond to when the changes appear in the configuration files. One configuration change was not recorded at all, so it appears it may have been changed accidentally.

There were errors in dates of calibration in the calibration control files and an offset was entered in the pressure calibration that was found to be incorrect.

The system upload time in the CTD files was out by one hour, presumably due to an error in the computer clock. Not remembering that the local time is -7 hours in summer would account for it.

There was no NMEA recorded for cast #2 and both the times in the log and the system upload times appear to be wrong so there is some uncertainty about this..

The descent rate was a little low for some shallow casts, but was generally kept high, minimizing the effects of shed wakes.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l from 0db–350db.

- ±0.05ml/l from 350 – 1500db

- unreliable below 1500db

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. There were changes to cables and configuration during the cruise so several configuration files will be needed to convert this data. The details of the changes are unclear so extra care will be needed to ensure all channels are converted.

Salinity and dissolved oxygen bottle calibration data were obtained. Flags and comments had been added as needed. There was no nutrient or chlorophyll sampling.

The cruise summary sheet was completed.

The history of the dissolved oxygen and conductivity sensors was obtained.

The calibration constants were checked for all instruments. The dates of the pressure and transmissivity calibrations were wrong. There was an offset of -0.6db in the pressure configuration. During 2004-04, which immediately followed this cruise, a pressure test was done which showed that there should be an offset of +0.2db. Initial checks show that an offset of -0.6db is clearly inappropriate, producing negative pressures for conductivity values that are clearly in-water. A choice of +0.2db looks reasonable so will be applied.

Five configuration files were prepared with the corrections, to be used as follows:

2004-07-CTD1.con – casts 2-5

2004-07-CTD2.con – casts 6-8

2004-07-CTD3.con – casts 9-37

2004-07-CTD4.con – casts 38-77 & 80-92

2004-07-CTD5.con – casts 78 & 79

##### Conversion of Raw Data

All data were converted using the configuration files noted above.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. There are problems in the transmissivity data for many casts; these should be examined carefully later and the channel removed from some casts if necessary.

Rosette files were converted using a start time of -2s and duration of 5s.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2

Pass 2    Std Dev = 5P

Points per block = 50

##### ALIGNCTD

The deck unit was one of the newer ones that advance the secondary conductivity channels so ALIGNCTD was not run.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.01,7) , (0.01,9),  (0.02,9), (0.03,9), (0.02,7) and (0.03,7. The setting of (0.02,7) improved the primary data. The secondary sensors have a lot of fine-scale noise in their signals and it was hard to see much effect from running CELLTM, but (0.02, 7) improved one cast a little and did not make others look worse.

All casts were put through CELLTM using (0.02,7) for both conductivity sensors.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

The differences between T and C sensors were unusually noisy. The descent rate was kept very high. The secondary salinity is full of fine-scale noise with average excursions of ±0.002psu at 2000db, while the noise level of the primary is about ±0.0002psu. The secondary sensors were used during 2004-05 when large differences between sensors were also noted, but the primary sensors are different for this cruise.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 7 | 375 | +0.0007 | +0.0002 | +0.0013 | Very high, ok |
| 51 | 375 | -0.001 | +0.0002 | +0.0045 | Very high, ok |
|  | 1150 | -0.0015 | +0.00022 | +0.0044 | Very high, ok |
| 55 | 375 | -0.0023 | +0.0006 | +0.0045 | Very high, ok |
|  | 1150 | -0.0012 | +0.00025 | +0.0044 | Very high, ok |
| 66 | 375 | -0.005 | -0.00001 | +0.005 | Very high, ok |
|  | 1150 | -0.0015 | +0.0002 | +0.0037 | Very high, ok |
|  | 1950 | -0.0044 | +0.0002 | +0.003 | Low, very noisy |

There is no clear evidence of pressure-dependence or time-dependence in these differences.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

A spreadsheet was produced with altimeter reading taken from the header. All values were less than 15m. Ten casts were checked to see that the data was interpreted correctly and no problems were found.

The rosette files were converted to IOS files and put through CLEAN to add event numbers; the output was named BOT.

The system upload times and NMEA times differ by 1 hour for all casts except #2 for which there is no NMEA time. The NMEA time appears to be correct (except possibly for cast #7.) It is likely that the computer clock was set based on PST instead of PDT.

The NMEA time and positions for cast #7 were scrambled in the SeaBird headers and had to be fixed in the ROS file before conversion. The NMEA time is 6:29:30, but the log says 07:24. There is a note in the log that the time stamp from NMEA was 7 minutes slow at the end of the cast. However, the log time is clearly wrong since it is later than the cast that follows. The log time was probably taken from the computer and the 7 minute difference is more likely due to computer clock error than NMEA error. It will be assumed that the NMEA time is correct.

There is no NMEA time or position for cast #2. The log positions were put into the file after conversion. The time in the header was the system upload time of 3:00 which is also what is in the log. This time is clearly wrong since the end of the cast is later than the beginning of the next cast. It seems likely that the time should be 2:00 since all other system upload times differ from the NMEA by 1 hour. The header was edited to change the time to 2:00, but a note should be put in the header to indicate that the time is estimated and the position taken from the log.

All BOT files were plotted and significant outliers were found only in cast #71 for the secondary salinity in bottle #5. CTDEDIT was used to clean the outliers in the IOS files; it was rerun through CLEAN with the output file named BOT.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The header check and header summary were run and the following problems were found:

- The time for cast #2 is suspect with a very low speed estimate between cast #2 and 3. Using a time of 3:00 leads to an impossibly high speed. Using 2:00 looks ok. The ship was hove-to before this cast so there are no other time checks available. There was no bridge log. The header time was changed from 0300 to 0200 UTC.

- The station name was missing from cast #38; this was corrected.

- Cast #38 rosette file was missing – it was reconverted.

- Cast #39 was missing. The DAT file on the CD-ROM appears to be corrupted and could not be copied. Steve Mihaly provided an uncorrupted version of the file. This was put through steps #3 through #9.

The speed check confirms that the NMEA time is reasonable for cast #7.

The cruise track was plotted and no further problems noted.

The average surface pressure is 2.4db which is reasonable.

A few CLN files were examined on-screen. There was no evidence of serious problems in T and S; both salinity channels have occasional spikes. The transmissivity was bad for a few casts and fluorescence looked bad in part of cast #72. Dissolved oxygen generally looked ok, but spiky for one cast The data will be examined more carefully after bin-averaging to decide what data should be removed.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet 2004-07.csv was edited to remove loop samples and to enter cast numbers and was saved as 2004-07-salinity.csv. There was a flag channel, but no flags entered.

There is an inconsistency about the bottle at cast #67. It is entered as #213 (10m) in the spreadsheet, but the rosette log indicates the sample was taken for #214 at 5m. The spreadsheet was edited to remove the loop samples and add the cast numbers. That file was converted to individual SAL files.

The dissolved oxygen files (*.add) were created by the analyst and had a quality channel but no comments for flagged values. The files were copied to the HYDRO folder and comments were added as required based on information on the rosette sheets.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets.

The SAL and DOX files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 11. COMPARE

Salinity

COMPARE was run and differences above 500db were excluded. There is a lot of scatter above 500db. Below 500db the salinity appears to be slightly pressure-dependent and time-dependent. However, if one outlier is removed the trend lines look much flatter. The primary salinity is low by about 0.0003psu and the secondary salinity is high by about 0.0025psu. No recalibration is necessary for the primary salinity. The secondary salinity is very noisy so will probably not be selected; if it is, it should be recalibrated. (See 2004-07-sal-comp1.xls)

Dissolved Oxygen

COMPARE was run using titrated dissolved oxygen data and plots were prepared of differences between CTD and titrated values versus pressure, DOX and cast number. The best fit was versus DOX value. As expected the sensor does not perform well below 1500db so 2 deep values were dropped from the comparison as were 2 values flagged by the analyst. Four other outliers were removed.

Cast #78 contained many severe outliers; the problem appears to be with the CTD data, not the bottles. Looking at consecutive casts it became evident that DOX for #78 and #79 were not like their neighbours with values increasing with depth and staying fairly steady at depth. One of the casts is off the West Coast (LC01) and the other in Juan de Fuca Strait (JFD7) so no good explanation for such variations is likely. The transmissivity signal was all zeros for those two casts as well. The configuration file was different for those two casts with no explanation in the log book, so it seems likely this was a mistake. Attempts were made to convert the data using different configuration files, and replacing the dissolved oxygen channel with transmissivity produced sensible values. Entering dissolved oxygen in other voltage spots did not produce anything sensible. The data was reconverted and put through the preliminary steps. COMPARE was not rerun since there will be no useful information in the new files.

From COMPARE the fit is:

DOX (Titrated) = 1.0336 * DOX (CTD) + 0.0751

The fit against cast number suggests no significant time dependence. (See 2004-07-dox-comp1.xls)

For 2004-05 the results were:

DOX (Titrated) = 1.0526 * DOX (CTD) + 0.0291

Outliers from all these comparisons were examined; all had been flagged already, or were only mild outliers or it appeared that the CTD value might be wrong rather than the bottle value. The CTD dissolved oxygen channel for pressure > 1500db is affected by hysteresis with errors on the order of 0.3ml/l at 2000db.

##### 12. SHIFT

Conductivity

Tests were run on the primary conductivity channel to determine what shift reduces instabilities best without oversmoothing; results were examined in T-S plots after running DELETE. The best results overall came from an advancement of -0.2 records. That would mean a net advancement of about +0.058s, since the channel was already advanced by +0.073s earlier.

All casts were put through SHIFT using -0.2 records for the primary conductivity. (Output *.SHFC)

Dissolved Oxygen

Tests were run using advancements of +90 records to +130 records and the best overall choice was +120 records (~5s). That shift was applied to all casts. (Output: SHFO)

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. Tests were run using shifts of +24, +36 and +48 records and the best results overall were with +36 records. A shift of +36 records (1.5s) was applied. (Output: SHFFL)

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing since the primary salinity is closer to the bottles and there is a lot of fine-scale noise in the secondary. There are some areas of spikiness in the primary salinity that is not in the secondary, but editing will remove most of the problems.

Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

The descent rate was fairly low for a few casts but was kept high for most. There was some spiking in the salinity, but these were fairly small and easily cleaned within CTDEDIT.

All casts required some editing, but only casts #2, 38, 44 and 54 were edited heavily. Those casts had very noisy descent rates leading to serious corruption of data by shed wakes.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors –

The primary sensors have been recalibrated since their last use. The secondary sensors were last used for 2003-39 and 2003-41 (October and December 2003) when the salinity was within 0.001psu of bottles. On an earlier cruise the salinity was low by 0.003psu. The dissolved oxygen sensor has been giving good results since it was repaired in January 2004.

Historic ranges – The temperature data was above the historic minima for some casts in the Strait of Georgia between 10 and 50m, especially on the western side. Similar excursions were noted at about the same time during 2004-08, and different equipment was used then so this is not indicative of instrumental problems. There were very high winds at about the time of these observations.

##### Initial Recalibration

File 2004-07-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun. The fit of differences was much improved versus pressure, time and DOX values, with the differences looking quite flat against all variables. (See 2004-07-dox-comp2.xls.)

##### Final DOX comparison

Another run of COMPARE was made using metre-averaged, thinned downcast files. There is a lot of scatter so the choice of what to exclude does affect the results. But when differences greater than 0.2ml/l are excluded, the distribution shows little pressure dependence and an average offset of about 0.03ml/l with the CTD reading a little high. The results for 2004-05 were the same. (See 2004-07-dox-comp3.xls.)

Since this difference does not show pressure-dependence it seems not to be due to alignment problems or hysteresis. It is presumed that this error is due to a slight, systematic difference in flow rate between upcast and downcast. The COR1 files were recalibrated by subtracting 0.03ml/l using file 2004-07-recal2.ccf. (Output: COR2)

COMPARE was rerun using the recalibrated thinned files and the average difference was +0.001ml/l. (See 2004-07-DOX-comp4.xls)

##### Special Fluorometer Processing

The COR2 files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. The FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles. Separate profile plots were prepared with Temperature, DOX, FL and Transmissivity versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts. For casts #1-37 the transmissivity channel was also removed.

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

For casts #1 through #37 the transmissometer was not working properly and

was removed.

The SBE Dissolved Oxygen data in the CTD files should be considered:

•	±0.2ml/l from 0db–350db.

•	±0.05ml/l from 350db – 1000db

•	unreliable below 1500db

For cast #2 an extra note was added as follows:

There was no NMEA data for this cast. The positions were taken from

the log book and the time was estimated, based on the system upload

time minus one hour, since an error of one hour was found in the system upload time for all other casts.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag. For casts #1-27 the transmissometer channel was also removed.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods (2004-07-bot-hdr.txt). Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1-11. Log notes problems with transmissometer; changes in configuration tried, still problems.

29. CTD dropped on deck.

30. Cables changed before this cast and altimeter added, con file changed according to log, but altimeter does not appear in con file until cast #38.

39. The DAT file was corrupted on the CD-ROM in the cruise file; a clean version was provided by Steve Mihaly.

78/79. Altimeter not working. (channel not in con file). Con file had DOX and Transmissometer reversed – no DOX data logged.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-07

Dates: Start: April 26, 2004 End: May 8, 2004

Location: West Coast Vancouver Island/Juan de Fuca Strait/Strait of Georgia

Vessel: John P. Tully

Party Chief: Juhasz, T.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-07

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 20/12/03 | Factory “ |  |  |
| Conductivity | 1763 | 19/12/03 | “ |  |  |
| Secondary Temp. | 2095 | 05/09/03 | “ |  |  |
| Secondary Cond. | 1764 | 05/09/03 | “ |  |  |
| Transmissometer | 333DR | 06/01/04 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13/01/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |