## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 19-May-2006 | Added CHL data to the rosette files G.G. |
| 29-Nov-2004 | Added nutrient data to the rosette files. J.L. |

## PROCESSING NOTES

Cruise: 2004-08

Agency: OSAP

Location: Juan de Fuca Strait / Strait of Georgia

Project: SoG / JdeF

Party Chief: Peña A

Platform: Vector

Date: April 26, 2004 – May 1, 2004

Processed by: Germaine Gatien

Date of Processing: 7 May 2004 – 8 June 2004

Number of original CTD casts: 	69

Number of casts processed: 69

Number of rosette casts: 25

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea transmissometer (#498DR), a Seabird Dissolved Oxygen Sensor (#0615), Altimeter (#1024), a PAR sensor (#4615), a SPAR sensor (#16504) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0424). The salinometer was a Portasal model 8410 (#59724).

# SUMMARY OF QUALITY AND CONCERNS

The pressures appear to be within ±1db, but there was no pressure test to confirm this. It is recommended that a pressure test be done once per cruise, where possible.

Most of the sample numbers for cast #59 were changed since they had been used for two different casts. When nutrients and chlorophyll are added to the CHE files, care will be needed to ensure the data go with the right bottles.

The secondary temperature data were very spiky for many of the casts and the secondary conductivity had huge excursions to unbelievable values; the return to normal values was gradual and slow, rendering large segments of data unusable.

The transmissivity, dissolved oxygen and fluorescence are unedited, except where records were removed in the editing of temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l in the top 150m

- ±0.1ml/l from 150m down

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity and dissolved oxygen bottle data were obtained. Flags and comments had not been added to the oxygen files as needed, but were provided at a later stage. The extracted chlorophyll and nutrient data were not yet available.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The only errors found were in the dates of the pressure and transmissometer calibrations. Those were corrected and the resulting file named 2004-08-CTD.con.

The history of the conductivity and pressure sensors was found. The dissolved oxygen sensor has been used only once before.  The pressure offset for this CTD drifted through 2003, so checks will be done on whether an offset is needed.

##### Conversion of Raw Data

A pressure test was not done during the cruise. A few files were converted to check on what pressure offset is needed. During 2004-11 using the same sensor there were some negative surface values that looked to be “in water” and an offset of +0.5db was applied. For this data there are no negative values for pressures; when pressure is about 0.2db T and C have values which appear to be just barely “in water” or possibly out of water, so the pressures are close to correct. No offset will be applied.

The raw data were converted using the configuration file 2004-08-ctd.con. A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. However, there is severe spiking in the secondary temperature and conductivity for many casts.

Rosette files were converted using a start time of -4s and duration of 5s. Usually we use -2s as the start time, but there have been problems with other VECTOR cruises due to the CTD starting up so quickly that upcast data were captured in the rosette file.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channels only. (Wildedit was ineffective at reducing the problems in the secondary temperature channel.) Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

In recent use of this equipment the secondary conductivity appears to have been advanced, so ALIGNCTD will not be run to do that. The alignment will be fine-tuned later using SHIFT.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.1, 7), (0.01, 9), (0.02, 9), (0.03, 9), (0.02, 7) and (0.03, 7). The best results varied from feature to feature and cast to cast. For the primary channels the best results overall were with (0.03, 7) and for the secondary channel (0.02, 7).

CELLTM was run using (0.03, 7) on the primary channels and (0.02, 7) on the secondary.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

Most of the casts were relatively shallow.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 27 | 300 | 0.0007 | -0.00003 | -0.0009 | Good |
| 41 | 300 | ~0 noisy | -0.00005 | -0.0005 | Good |
| 41 | 350 | 0.0003 | ~0 noisy | -0.0005 | Good |
| 71 | 300 | 0.0005 | -0.0001 | -0.0015 | Good |
| 71 | 350 | 0.0005 | -0.00008 | -0.0015 | Good |
| 71 | 430 | ~0 x noisy | -0.0002 noisy | -0.0015 | Good |

The differences are fairly constant and depth-independent below 300db. The differences are similar to those observed during 2004-11.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers.

CLEAN was run to add event numbers to the headers and to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files.

CLEAN was run to add event numbers and the output file was named BOT.

All BOT files were plotted and many outliers were found in the secondary data only. It was decided not to edit these files since the secondary salinity was generally very poor and unlikely to be archived. If a choice is later made to use the secondary data, the BOT files should be edited first.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary, a header check and cross-reference listing were produced. The only problem noted is a discrepancy in time between the Daily log and headers for cast #33. The Bridge log agrees with the headers and there is mention in the log of a computer clock error. The header is presumed correct.

The average surface pressure is 1.5db which is reasonable.

##### BOTTLE FILE PREPARATION

The BOT files were averaged to enable an addsamp file to be created.

The addsamp file was edited to add sample numbers taken from the rosette sheets. In doing that the following problems were noted:

- There were 24 bottles in the BOT file for cast #1, but only 13 bottles recorded in the Daily log and Rosette log. The “phantom” bottles were all at the surface. Those bottles were removed from the addsamp file.

- For casts #7, 10, 15, 18 and 19, bottle #7 was fired, but no sample was taken because the bottle was broken; those bottles were removed from the addsamp file.

- The sample numbers given for cast #61 overlap those given for cast #59. Those that were repeated were renamed in the addsamp list for cast #59 by putting a “9” before the other digits. There was no salinity sampling for cast #59 so the sal spreadsheet does not need adjusting. The sample numbers were changed for DOX cast #59, in the initial comparison. The sample numbers should later be fixed in the other sampling files – the rosette sheet indicates there was sampling for both casts for chl and nuts.

An initial run of COMPARE turned up some other problems in the ADD files:

- There is no oxygen data for cast #56, 61 and 64, although the rosette sheets indicate such sampling was done. Sheila Toews could not find any data – she will check with Doug Anderson. This data will be assumed to be lost, but can be added later if found.

- Cast #7 dissolved oxygen data were incorrectly named as “5”; that was changed in the ADD file before merging with the BOT files..

- Cast #34 had a sample #1230 – this was changed to #123.

- An error in cast #2 format was fixed.

- An error in cast #52 (wrong sample number at 25m) was fixed.

- Errors in sample numbers in casts #41 and #59 were fixed.

The salinity analysis spreadsheets was edited to add event numbers and sorted on sample number; they were converted to individual SAL files. There was only one flagged value but no comment to explain it. The salinity analysis sheets were found and the following flags assigned based on comments found on the sheets:

- Cast 47 – samples #179, 180 and 181 – “d” no liner

- Cast 50 -  sample #198 -  “d” liner not in all the way

- Cast 61 – sample #238 – “d” liner not in all the way

- Cast 69 – sample #299 – “d” liner not in all the way

- Cast 69 – sample #300 – “d” liner not in all the way

The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files. Those files were edited to remove the bottles that had not been sampled and to reorder cast #59 on sample number. They were then merged with the SAL and ADD files in two steps. (Output: MRG1 and MRG)

##### 11. COMPARE

Salinity

COMPARE was run. When a few outliers and all values from above 200db were excluded, the primary salinity was found to be high by 0.0002 and the secondary low by 0.0006psu. There is a lot of scatter. While there is some suggestion of drift with pressure and time, the evidence is slight, and the time variations more likely due to geography than instrumental. Both the primary and secondary differences vary in the same way. The only two outliers that were large enough to flag had already been flagged based on the analyst’s comments (cast #47, samples 180 and 181); a note was added to the header of the SAL files about the comparison and Merge was rerun. There were two minor outliers from cast #44 but they were not considered large enough to flag the data given the local variability. (See 2004-08-sal-comp1.xls.)

Dissolved Oxygen

COMPARE was run and plots were prepared of differences between CTD and titrated values versus pressure, DO and cast number. When values from the top 5db were excluded the differences were fairly flat against pressure and time. The most useful fit is probably that versus DO value. A few outliers were identified from that plot. After a few outliers were excluded and all data above 5db the fit versus DO is:

DOX (Titrated) = 1.0154* DOX (CTD) + 0.0608

This sensor has been used only once before, during 2004-11, when the fit was:

DOX (Titrated) = 1.0212* DOX (CTD) -0.0168

Some of the difference between the two cruises would be due to the fact that the waters sampled were quite different. No outliers were extreme enough to justify flagging. (See 2004-08-dox-comp1.xls)

The above fit will be used to recalibrate the CTD DOX channel.

##### 12. SHIFT

Conductivity

During 2003-26, -27, -31, -32 and 2004-11 when the same CTD and similar attached instruments were used, the primary conductivity was shifted by -0.2 records. Tests were run on several casts with shifts of from 0 to -0.8 records, and the best results overall were with settings of -0.4 or -0.5 records.

All casts were put through SHIFT using -0.4 records for the primary sensor.

Tests were not done for the secondary channel because the data quality was very poor and unlikely to be selected for the archive.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The difference between the two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the results varied from +1s to +1.5s. Overall a shift of 1s (+24 records) looked best. This is the value that has been used for most cruises in the past, but during 2004-11 +36 records was used. (The average descent rate was lower than usual during 2004-11 which may have affected the flow through the fluorometer.)

The fluorescence channel was advanced by 24 records for all casts.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. Tests were done shifting the DO by +80 to +140 records and the best results were found using +110 or +120 records. Since the best results for 2004-11 (with the same sensor) was +120 records (+5s) that shift was applied to all casts.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min

Maximum Surface Pressure (relative): 10.00

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing because of the severe spiking in the secondary.

Page plots were produced using (T0,S0). These were used to guide the editing. The descent rate was extremely noisy from cast #5 to #11 due to very heavy seas. For other casts it was fairly steady.

Most casts had spikes in salinity. If the temperature data looked ok at the level of the spike the salinity was cleaned. If the temperature was also spiky the records were removed.

Cast #48 was marked by many small instabilities caused by salinity noise. Most of the noise is two-sided and should largely disappear with metre-averaging. Editing of these features was minimal as it was rarely clear what was good data and what bad.

The following cast required extensive editing: #5-11

All other casts required only light editing. Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – Both sensors were used for 2004-11 in April 2004 when the primary salinity was found to be low by 0.0003 and the secondary low by 0.0012psu. During 2003-31 in September 2003 the primary salinity was found to be high by about 0.0035psu and during 2003-37 in August when it was high by 0.004psu. The secondary conductivity during 2003-31 was found to be high by 0.0032.

The dissolved oxygen sensor has been used only once before during 2004-11 when it was recalibrated using the fit of differences versus DOX:

DOX (Titrated) = 1.0212* DOX (CTD) -0.0168

That cruise included many samples with DO < 2ml/l whereas 2004-08 has only one such sample, so there is reason to expect quite different fits.

Historic ranges – Historic ranges were available for casts #2 to #48 only; some temperatures were above the historic maximum between the surface and 50m for casts # 26, 27, 28, 35, 42 and 48 and salinities were sometimes below the historic minimum for casts #22, 23, 26, 27 and 28 in the top 50m. There was a big wind storm shortly before these observations and the excursions from the historic range are not considered indicative of instrument malfunction.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

No recalibration is needed for the primary salinity.

File 2004-08-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 11 (corrected DOX = 1.0154* DOX (CTD) + 0.00608); it was applied to the EDT, MRG3 and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for DOX and the corrections were found to be made correctly. (See 2004-08-dox-comp2.xls.)

### Final Dissolved Oxygen Recalibration

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but in the past there have sometimes been significant errors due to response-time problems. To check for such errors the downcast CTD data (after SHIFT and CALIBRATE) was compared with the bottles from the upcast. A set of downcast files were prepared by bin-averaging (0.5db bins) the COR1 files and thinning the data to the depths generally used for bottles. COMPARE was run comparing those files with the bottle data in the MRG files. The fit of differences versus pressure was found excluding a few outliers and values from above 10db. The CTD is higher than the bottles by an average of +0.01ml/l with only slight pressure dependence. As found for this instrument during 2004-11 time-response is not a serious problem and no further recalibration will be done. (See 2004-08-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

A spreadsheet was prepared with altimeter readings taken from the header to ensure that the calculation worked properly. All values look reasonable including that for cast #53 which was reported to have had noisy data near the bottom. A few casts were checked in detail and all appeared to be correct.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T & S profiles. Separate profile plots were prepared with Temperature, Transmissivity, DOX and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment using file 2004-08-hdr.hdr and to fix formats and units.

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats and to add a standard comment about quality flags. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

3. Problem with primary conductivity.

4. Problem with primary conductivity, connections cleaned and re-greased.

5. Very bad weather, slow descent due to slack in wire.

6. Rosette hit side of ship, bottle #7 broke.

8. High descent rate for first 75m

13. First attempt pumps off. Redone, overwriting original file.

33. Computer crashed – restarted file. Clock was erratic.

36. Altimeter noisy at depth.

37. S and T differences high at beginning.

40. Secondary T and C cables cleaned after this cast.

53. Altimeter noisy at bottom.

65. Triton plungers were left on, so came back up from 75db and reran the cast, overwriting original file.

68. Altimeter noisy at bottom.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-08

Dates: Start: April 26, 2004 End: May 1, 2004

Location: Juan de Fuca Strait / Strait of Georgia

Vessel: Vector

Party Chief: Peña A

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550

Cruise ID#:		2004-08

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 22/04/03 | Factory |  |  |
| Conductivity | 2173 | 24/04/03 | “ |  |  |
| Secondary Temp. | 2968 | 22/04/03 | “ |  |  |
| Secondary Cond. | 1729 | 24/04/03 | “ |  |  |
| Transmissometer | 498DR | 07/08/03 | IOS |  |  |
| Dissolved Oxygen | 0615 | 09/03/04 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| PAR | 4615 | 15/12/00 | ? |  |  |
| SPAR | 16504 | 02/01/04 | ? |  |  |
| Pressure Sensor | 75636 | 06/04/99 | Factory |  |  |