## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |

## PROCESSING NOTES

Cruise: 2004-10

Agency: OSAP

Location: Gulf of Alaska

Project: Line P

Party Chief: Robert M.

Platform: John P. Tully

Date: June 1, 2004 – June 18, 2004

Processed by: Germaine Gatien

Date of Processing: July 30, 2004 – August 18, 2004

Number of original CTD casts: 86

Number of CTD casts processed: 85

Number of rosette casts: 32 (multiple bottles) 27 (single Niskin at 5m)

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0585) was mounted with Wetlab transmissometers (#333DR for casts 1-34 and #498DR for casts 34-109), a Seabird Dissolved Oxygen Sensor (#0047), Altimeter OA-916D (#1024), a Seapoint Fluorometer (#2229) with a 10X cable and a Biospherical PAR sensor (#4615). The deck unit was a model 911 (#0424) and the logging computer was #FS02. The salinometer was a Portasal model 8410 (#58879). The deck unit was #508. A mid-ship winch was used.

# SUMMARY OF QUALITY AND CONCERNS

The CTD log was in good order. Notes from the chief scientist on problems during the cruise were very helpful.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l from 0db–150db.

- ±0.1ml/l below 150db – 300db

- ±0.05ml/l from 300 – 1500db

- unreliable below 1500db (that channel was removed from the final CTD files for P>1500)

The transmissivity has a lot of spikes especially at depth. Data was removed from casts #27, 32 and 89 below 1500, 3400 and 1400db respectively, because no useful signal could be found in the noise.

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained. Together with the notes from the Chief Scientist this gave an excellent summary of problems encountered during the cruise.

Bottle salinity and dissolved oxygen data were obtained; flags and comments had been added as needed. Final chlorophyll and nutrient bottle calibration data were not yet available. There is confusion about sample #50 – this will be investigated later.

The cruise summary sheet was completed.

The histories of the conductivity, pressure and dissolved oxygen sensors were obtained.

The calibration constants were checked for all instruments. The dates of the calibrations for the pressure sensor and transmissometer #333DR were wrong. There was an offset of -0.6db in the pressure configuration. During 2004-04, 2004-05 and 2004-07 pressure offsets of +0.2db were found appropriate; in the case of 2004-04 a careful surface pressure test was done. An offset of +0.2db will be applied and the results checked later.

Four calibration files were prepared with corrections to the dates and pressure offset.

2004-10-CTDa.con applies to casts 1-28 (PAR, Trans #333DR)

2004-10-CTDb.con applies to casts 29-34 (No PAR and Trans #333DR)

2004-10-CTDc.con applies to casts 35-36 (No PAR, Trans #498DR)

2004-10-CTDd.con applies to casts 37-109 (PAR, Trans #498DR)

For many casts the PAR channel existed but the instrument was not actually mounted. The channel should be removed from those casts later in the processing.

##### Conversion of Raw Data

All data were converted. A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. The altimeter data is noisy near the bottom, though a useful signal is apparent, at least visually, amidst the noise.

Rosette files were converted using a start time of -2s and duration of 5s. The station name for cast #12 was changed to P8 in both the CNV and ROS files.

The rosette files were then converted to IOS SHELL files. In doing this problems were noted for 4 files. The NMEA headers were scrambled for casts #4 and 38. These were corrected in both the CNV and ROS files, and the ROS files were then reconverted. There are also problems in casts 104 and 109 with the pump status channel.

- For the ROS file of cast 109 this occurs only in 2 records at the end of the file. It is not known why the data overflowed rather than convert as “0”. A text editor was used to replace the overflow with “0” in the two records.

- For the ROS file for cast #104 all data is corrupted. The cast was run with the pumps off on purpose. Again, a text editor was used to replace the ** with 0. It is assumed that this cast will not go into the archive, but it will be processed in case it is needed for calibration purposes.

##### WILDEDIT

Program WILDEDIT was used to remove spikes from the pressure channel only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit was one of the newer ones that advances secondary conductivity channels so ALIGNCTD was not run on conductivity.

The dissolved oxygen channel was aligned by +5s based on tests on a few casts. After this step DO for the up and downcasts are much closer. Again, this can be fine-tuned later if found necessary.

##### CELLTM

Tests were run with (alpha, 1/beta) set to (0.01,9),(0.02,7), (0.03,7), (0.02,9), (0.0245,9.5) and (0.03,9). All choices improved the data, but the optimal setting varied from cast to cast and even within a cast. A setting of (0.03,9) was best overall for both channels.

All casts were put through CELLTM using (0.03,9) for both conductivity channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 9 | 2000 | -0.0008 | +0.00032 | +0.005 | High |
| 32 | 2000 | -0.0005 | +0.00025 | +0.004 | High, noisy |
| 51 | 2000 | -0.0004 | +0.00028 | +0.004 | High |
| 94 | 2000 | -0.0006 | +0.00028 | +0.0045 | High |

All the differences were very noisy. The temperature difference was much more pressure-dependent than usual, but were not so large as to cause great concern. For cast #32 the temperature differences at 4000db were about -0.0002 compared with -0.0005 at 2000db. Conductivity showed little variation with pressure.  The secondary salinity is full of fine-scale noise with average excursions of ±0.002psu at 2000db, while the noise level of the primary is about ±0.0005psu. The secondary conductivity had a noise level of ±0.00015 S/m whereas the primary has noise on the order of ±0.00003.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Warnings in the log led to the discovery that the headers were scrambled for cast #74; this was fixed in the CNV file and reconverted.
The rosette files were converted to IOS files. CLEAN was run on the rosette files to add an event number and the output was named BOT.

All BOT files were plotted and significant outliers were found only in cast #34, bottle #7. CTDEDIT was used to clean the secondary salinity for that cast. The output file was renamed BOT.

##### Checking Headers

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The header check and header summary were run. The station name for cast #9 was incorrect; this was fixed in the IOS and CLN files.

The cruise track was plotted and no problems noted.

The average surface pressure is 2.1db, a reasonable value though slightly lower than usual for the Tully. A few casts were checked to see if conductivity looked reasonable near the surface. For cast #7 when pressure was 0.2db at the end of the cast, the conductivity looked like “in-water” values. For casts #9 the conductivity at 1db looked like near-surface “in water” values. The pressure appears to be close to correct and will not be adjusted further.

A few casts were examined on screen. No problems were noted.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheet was edited to remove the loop samples and to add a few cast numbers, and was named 2004-10sal.csv. The file was reordered on cast #. That file was converted to individual SAL files. Comments for 4 flagged values were added to the headers of the SAL files for casts #9, 10, 12 and 14. The sample for cast #12 is said to be from the wrong Niskin bottle. After running COMPARE this will be revisited to see if it is obvious where this bottle came from.

The dissolved oxygen files (*.add) were created by the analyst and have quality channels and comments for flagged values.

The extracted chlorophyll and nutrient data was not yet available.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. Cast #9 had 2 bottles but there is a sample number for only the shallow one. The other was taken for UBC. Similarly, for cast #38, the first 5 bottles were for UBC and do not have sample numbers. For cast #96, 7 bottles were fired as a test. The log indicates that one sample was taken, but no data was found, so it was probably not analyzed.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1, MRG)

##### 11. COMPARE

Salinity

COMPARE was run. There is a lot of noise in the comparison for both sensors with suggestions of pressure-dependence and time-dependence, but not of a systematic nature. Depending on what data is selected the differences run from -0.0022 to -0.0009psu for the primary and -0.0008 to +0.0027psu. The results from the deepest samples shows the primary low by about 0.0009 and the secondary high by 0.0027psu. (See 2004-10-sal-comp1.xls)

There were two extreme outliers. Sample #50, cast #12 had been flagged by the analyst; the flag was changed to “e”, the value was replaced with -99, but the original value was reported in the header comment. In the other case, cast #109, the problem appears to be in the CTD data, not the titrated sample.

Dissolved Oxygen

COMPARE was run using titrated dissolved oxygen data and plots were prepared of differences between CTD and titrated values versus pressure, DOX and cast number. The best fit was versus DOX value. Seven outliers were identified and removed from the comparison, as were values from 1500db downwards since the sensor is not expected to give good values at those pressures. Using those values the fit is:

DOX (Titrated) = 1.0772 * DOX (CTD) + 0.0331

The fit against cast number suggests time dependence, but this is probably due to geographic variations rather than calibration drift, and deep casts versus shallow affecting the range of DO. When only the casts from Line P itself were included the result was:

DOX (Titrated) = 1.0698 * DOX (CTD) + 0.0143

The calibration from other recent cruises using this equipment were:

DOX (Titrated) = 1.0336 * DOX (CTD) + 0.0751 (April, near-shore)

DOX (Titrated) = 1.0526 * DOX (CTD) + 0.0291 (February, Line P)

The results are reasonably close to the Line P results, but some temporal drift is likely occurring.

(See 2004-10-dox-comp1.xls)

The above fit will later be used to recalibrate the CTD DOX channel.

A few outliers were investigated. Samples #50 and 227 had already been flagged, but the flags were changed from “d” to “e” and the values were replaced by pad values. The original values were written in comments in the headers. Two others were from near the surface and the problems might well be with the CTD sensor rather than the titrated samples.

##### 12. SHIFT

Conductivity

Tests were run on the primary conductivity channels to determine what shift reduces instabilities best without oversmoothing; results were examined in T-S plots after running DELETE. The best results came from an advancement of -0.2 records. That would mean a net advancement of about -0.065s, since the channel was already advanced by +0.073s earlier.

All casts were put through SHIFT using -0.2 records for the primary conductivity. (Output *.SHFC)

No shift was applied to the secondary channel since it is unlikely to be used for the archive.

Dissolved Oxygen

A shift equivalent to +120s had been applied earlier using ALIGNCTD. Tests were run using advancements of -20 records to +20 records to determine if further shifting the data would improve the data. The best choice varied from cast to cast and from one depth to another. Overall the choice of +120 records (~5s) was found best, so no further shift was applied.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets are treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. The value found was 1 to 1.5s. A shift of +24 records (1s) was applied. This is the shift that has been used in most other cruises. (Output: SHFFL)

##### DELETE

During the downcast sections of casts #19, 42, 79 the pump was not on and the CTD was returned to the surface and the cast rerun. DELETE patched the two sections together. A text editor was used to remove the initial sections and DELETE was rerun.

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing since the secondary sensors have a lot of fine-scale noise.

Page plots were produced using (T0,S0). These were used to guide the editing. On-screen plots of descent rate and pump status were also used.

All casts required some editing, and about half required heavy editing. The descent rate was extremely noisy for many casts, but the average CTD descent rate was kept very high for the deep casts which reduced the incidence of shed wakes. The shallow casts in the eddy region required the heaviest editing.

Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary sensor has produced salinity within 0.001psu on two other cruises in 2004, but there were not many bottles for one of those cruises. The secondary sensors gave salinity high by 0.0017 and 0.0025psu on the same two cruises.

Historic ranges – Some data from a region of eddy activity fell outside the historic ranges, with high temperatures and low salinity at depths between 400 and 2000db depending on the cast. This is not considered indicative of instrumental problems. Temperatures at stations P16 & P17 were at the low end of the climatology around 1000db. Temperatures fell above the historic ranges near the surface and between 25 and 50db for some near-shore casts; this is considered real and has been reported by others.

##### Initial Recalibration

COMPARE indicates that the primary salinity is low by from 0.0009psu to 0.0022psu. For bottles from 500db down the result was -0.0017 and for 2500db down it was -0.0009psu. In March and April it was found to be high by 0.0004 and low by 0.0003 which suggests a slow drift to lower values. Recalibration by adding 0.001psu looks like a reasonable choice.

File 2004-10-recal1.ccf was prepared to apply an offset of +0.001 to the primary salinity and to recalibrate the dissolved oxygen using the results of section 11; it was applied to the EDT, MRG and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for salinity and DOX and the calibration was found to have been done correctly. (See 2004-10-dox-comp2.xls.)

##### Final Dissolved Oxygen comparison

Another run of COMPARE was made using metre-averaged, thinned downcast files. The differences versus pressure show an offset of about 0.026 ml/l with the CTD reading a little high and no significant pressure-dependence. When the differences are plotted against DOX values the results are similar. When plotted against cast number, there appears to be some time dependence. But as noted above, this is probably due to geographic variations.

The offset does not seem likely to be due to an imperfect alignment since a small error would cause little change at depth. It is likely due to the response time being a little slow, so the sensor gives slightly higher values than expected on the downcast. Only the downcast files are affected by this, the bottle files do not require recalibration because the sensor had time to equilibrate. During 2004-04, -05 and -07 when the same instrument was used an offset of -0.03ml/l was found appropriate. This will be applied to this data set too.

The COR1 files and the thinned downcast files were recalibrated by subtracting 0.03ml/l using file 2004-10-recal2.ccf. (Output: COR2)

COMPARE was rerun using the recalibrated thinned files and the average difference was +0.0027ml/l. (See 2004-10-DOX-comp4.xls)

##### Special Fluorometer Processing

The COR2 files were clipped to 100db and put through REMOVE and HEADEDIT to produce files FCTD. The FCTD files were written on a CD-ROM for the use of Angelica Peña.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR2 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000

Minimum bin value =   .000

Average value will be used.

Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

##### Replacement of bad data with pad values

CLEAN was run to replace the dissolved oxygen data for pressures greater than 1500db. (Output: CLN1)

For casts 27, 32 and 89 the transmissivity channel was blanked below 1500, 3400 and 1400db respectively. (Output: CLN2) The *.CLN2 files were copied to *.CLN1.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Primary, Salinity:T0:C0, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag channels were removed from all casts. (Output: *.REM)

The PAR channel was removed from casts 1-5, 9, 10-16, 19-38, 42-55, 83-102. (Output: REM2 was copied to *.REM.)

HEADER EDIT was used to fix formats and channel names and to add the following comments:

The fluorescence and transmissivity data are nominal and unedited

except that some records were removed in editing T and S.

The dissolved oxygen data was removed for pressures greater than

1500db because the sensor is unreliable at depth. The DO data in

the CTD files should be considered:

•	±0.2ml/l from 0db–150db.

•	±0.1ml/l below 150db – 300db

•	±0.05ml/l from 300 – 1500db

For casts 27, 32 and 89 it was noted in the header that the transmissivity channel was blanked below 1500, 3400 and 1400db respectively.

The Standards Check routine was run and HEADEDIT adjusted until no further problems were found. The final files were named CTD.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T and S profiles. Separate profile plots were prepared with PAR, DOX, FL and Transmissivity versus pressure.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Oxygen:Voltage:SBE, Status:Pump, Descent_Rate, Altimeter and Flag.

HEADER EDIT was run to fix formats and units and to add a comment about quality flags and analysis methods. Standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

##### 25. Check of altimeter readings in headers

A list was made of the altimeter reading in the headers of these files to ensure the information is reasonable. It was found that for cast #53, noise has probably been interpreted as a signal. And for casts #56 through #82 it appears that the altimeter either was not mounted or no signal was recorded; since it was included in the configuration file a value of “0” was recorded and that appears in the headers. For casts #53 through #82 the header entry was removed from the final files using a text editor.

Only one CHE file exists in that range; cast 2004-10-0053.che was edited to remove the altimeter header.

Particulars

7. PAR on

9. Surface data on upcast may show “funny numbers” due to confusion re heave compensation “surface”

12. Sample #50 believed to be from a different Niskin bottle

12. Station name wrong in header – should be P8.

17. PAR on

19. Initial partial cast with pumps off followed by full cast with pumps on.

24. CTD down to 200, up to 175 and then down again to close bottle.

27. Transmissivity replaced with pad values below 1500db.

32. Transmissivity replaced with pad values below 3400db

35. Transmissometer changed before this cast

40. PAR on

42. CTD returned to surface from 45m because pump was not on. Full cast then run.

51-82. PAR on

79. CTD down to 33db with pumps off; returned to surface and pumps turned on for full cast.

89. Transmissivity replaced with pad values below 1400db

90. Long stay at surface due to heave comp. problem

96. 7 bottles closed to test Niskin #5, but only Niskin #1 sampled

102. Drift in temperature sensors noted ±0.04.

103-109. PAR on

104. Calibration cast without pump

109. Bottom bottles tripped quickly to get off bottom – steep-sided inlet

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-10

Dates: Start: June 1, 2001 End: June 18, 2004

Location: Gulf of Alaska

Vessel: John P. Tully

Party Chief: Robert M.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0585 | Yes | Yes |

Institute of Ocean Sciences

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0585

Cruise ID#:		2004-10

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2023 | 20/12/03 | Factory “ |  |  |
| Conductivity | 1763 | 19/12/03 | “ |  |  |
| Secondary Temp. | 2095 | 05/09/03 | “ |  |  |
| Secondary Cond. | 1764 | 05/09/03 | “ |  |  |
| Transmissometer | 333DR/ 498DR | 06/01/04 07/04/04 | IOS |  |  |
| Dissolved Oxygen | 0047 | 13/01/04 | Factory |  |  |
| Altimeter | 1024 | ? | ? |  |  |
| Fluorometer | 2229 | July 01 | IOS |  |  |
| Pressure Sensor | 77511 | 13/03/00 | Factory |  |  |