## REVISION NOTICE TABLE

| ### DATE | ### DESCRIPTION OF REVISION |
| --- | --- |
| 7-Jul-2005 | Cast 0012.CHE - Oxygen values for samples 77 to 81 were missing from original OXY file. Values from rosette log were inserted using a text editor. |

## PROCESSING NOTES

Cruise: 2004-11

Agency: OSAP

Location: B.C. Inlets

Project: B.C. Inlets

Party Chief: Spear D.

Platform: Vector

Date: April 14, 2004 – April 23, 2004

Processed by: Germaine Gatien

Date of Processing: 7 May 2004 – 20 May 2004

Number of original CTD casts: 	57

Number of casts processed: 57

Number of rosette casts: 39

# INSTRUMENT SUMMARY

A SeaBird Model SBE 911+ CTD (#0550) was mounted with a Chelsea transmissometer (#498DR), a Seabird Dissolved Oxygen Sensor (#0615), Altimeter (#1024) and a Seapoint Fluorometer (#2229) with a 10X cable. The deck unit was a model 911 (#0424). The salinometer was  Portasal model 8410 (#59724).

# SUMMARY OF QUALITY AND CONCERNS

The sensor numbers were entered into the first page of the log incorrectly; they were then erased but not replaced. The serial number of the CTD was not recorded. For many casts the rosette sheets were only partly filled in; no indication is given of what sampling was done. This can be a problem when we try to track down errors and there is no way to determine if any bottle data were missing.

The pressures may be low by up to 1db. It is recommended that a pressure test be done once per cruise.

The CTD hit bottom during cast #2; the altimeter was working well and the data shows that the descent rate was >0.6m/s when the altimeter indicated that the bottom was within 2m. Shoaling may have also contributed to the problem, but paying attention to the altimetry or the sounder might have at least ensured a less violent touch-down.

The secondary temperature data was very spiky for many of the casts and the secondary conductivity had huge excursions to unbelievable values; the return to normal values was gradual and slow, rendering large segments of data unusable.

The transmissivity, dissolved oxygen and fluorescence are unedited, except where records were removed in the editing of temperature and salinity.

The dissolved oxygen data in the CTD files should be considered:

- ±0.2ml/l in the top 150m

- ±0.1ml/l from 150m down

# PROCESSING SUMMARY

##### Seasave

This step was completed at sea; the raw data files have extension DAT.

##### Preliminary Steps

The Log Book was obtained.

Salinity and dissolved oxygen bottle data were obtained. Flags and comments had been added to the oxygen files as needed. The nutrient data was not yet available.

There was no chlorophyll sampling.

The cruise summary sheet was completed.

The configuration files were obtained and the calibration constants were checked. The only error found was in the date of the pressure calibration. That was corrected and the resulting file named 2004-11-CTD.con.

The history of the conductivity and pressure sensors were found. The dissolved oxygen sensor is new.  The pressure offset for this CTD drifted through 2003, so checks will be done on whether an offset is needed.

Note: The transmissivity calibration information was not available at the time of conversion. It was later obtained and it was found that the date was wrong. This was changed in 2004-11-CTD.con after conversion. A comment will be placed in the final files to correct the date. The co-efficients were correct.

##### Conversion of Raw Data

A few files were converted to check on what pressure offset is needed. The pressures at the surface were often <0 while salinity shows “in-water” values. The lowest such pressure noted was -0.4db. Normally we would expect the soaking period pressure to be about 1.5db, but Sheila Toews says that some people were starting the pumps when the CTD was in the water, but not the whole rosette. So lower than usual values may be correct. For 2003-02 an offset of +0.5db was applied. For 2003-26 the offset was +1.5db, but there were some doubts about this choice. The configuration files was edited to add an offset of +0.5db to the pressure channel; that ensured that there were no negative pressures. The pressure may still be a little low, but the error should be no more than 1db.

The pressures for cast #1 do not fit this pattern, but possibly this is because it was the first use of the CTD in 9 months. The CTD was put into the water with the syringes on. It was brought out, the syringes removed and then put back in, without starting a new file. The pressures when it was presumably on deck are about -0.3db which seems too high. Adding the +0.5db offset results in on-deck pressures of about +0.2db. Given that the data in other casts contradicts this result, it will be ignored as far as deducing an offset goes.

The raw data were converted using the configuration file 2004-11-ctd.con. The file name for cast #19 was corrected to standard format.

A few casts were examined and all expected channels are present. The up and downcasts are similar and the pairs of sensors are reasonably close. However, there is severe spiking in the secondary temperature for some casts.

Rosette files were converted using a start time of -4s and duration of 5s. Usually we use -2s as the start time, but there have been problems with other VECTOR cruises due to the CTD starting up so quickly that upcast data were captured in the rosette file.

##### WILDEDIT

Tests were run on cast #4 using WILDEDIT to remove the spikes, but this was unsuccessful.

Program WILDEDIT was used to remove spikes from the pressure channels only.  Parameters used were:

Pass 1    Std Dev = 2;	Pass 2    Std Dev = 5;	Points per block = 50

##### ALIGNCTD

The deck unit model is one of the older ones that was believed to not advance the secondary conductivity. However, recent results look as though the secondary was advanced, so there may have been some change to the equipment or software. In any case, the alignment will be fine-tuned later using SHIFT. ALIGNCTD was not run.

##### CELLTM

Tests were run on a few casts with (alpha, 1/beta) set to (0.1,7), (0.01,9), (0.02,9), (0.03,9), (0.02,7), (0.03,7) and (0.0245, 9.5). For one cast there was no noticeable difference and for two casts the best results were with (0.02, 7) for the primary channels. For the secondary channel the same setting improved one cast and had no effect on the other two. Since (0.02, 7) does improve some data and did not make any cast worse, that parameter will be used for both channels.

CELLTM was run using (0.02, 7) on all casts and both channels.

##### DERIVE

Program DERIVE was run twice:

on all casts to calculate primary and secondary salinity.

on all casts to calculate the differences between primary and secondary channels for temperature, conductivity and salinity. These were placed in a test directory and will not be archived.

##### Test Plots and Channel Check

A sample of casts was plotted to check for agreement between the pairs of T and C sensors. Because there was a hint of drift the number of casts examined was increased.

Most of the casts were relatively shallow.

| Cast # | Press | T1-T0 | C1-C0 | S1-S0 | Descent Rate |
| --- | --- | --- | --- | --- | --- |
| 22 | 350 | 0.0009 | -0.0001 | -0.0017 | Steady/Med |
| 22 | 400 | 0.0009 | -0.0001 | -0.0017 | Steady/Med |
| 23 | 350 | 0.00095 | -0.000095 | -0.0018 | Steady/Med |
| 23 | 500 | 0.00085 | -0.0001 | -0.0017 | F.Steady/low |
| 32 | 350 | 0.00100 | -0.000105 | -0.00185 | Steady/med |
| 32 | 500 | 0.00095 | -0.0001 | -0.00175 | Steady/med |
| 57 | 340 | 0.0006 | ~0 | -0.0007 | Steady/med |
| 58 | 300 | 0.00035 | ~0 v.noisy | -0.0007 | Steady/med |

The differences are fairly constant and depth independent below 250db, except for casts #57 and 58. The secondary conductivity had a huge excursion for cast #57 that may explain some differences. And the conductivity was very noisy for cast #58. There are no deep casts between 32 and 57, so no check could be made of whether there was a steady change. The variations may be geographic as casts #57 and 58 are in a different region from casts 22, 23 and 32.

##### Conversion to IOS Headers

The IOSSHELL routine was used to convert SEA-Bird 911+ CNV files to IOS Headers. Three casts had rotated Sea-bird headers that had to be fixed in order to convert properly. (#20, 41 and 42)

CLEAN was run to replace pad values in the Pressure channel using linear interpolation based on scan number.

The rosette files were converted to IOS files.

CLEAN was run to add event numbers and the output file from this was renamed BOT.

All BOT files were plotted and many outliers were found in the secondary data only. It was decided not to edit these files since the secondary salinity was generally very poor and unlikely to be archived. If a choice is later made to use the secondary data, the BOT files should be edited first.

##### Checking Headers

The cruise track was plotted and no problems noted.

A header summary, a header check and cross-reference listing were produced and errors ????found and corrected in the headers of the CLN files.????

The average surface pressure is 1db. As discussed earlier this is lower than usual, but may be correct. The CTD was sometimes started with the sensors in the water but parts of the rosette out. There are no negative surface pressures. No further recalibration will be done.

##### BOTTLE FILE PREPARATION

The salinity analysis spreadsheets was edited to add event numbers and sorted on sample number; they were converted to individual SAL files. There were no flagged values.

One sample was removed from the *.SAMAVG file for cast #15 and another from #34 because no sampling was done from those bottles.

A first run of COMPARE turned up some severe outliers. Investigation of these problems led to the following discoveries and actions:

- Cast #5 – The bottle at 275m is higher than the CTD by 0.047 psu. This sample was flagged “d” with a comment noting it was a severe outlier in COMPARE and that there was no liner on the sample bottle.

- Cast #29 – There is only 1 salinity sample noted on the rosette sheet and that is from 50db. That sample compares reasonably well with the CTD. But the salinity sheet shows 2 other samples from this cast and they are severe outliers in COMPARE. It is also interesting that they are said to be from 75db and 100db, but the values are almost identical. It seems doubtful that they come from this cast. Checking with the original salinity analysis sheet showed that the two samples had no sample numbers on the labels, but were labelled “station SE05”. The spreadsheet was changed so those two samples were not assigned to cast #29.

- Cast #24 – This cast is from station “SE05” and has salinity at the bottom (Niskin #1) that is close to the two bottles labelled “station SE05”. The CTD salinity from the 2nd Niskin bottle is lower than those bottles by about 0.002psu, so neither sample is likely to be from that level. So one sample was assigned to bottle #1, sample #168, cast #24 and flagged “c” with a comment to indicate uncertainty about the sample number. The other value was entered in the header comment only.

- Cast #34 - There is no salinity sampling noted on the rosette sheet. Two samples are given in the salinity analysis and are mild outliers in COMPARE, but given they are from 75 and 100db it is not clear that there is an error. They will be flagged “c”

- Cast # 36 – There is no salinity sampling noted on the rosette sheet. One sample is given in the salinity analysis and is an outlier in COMPARE but given that it is from only 67db it is not clear that this is an error. It will be flagged “c”

After these corrections the SAL files were re-merged with the SAMAVG files to produce MRG1 files.

The dissolved oxygen files (*.add) have comments in the headers to explain flagged values. The flag in cast #5 was entered in the wrong column and the comment format was wrong. These errors were corrected. The file 2004-11-0008.add was changed to 2004-11-0009.add since that is where those samples were really collected. The ADD files were merged with the MRG1 files. Running COMPARE turned up a problem in file 2004-11-0027.ADD; the  sample numbers were wrong – one was repeated and those that followed got out of order. This was corrected and the merge rerun.

There was no chlorophyll sampling.

The nutrient data was not available and will need to be added later.

The BOT files were averaged to enable an addsamp file to be created. This file was edited to add sample numbers taken from the rosette sheets. In doing that it was discovered that sample #285 had been used twice, for bottle #7, cast #46 and for bottle #1, cast #47. The former (cast #7) was renamed #9285 in the addsamp file and in the ADD file; only DO sampling was done for that bottle.

The sample numbers were added to the BOT files to create SAM files, which were averaged to create SAMAVG files.

The SAL and ADD files were merged with SAMAVG in two steps. (Output: MRG1 and MRG3)

##### 11. COMPARE

Salinity

COMPARE was run. A number of errors were found in the salinity bottle data, as indicated in the previous section. When one flagged value and all values from above 200db were excluded, the primary salinity was found to be low by 0.0003 and the secondary low by 0.0012psu. For the deepest sample at about 600m the primary is low by <0.0001psu and the secondary low by 0.0017psu. There is a lot of scatter. While there is some suggestion of drift with time, the trend can probably be explained as due to geography as different water types were sampled. Both the primary and secondary differences vary in the same way.

Dissolved Oxygen

There was a lot of DO sampling from this cruise. COMPARE was run and plots were prepared of differences between CTD and titrated values versus pressure, DO and cast number. When values from the top 10db were excluded the differences were quite flat against pressure and time. The most useful fit is probably that versus DO value. A few outliers were identified from that plot. Studying the results led to the discovery of an error in the ADD file for cast #27. After that was fixed the only serious outlier was for a sample from cast #5 that had been flagged by the analyst. The fit versus DO is:

DOX (Titrated) = 1.0212* DOX (CTD) -0.0168

This is a new DO sensor. (See 2004-05-dox-comp1.xls)

The above fit will be used to recalibrate the CTD DOX channel.

##### 12. SHIFT

Conductivity

During 2003-26, -27, -31 and -32 when the same CTD and similar attached instruments were used, the primary conductivity was shifted by -0.2 records. Tests were run on several casts, and the best results were settings between 0 and -0.3 records.

All casts were put through SHIFT using -0.2 records for the primary sensor.

Tests were not done for the secondary channel because the data was very poor and unlikely to be selected for the archive.

Fluorescence

To find what shift is needed for the fluorescence, upcast and downcast profiles were examined to determine the vertical offset of the temperature and fluorescence traces. The differences between these two offsets is treated as a measure of how much the fluorescence needs to be shifted. The “excess” offset for the fluorescence was divided by the averaged descent/ascent rate and divided by 2(since the shift will be applied to both up and downcast) to find the shift (in seconds) to remove that offset. For this cruise the results varied from +1s to +2s. Overall a shift of 1.5 s (+36 records) looked best so this was applied to all casts. For all previous cruises a setting of +24 records was used.

Dissolved Oxygen

Tests were done on a few casts to study the alignment of the oxygen data. A first estimate was made by comparing the offset between upcast and downcast for the DO trace and Temperature trace for a cast without bottle stops. The DO trace appears to have a delay of about 4s or about 100 records. Tests were done shifting the DO by +70 to +130 records and the best results were found using +120 records or 5s.

##### DELETE

The following DELETE parameters were used:

Surface Record Removal: Last Press Min and Low Salt

Maximum Surface Pressure (relative): 10.00

Minimum Surface Salinity: 5.0

Surface Pressure Tolerance: 1.0                  Pressure filtered over 15 points

Swells deleted. Warning message if pressure difference of 2.00

Drop rates <   0.30m/s (calculated over 11 points) will be deleted.

Drop rate applies in the range    10db to 10db less than the maximum pressure

Sample interval = 0.042 seconds. (taken from header)

COMMENTS ON WARNINGS: There were no warnings.

##### DETAILED EDITING

The primary sensors were chosen for further processing because of the severe spiking in the secondary.

Page plots were produced using (T0,S0). These were used to guide the editing. With a few exceptions the descent rate was fairly low, but was usually steady enough that it was not a major problem.

Casts #10 and 34 could not be plotted until the first record was removed from the DEL file. The files were then put through CLEAN and renamed DEL.

The following cast required no editing: #11.

The following cast required extensive editing: #1.

All other casts required only light editing. Note was made of the editing details in the headers of the relevant files.

##### Other Comparisons

Previous experience with these sensors – The primary conductivity sensor was last used during 2003-31 in September 2003 when the salinity was found to be high by about 0.0035psu and during 2003-37 in August when it was high by 0.004psu. The secondary conductivity was used during 2003-31 and was found to be high by 0.0032. The dissolved oxygen sensor has not been used before.

Historic ranges – Local climatology was available for only one cast; all data fell within the historic range for that cast.

Repeat Casts – There were no repeat casts.

##### Initial Recalibration

No recalibration is needed for the primary salinity.

File 2004-11-recal1.ccf was prepared to recalibrate the dissolved oxygen using the results of section 11 (corrected DOX = 1.0212* DOX (CTD) -0.0168); it was applied to the EDT, MRG3 and SAM files. (Output: COR1, MRGCOR1, SAMCOR1)

COMPARE was rerun for DOX and the corrections were found to be made correctly. (See 2004-11-dox-comp2.xls.)

### Final Dissolved Oxygen Recalibration

The first recalibration corrects for the in situ errors in the sensor and the SHIFT routine corrects for transit time, but in the past there have sometimes been significant errors due to response-time problems. To check for such errors the downcast CTD data (after SHIFT and CALIBRATE) was compared with the bottles from the upcast. A set of downcast files were prepared by metre-averaging the COR1 files and thinning the data to the depths generally used for bottles. COMPARE was run comparing those files with the bottle data in the MRG files. The fit of differences versus pressure was found excluding a few outliers and values from above 10db. The CTD is higher than the bottles by an average of +0.04ml/l with some pressure dependence. If only the bottom bottle for each cast is considered the average is -0.05ml/l. Since the difference in time between the downcast CTD and bottle data is least for the deepest bottle, that might seem the best estimate. However, the DOX value tends to be lowest there so that may not be the best choice. Given the roughness of the method and the lack of a clear pattern, no recalibration will be attempted for what is, at worst, a small error. Time-response is not a serious problem for this instrument. (See 2004-11-dox-comp3.xls.)

##### Special Fluorometer Processing

The COR1 files were clipped to 100db and stored in a separate directory for the use of Angelica Peña. They were put through REMOVE and HEADEDIT to produce files FCTD and saved to a CD-ROM.

A median filter, fixed size=11, was applied to the fluorescence channel in the COR1 files to reduce spikiness. One cast was examined before and after this step and showed that the filter was effective. (Output: FIL)

##### BIN AVERAGE of CTD files

The following Bin Average values were applied to the FIL files (output AVG):

Bin channel = pressure

Averaging interval = 1.000	Minimum bin value =   .000

Average value will be used.	Interpolated values are NOT used for empty bins.

After averaging, page plots were examined on screen and no further editing was deemed necessary.

A spreadsheet was prepared with altimeter readings taken from the header. This was checked to ensure that the values were reasonable. Cast #17 had a value of 28.5m. Looking at the full data file, a value of about 5m looks appropriate. There was a lot of noise in the altimetry near the bottom with high values, but not high enough to be excluded from the median calculation. The altimeter reading was removed from the header and a note put in the REMARKS table of the LOCATION section to explain why.

##### Final Plots

THIN and DERIVE were run to obtain values for tables and page plots were prepared using the edited data and displaying T & S profiles. Separate profile plots were prepared with Temperature, Transmissivity, DOX and Fluorescence versus pressure.

##### FINAL CTD files steps (REMOVE and HEADEDIT)

The Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag channels were removed from all casts.

HEADER EDIT was used to add a comment (including a correction to the transmissometer calibration date) using file 2004-11-hdr.txt and to fix formats and units.

The Standards Check routine was run and HEADER EDIT adjusted and rerun until no further problems were found.

The final files were named CTD.

##### Final Bottle Files

The MRGCOR1 files were put through CLEAN to remove the SeaBird headers, SORT to put the data in order of increasing pressure and REMOVE to remove Scan_Number, Temperature:Secondary, Salinity:T1:C1, Conductivity:Primary, Conductivity:Secondary, Status:Pump, Descent_Rate, Oxygen:Voltage:SBE, Altimeter and Flag.

HEADER EDIT was run to fix channel names and formats and to add a standard comment about quality flags and a comment about the date of the transmissivity calibration. The standards check was run on all files and HEADEDIT adjusted until all format problems were resolved. The final files were named CHE.

##### Producing final files

A cross-reference listing was produced.

The sensor history was updated.

Particulars

1. CTD went into water with syringes on. It was brought back on deck and returned to water without starting a new file.

2. CTD hit bottom

5. Secondary temperature oscillating in value.

8. Identified in Daily Log and on rosette sheet as rosette cast. In fact, the rosette cast was cast #9.

20. Daily log notes problem with NMEA time and date; Lat/Long/Time were rotated in the Sea-Bird headers. This was corrected to enable conversion.

26. Bottom 4 bottles had strong H2S odour.

27. Bottom 3 bottles had strong H2S odour.

29. Bad secondary temperature.

38. Significant ship drift.

39. Significant ship drift.

41 & 42. According to Daily Log the NMEA date was wrong, but in fact the problem was rotated Sea-Bird headers. The log also notes that the deck computer was 3 minutes slow.

43. GPS time ok.

47. Bottle #1 had H2S odour.

48. Bottle #1 had H2S odour; mud in water, bottle #3 misfired at bottom.

55. Spikes in secondary salinity in upcast.

# Institute of Ocean Sciences

CRUISE SUMMARY

Cruise ID#: 2004-11

Dates: Start: April 14, 2004 End: April 23, 2004

Location: B.C.Inlets

Vessel: Vector

Party Chief: Spear D.

| CTD# | Make | Model | Serial# | Used with Rosette? | CTD Calibration Sheet Competed? |
| --- | --- | --- | --- | --- | --- |
| 1 | SEABIRD | 911+ | 0550 | Yes | Yes |

# CTD Calibration Information

Make/Model/Serial#:	SEABIRD/911+/0550

Cruise ID#:		2004-11

| Calibration Information |  |  |  |  |  |
| --- | --- | --- | --- | --- | --- |
| Sensor |  | Pre-Cruise |  | Post Cruise |  |
| Name | S/N | Date | Location | Date | Location |
| Temperature | 2038 | 22/04/03 | Factory |  |  |
| Conductivity | 2173 | 24/04/03 | “ |  |  |
| Secondary Temp. | 2968 | 22/04/03 | “ |  |  |
| Secondary Cond. | 1729 | 24/04/03 | “ |  |  |
| Transmissometer | 498DR | 07/04/04 | IOS |  |  |
| Dissolved Oxygen | 0615 | 09/03/04 | Factory |  |  |
| Fluorometer | 2356 |  | IOS |  |  |
| Pressure Sensor | 75636 | 06/04/99 | Factory |  |  |