Information on post-cruise calibration provided by Bon van Hardenberg

temperature sensor   SBE03-4115

new 21c March 2001

recalibrated December 31 2002: drift (degC/yr) +0.00003

recalibrated 15 May 2004: drift= +0.00017 (pre-cruise)

recalibrated 10 December 2004: drift= +0.00026.

conductivity sensor  SBE04-2607

new 11 April 2001

recalibrated 31 December 2002: drift= 0.0000

recalibrated 14 May 2004: drift= -0.0001

post-cruise calibration 20 January 2005: drift= 0.0000 but cracks in the glass were found so the sensor was repaired by replacing the glass-with-electrodes, and a new set of calibration coefficients as pre-cruise dated 20 January 2005 was supplied (these coefficients were entered used to update the CON file used during the 2005-05 cruise.